373 | CN121196724A | Pulmonary nodule diagnosis and treatment system based on multi-modal fusion and self-adaptive regulation and control | The invention discloses a pulmonary nodule diagnosis and treatment system based on multi-modal fusion and adaptive regulation and control, and relates to the technical field of medical equipment and clinical diagnosis and treatment. The system comprises a multi-mode image assembly, an intelligent puncture execution assembly, a self-adaptive ablation assembly and a central control unit. The multi-modal image assembly obtains an ultrasonic image, a CT image and an OCT image of the lung, the central control unit fuses the ultrasonic image, the CT image and the OCT image of the lung, plans an initial puncture path, controls the intelligent puncture execution assembly to drive the puncture needle to conduct lung puncture, and dynamically regulates and controls the motion trail of the puncture needle; after the puncture needle reaches the pulmonary nodule position, the central control unit automatically adjusts ablation parameters and controls ablation treatment of the self-adaptive ablation assembly until ablation is finished. Accurate positioning, intelligent puncture and efficient ablation of pulmonary nodules can be achieved. | 1 . A lung nodule diagnosis and treatment system based on multimodal fusion and adaptive control is characterized in that the system comprises: a multimodal imaging component, an intelligent puncture execution component, an adaptive ablation component, and a central control unit; the multimodal imaging component is used to acquire preoperative ultrasound, CT, and OCT images of the lung; the central control unit is used to fuse the preoperative ultrasound, CT, and OCT images of the lung to construct a three-dimensional visualization model of the lung including the location of the lung nodules, and to plan the initial puncture path based on the three-dimensional visualization model of the lung; and to determine the initial puncture parameters based on the planned initial puncture path; the intelligent puncture execution component is used to drive the puncture needle to perform lung puncture according to the initial puncture parameters under the control of the central control unit, and to sense the puncture force and spatial posture information of the puncture needle in real time during the puncture process; the multimodal imaging component is also used to acquire real-time ultrasound, real-time CT, and real-time OCT images of the lung during the puncture process; The central control unit is also used to dynamically adjust the puncture path and determine the puncture parameters after real-time adjustment by combining real-time sensed puncture force, spatial posture information, real-time ultrasound images, real-time CT images and real-time OCT images; the intelligent puncture execution component is also used to control the movement trajectory of the puncture needle according to the real-time adjusted puncture parameters under the control of the central control unit until the puncture needle reaches the lung nodule.The adaptive ablation component is used to initiate ablation treatment on the lung nodule after the puncture needle reaches the location of the nodule, and to monitor the temperature and tissue impedance of the ablation area in real time during the ablation treatment. The central control unit is also used to automatically adjust the ablation parameters according to the temperature and tissue impedance of the ablation area, and to control the ablation treatment of the adaptive ablation component according to the adjusted ablation parameters until the ablation is completed. 2 The lung nodule diagnosis and treatment system based on multimodal fusion and adaptive control according to claim 1 is characterized in that the multimodal imaging component includes: an ultrasound imaging component, a CT image import interface, and an optical coherence tomography (OCT) scanning component; the ultrasound imaging component, the CT image import interface, and the OCT scanning component are all connected to the central control unit; the ultrasound imaging component uses an ultrasound probe to acquire ultrasound images of the lungs and transmits them to the central control unit; The CT image import interface is used to import CT images of the patient's lungs and transmit them to the central control unit; the optical coherence tomography (OCT) component is used to acquire OCT images of lung tissue and transmit them to the central control unit. 3 The lung nodule diagnosis and treatment system based on multimodal fusion and adaptive regulation according to claim 1, characterized in that the intelligent puncture execution component includes: an electric puncture needle assembly, a multidimensional force sensor array, and a micro inertial measurement unit; the multidimensional force sensor array is distributed on the surface of the puncture needle; the micro inertial measurement unit is disposed on the puncture needle; the electric puncture needle assembly, the multidimensional force sensor array, and the micro inertial measurement unit are all connected to the central control unit; the electric puncture needle assembly is used to drive the puncture needle to perform lung puncture according to the initial puncture parameters under the control of the central control unit; the multidimensional force sensor array is used to sense the puncture force of the puncture needle in real time during the puncture process and transmit it to the central control unit; the micro inertial measurement unit is used to monitor the spatial attitude information of the puncture needle in real time during the puncture process and transmit it to the central control unit. 4 The lung nodule diagnosis and treatment system based on multimodal fusion and adaptive regulation according to claim 1 is characterized in that the adaptive ablation component includes: an ablation electrode, a temperature sensor array, and an impedance monitoring unit; the temperature sensor array is distributed around the ablation electrode and within the lung tissue; the ablation electrode, the temperature sensor array, and the impedance monitoring unit are all connected to the central control unit; The ablation electrode is used to ablate lung nodules according to the ablation parameters sent by the central control unit; the temperature sensor array is used to monitor the temperature of the ablation area in real time during the ablation treatment and transmit the data to the central control unit; the impedance monitoring unit is used to detect the tissue impedance of the ablation area in real time during the ablation treatment and transmit the data to the central control unit. 5 The pulmonary nodule diagnosis and treatment system based on multimodal fusion and adaptive regulation according to claim 1 is characterized in that the pulmonary nodule diagnosis and treatment system based on multimodal fusion and adaptive regulation further includes: a human-computer interaction interface; the human-computer interaction interface is connected to the central control unit; the human-computer interaction interface is used to input basic patient information and surgical requirements, display ultrasound images, CT images and OCT images, and display the planned initial puncture path and the real-time status of the ablation process; the human-computer interaction interface is also used to receive the doctor's confirmation instruction after displaying the planned initial puncture path and generate a puncture start instruction. 6 The lung nodule diagnosis and treatment system based on multimodal fusion and adaptive regulation according to claim 1 is characterized in that, in terms of fusing preoperative ultrasound images, CT images, and OCT images of the lung, the central control unit specifically includes: removing speckle noise from the ultrasound image using a nonlocal mean filtering algorithm, and then enhancing the contrast of the ultrasound image using a histogram equalization method to obtain a preprocessed ultrasound image; performing grayscale normalization on the CT image and removing salt-and-pepper noise from the CT image using median filtering to obtain a preprocessed CT image; using a bilateral filtering algorithm to denoise and enhance the OCT image, and adjusting the brightness and contrast of the OCT image to obtain a preprocessed OCT image; using a deep learning-based convolutional neural network to extract features from the preprocessed ultrasound image to obtain a feature map of the lung nodule; using a residual network model to extract features from the preprocessed CT image to generate a feature representation of the CT image; and using a method based on local feature descriptors to extract features from the preprocessed OCT image to obtain a feature representation of the OCT image. Based on the feature maps of lung nodules, the feature representations of CT images, and the feature representations of OCT images, coarse registration is performed on ultrasound images, CT images, and OCT images. A registration algorithm based on mutual information is then used to perform fine registration on the coarsely registered ultrasound images, CT images, and OCT images to obtain the registered ultrasound images, CT images, and OCT images.Wavelet transform is used to decompose the registered ultrasound, CT, and OCT images into low-frequency and high-frequency sub-band images. Based on a weighted average fusion rule, the low-frequency sub-band images from the ultrasound, CT, and OCT images are fused using a weighted average to obtain the fused low-frequency sub-band image. Then, a region energy-based fusion rule is used to fuse the high-frequency sub-band images from the ultrasound, CT, and OCT images to obtain the fused high-frequency sub-band image. Finally, inverse wavelet transform is performed on both the fused low-frequency and high-frequency sub-band images to reconstruct the fused image. 7 The lung nodule diagnosis and treatment system based on multimodal fusion and adaptive regulation according to claim 1, characterized in that, in planning the initial puncture path according to the three-dimensional visualization model of the lung, the central control unit specifically includes: Locate the lung nodule in the 3D visualization model of the lung and mark the contraindication area; screen candidate needle insertion points on the body surface and generate multiple straight lines between the candidate needle insertion points and the lung nodule location to form multiple initial straight line trajectories; perform collision detection on each initial straight line trajectory in the contraindication area, delete the initial straight line trajectory that crosses the contraindication area, and obtain multiple candidate straight line trajectories; select the candidate straight line trajectory with the shortest path distance under the puncture angle from the multiple candidate straight line trajectories as the planned initial puncture path. 8 The lung nodule diagnosis and treatment system based on multimodal fusion and adaptive regulation according to claim 1 is characterized in that, in determining the initial puncture parameters according to the planned initial puncture path, the central control unit specifically includes: determining the initial puncture parameters, including needle insertion speed, needle insertion angle, and needle insertion depth; layering the initial puncture path according to tissue type to obtain the layered initial puncture path; the layered initial puncture path includes the initial puncture path of the skin layer and the initial puncture path of the lung parenchyma layer; predicting the future displacement of the lung nodule through real-time respiratory phase combined with a Kalman filter, and dynamically correcting the target point position on the lung nodule according to the predicted future displacement of the lung nodule; discretizing the layered initial puncture path into multiple spatial nodes; dividing the time step according to the respiratory cycle and determining the time step of each spatial node to obtain the spatiotemporally matched path node; constructing a mechanical differential equation for each spatiotemporally matched path node, and determining the needle bending amount of each path node through finite element simulation. Based on the principle that the needle insertion speed in the initial puncture path of the skin layer is greater than the speed threshold, the needle insertion speed in the initial puncture path of the lung parenchyma layer is less than the speed threshold, and the change in needle insertion angle between adjacent path nodes is less than the angle deviation threshold, a sequential quadratic programming algorithm is used to solve for the optimal needle insertion speed and the planned needle insertion angle for each path node; the planned needle insertion angle is corrected by the amount of needle bending at each path node to obtain the optimal needle insertion angle.The optimal needle depth for each path node is determined based on the optimal needle advance speed of each path node and the time step of each spatial node. 9 According to claim 1, the pulmonary nodule diagnosis and treatment system based on multimodal fusion and adaptive regulation is characterized in that the central control unit, in combining real-time sensed puncture force, spatial posture information, real-time ultrasound images, real-time CT images, and real-time OCT images, dynamically adjusts the puncture path and determines the real-time adjusted puncture parameters, specifically including: determining the initial puncture parameters, including needle insertion speed, needle insertion angle, and needle insertion depth; determining the real-time spatial coordinates of the puncture needle based on the real-time sensed spatial posture information, real-time ultrasound images, real-time CT images, and real-time OCT images; determining the distance deviation and angle deviation based on the real-time spatial coordinates of the puncture needle and the initial puncture path; when the distance deviation is less than or equal to a distance threshold and the angle deviation is less than or equal to an angle threshold, determining the needle insertion angle adjustment value according to the formula a<sub>a</sub>/a<sub>b</sub> + g; where a<sub>a</sub> is the needle insertion angle adjustment value, a<sub>b</sub> is the first proportional coefficient, a<sub>b</sub> is the distance deviation, a<sub>c</sub> is the second proportional coefficient, and a<sub>c</sub> is the angle deviation; when the distance deviation is greater than the distance threshold or the angle deviation is greater than the angle threshold, determining the needle insertion angle adjustment value according to the formula a<sub>c</sub>/a<sub>b</sub> + g; Δa = min(3 x VZJ + Storagex"2,]o. ), determines the adjustment value of the insertion angle of the puncture needle; where, small is the third proportionality coefficient, and h is the fourth proportionality coefficient; Based on the real-time spatial coordinates of the puncture needle at adjacent moments, the actual insertion speed of the puncture needle is determined. If the actual insertion speed of the puncture needle is not equal to the insertion speed in the initial puncture parameters, the intelligent puncture execution component is adjusted to make the actual insertion speed equal to the insertion speed in the initial puncture parameters. When the real-time sensed puncture force is greater than the upper limit of tissue force, the insertion speed is reduced to 30% of the insertion speed in the initial puncture parameters, the insertion is paused for 0.5 seconds, and the multimodal imaging component is called simultaneously to confirm whether it is bone or calcification. If it is confirmed to be bone or calcification, an avoidance path is planned with the current actual position of the puncture needle as the starting point and the real-time position of the lung nodule as the target point. When the real-time perceived fluctuation of the puncture force exceeds the force fluctuation threshold within a preset time period, and vascular image signals are acquired through the multimodal imaging component, the needle insertion speed is reduced to 10% of the initial puncture speed in the puncture parameters. The needle insertion angle is adjusted along the direction deviating from the blood vessel, and the needle is paused for 0.3 seconds every 0.1 mm advance. Simultaneously, the CT images acquired through the multimodal imaging component are used to verify in real time whether the needle is far from the blood vessel. Once it is confirmed that the needle is far from the blood vessel, the needle insertion speed is resumed. When the real-time perceived puncture force is within the preset force range and continues to rise, the needle is withdrawn by 0.5 mm to 1 mm, and then puncture is performed at 50% of the initial puncture speed. If the real-time perceived puncture force continues to rise, the puncture is terminated and the doctor is alerted to intervene. The upper limit of the preset force range is less than the upper limit of the tissue force. 10 The lung nodule diagnosis and treatment system based on multimodal fusion and adaptive control according to claim 1 is characterized in that the central control unit automatically adjusts the ablation parameters according to the temperature and tissue impedance of the ablation area, and controls the ablation treatment of the adaptive ablation component according to the adjusted ablation parameters until the ablation is completed, specifically including: determining the Pennes equation as a biological heat transfer model; discretizing and solving the biological heat transfer model using the finite element method according to the temperature of the ablation area to obtain a three-dimensional temperature field distribution; the three-dimensional temperature field distribution includes the temperature value of each discretized grid cell and the temperature of the ablation area at each time step; calculating the temperature field gradient of each grid cell using the second-order central difference method according to the three-dimensional temperature field distribution, and constructing effective ablation boundaries for grid cells with temperature field gradients greater than a preset gradient threshold; fitting the impedance spectrum using the Cole-Cole model according to the tissue impedance of the ablation area, and determining the spatial distribution of the rate of change of tissue resistivity within the ablation area according to the fitted impedance spectrum; constructing a three-dimensional ablation volume model according to the effective ablation boundary and the spatial distribution of the rate of change of tissue resistivity within the ablation area; and determining the current ablation volume according to the three-dimensional ablation volume model. Based on the three-dimensional temperature field distribution, the tissue damage distribution in the ablation region is determined using the Arrhenius equation; based on the tissue damage distribution in the ablation region, the current ablation volume, and the temperature field gradient of each grid cell, an optimization algorithm is used to determine the optimal ablation parameters;The optimal ablation parameters are used as the adjusted ablation parameters, and the ablation treatment of the adaptive ablation component is controlled according to the adjusted ablation parameters. When the preset ablation conditions are met, the ablation is terminated. The preset ablation conditions include the current ablation volume being equal to the target ablation volume, or the risk of carbonization occurring, or the ablation time being greater than or equal to the maximum safe time. The condition for determining the risk of carbonization is that the tissue resistivity is greater than the preset resistivity threshold, and the temperature value of any grid cell is greater than the preset temperature threshold. | 2025 | ZHOUSHAN HOSPITAL | Hospital | NA | https://www.zs-hospital.sh.cn | China | The patent focuses on a lung-nodule diagnosis-and-treatment platform that fuses multimodal images (ultrasound, CT, OCT) and uses an intelligent puncture executor and adaptive ablation component, all coordinated by a central control unit. It addresses the limitations of standalone ultrasound positioning systems for lung nodules, where diagnostic and therapeutic capabilities are fragmented. Image fusion drives both targeting and parameter adjustment. | הפטנט מתמקד בפלטפורמת אבחון וטיפול בקשריות ריאתיות המאחדת תמונות רב-מצביות (אולטרסאונד, CT, OCT) ומשתמשת במבצע ניקוב חכם וברכיב אבלציה מסתגל, הכל מתואם על ידי יחידת בקרה מרכזית. הוא מתמודד עם המגבלות של מערכות מיקום אולטרסאונד עצמאיות לקשריות ריאתיות, שבהן יכולות אבחון וטיפול מקוטעות. איחוד תמונות מניע הן את המיקוד והן את התאמת הפרמטרים. | 1 |  | 1 |  | 1 |  |  |  | 1 | [14 empty] | 1 |  |  | 1 | [5 empty] | ZHOUSHAN HOSPITAL | ZHOUSHAN HOSPITAL | ZHOUSHAN HOSPITAL | ZHOUSHAN HOSPITAL | ZHOUSHAN HOSPITAL | 2025-10-11 | 2025 | 2025-12-26 | 2025 | 2025-10-11 | 2025 | Application | Utility Invention | active | CN202511449738 | 325298555 | [NA] | CN
374 | CN119279752A | Real-time feedback pulse ablation system | The invention provides a real-time feedback pulse ablation system. The system comprises a power supply module; the human-computer interaction interface is used for displaying information fed back by the main control module; the high-voltage pulse module outputs a high-voltage pulse signal to the ablation instrument through the switching module; the impedance detection module is used for detecting impedance information of a target position to the ablation instrument through the switching module; the temperature detection module is used for detecting the temperature of the ablation instrument in real time through the switching module; the main control module controls switching of different functional modules; the main control module is used for displaying the data on a human-computer interaction interface; the electrical stimulation module simulates and outputs sine waves, square waves, triangular waves and other signals to the ablation instrument; the signal acquisition module is used for directionally acquiring signals through an electrode at the other end of the ablation instrument; an ablation instrument; according to the real-time feedback pulse ablation system, control over ablation tissue is improved, the ablation process is effectively monitored, and safety and reliability of the operation process are guaranteed. | 1. A real-time feedback pulse ablation system, characterized in that it includes a power module (1), a high-voltage pulse generating module, an electrical stimulation module (8), a pulse ablation device (10), a switching module (6), and a main control module (7); wherein, the power module (1) is electrically connected to the high-voltage pulse generating module, the impedance detection module (4), and the temperature detection module (5) for module power supply; the high-voltage pulse generating module is configured to send a high-voltage pulse signal to the pulse ablation device (10); the electrical stimulation module (8) outputs analog output signals such as sine waves, square waves, and triangular waves to the outer electrodes of the paired working electrodes of the ablation device (10); the pulse ablation device (10) includes a plurality of ablation electrodes, the ablation electrodes include outer electrodes and inner electrodes, and at least some of the ablation electrodes are configured as working electrodes that receive high-voltage pulse signals and perform ablation treatment on target tissues; The switching module (6) is connected to the control of the main control module (7), switches to different functional modules, and is configured to receive the high-voltage pulse signal sent by the high-voltage pulse generating module under the control of the main control module (7), and select the working electrode and switch the positive and negative electrodes; the main control module (7) displays the data on the human-computer interaction interface (2) through serial communication, collects information from different functional modules, controls the switching module (6), records the data of the detection module and the temperature module, can set and excite the high-voltage pulse module (3), and is configured to receive the impedance value and evaluate the degree of contact between the working electrode and the target tissue. 2.A real-time feedback pulse ablation system according to claim 1, characterized in that at least one of the paired ablation electrodes is the outer electrode. 3.A real-time feedback pulse ablation system according to claim 1, characterized in that the outer electrode is arranged close to the distal end of the electrode carrier, the inner electrode is arranged close to the proximal end of the electrode carrier, and the paired ablation electrodes are the inner electrode and/or the outer electrode. 4.A real-time feedback pulse ablation system according to claim 1 or 2, characterized in that the pulse ablation device (10) further includes an end electrode, and the end electrode and the outer electrode are configured as paired ablation electrodes. 5.A real-time feedback pulse ablation system according to claim 4, characterized in that the outer electrodes located on the same electrode carrier are divided into an outer electrode group, and the end electrodes are paired with the outer electrode group in turn to detect the impedance value between the end electrode and the outer electrode. 6.A real-time feedback pulse ablation system according to any one of claims 1 to 5, characterized in that the pulse ablation system also includes an impedance detection module (4), which detects the impedance information of the target position through a switching module (6) to the pulse ablation device (10), evaluates the degree of adhesion and the degree of ablation, and is configured to select paired ablation electrodes and detect the impedance value between the paired ablation electrodes, and at least one of the paired ablation electrodes is a working electrode. 7.A real-time feedback pulse ablation system according to claim 6, characterized in that the initial impedance value is detected before the working electrode is in contact with the target tissue. 8.A real-time feedback pulse ablation system according to any one of claims 1-3, 5, and 7, characterized in that the impedance detection module ((4)) also includes a first multi-way selection switch, the first multi-way selection switch includes a plurality of first selection switches corresponding one-to-one to and electrically connected to the ablation electrodes, and the first selection switch is configured to selectively open or close under the control of the main control module ((7)) so that the impedance detection module ((4)) is electrically connected to the corresponding paired ablation electrodes. 9.A real-time feedback pulse ablation system according to any one of claims 1-3, 5, and 7, characterized in that the pulse ablation system also includes a signal acquisition module (9), which directionally acquires signals through the inner electrode of the ablation device (10), and feeds back the collected information to the main control module (7) in real time, and displays the waveform in the human-computer interaction interface (2) to judge the ablation Whether the treatment is effective. 10.A real-time feedback pulse ablation system according to any one of claims 1-3, 5, and 7, characterized in that the pulse ablation system also includes a temperature detection module (5), which detects the temperature of the ablation device (10) in real time through a switching module (6), and feeds back to the main control module (7) in real time to avoid damage caused by excessive temperature during the ablation process. 11.A real-time feedback pulse ablation system according to claim 10, characterized in that the switching module (6) includes a first layer combination switch, a second layer combination switch and a third layer combination switch; wherein, the input end of the first layer combination switch is electrically connected to the positive and negative poles of the high-voltage pulse signal output respectively; the input end of the second layer combination switch can be switchably connected to the output end of the first layer combination switch respectively, and the second layer combination switch is configured to perform the switching of the positive and negative poles under the control of the main control module (7); the third layer combination switch includes a plurality of ablation selection switches corresponding to the ablation electrodes one by one, the input end of the ablation selection switch is electrically connected to the second layer combination switch, and the output end of the ablation selection switch is electrically connected to the corresponding ablation electrode, and the third layer combination switch is configured to select the working electrode and the positive and negative of each working electrode under the control of the main control module (7). 12.A real-time feedback pulse ablation system according to any one of claims 1-3 and 5, characterized in that the main control module (7) is also configured to control the high-voltage pulse generating module to pre-discharge according to the discharge mode, and the pulse ablation system also includes: a diagnostic module, electrically connected to the signal acquisition module (9) and the main control module (7), configured to determine whether the discharge of the working electrode is normal based on the collected current and voltage data, and feed back to the main control module (7). | 2024 | ZHULING SHANGHAI MEDICAL DEVICES | Medical and Healthcare | NA | NA | China | The patent focuses on a real-time-feedback pulsed ablation system that combines a power module, HMI, high-voltage pulse module with a switching unit, plus impedance and temperature detection at the target. It addresses the safety problems of traditional RF treatment in maze IV atrial fibrillation surgery (uncontrollable range, thrombus risk). Closed-loop signals from the catheter drive moment-to-moment pulse adjustments. | הפטנט מתמקד במערכת אבלציה בפולסים עם היזון חוזר בזמן אמת המשלבת מודול הספק, HMI, מודול פולס במתח גבוה עם יחידת מיתוג, בתוספת גילוי אימפדנס וטמפרטורה במטרה. הוא מתמודד עם בעיות הבטיחות של טיפול RF מסורתי בניתוח מבוך IV לפרפור פרוזדורים (טווח לא נשלט, סיכון תרומבוס). אותות לולאה סגורה מהקטטר מניעים התאמות פולסים מרגע לרגע. | 1 |  |  |  | 1 | [6 empty] | 1 | [18 empty] | 1 |  | ZHULING SHANGHAI MEDICAL DEVICES CO LTD | ZHULING SHANGHAI MEDICAL DEVICES CO LTD | ZHULING SHANGHAI MEDICAL DEVICES CO LTD | ZHULING SHANGHAI MEDICAL DEVICES CO LTD | ZHULING SHANGHAI MEDICAL DEVICES CO LTD | 2024-11-07 | 2024 | 2025-01-10 | 2025 | 2024-11-07 | 2024 | Application | Utility Invention | active | CN202411579855 | 318570788 | [NA] | CN
375 | CN120651489A | Medical mid-infrared 1720nm pulse fiber laser | The invention discloses a medical mid-infrared 1720 nm pulse fiber laser, which relates to the technical field of laser medicine, and comprises the following steps: collecting data of the pulse fiber laser, respectively calculating a thulium-doped ion concentration uniformity index, laser energy intensity and ablation indexes of small target tissue ablation areas, and calculating a thulium-doped ion concentration uniformity index and a thulium-doped ion concentration uniformity index, a thulium-doped ion concentration uniformity index and a thulium-doped ion concentration uniformity index; according to the method, the spectrum analysis technology, the multiple linear regression algorithm, the convolutional neural network algorithm and the modern information technology are closely combined, the multiple thulium-doped optical fiber sections are combined, and the anti-fake color coding graph of the target tissue ablation area is drawn. The thulium-doped ion concentration distribution uniformity degree and the target tissue ablation degree are comprehensively monitored in real time, monitored data become more accurate indexes under the same condition, and the intelligent degree of the medical intermediate infrared 1720n pulse fiber laser in the using process is remarkably improved. | 1.A medical mid-infrared 1720nm pulsed fiber laser, characterized in that it includes the following steps: collecting pulsed fiber laser data, the pulsed fiber laser data including thulium-doped fiber spectrum data, pulse fiber data, laser data and target tissue data; performing feature extraction on the thulium-doped fiber spectrum data, combining with the pulsed fiber data, obtaining thulium-doped fiber ion concentration distribution evaluation data, and performing normalization processing on it, thereby calculating the thulium-doped ion concentration uniformity index, and evaluating the distribution uniformity of the thulium-doped ion concentration in each small thulium-doped fiber segment; using the laser data to calculate the laser output wavelength accuracy, combining the pulsed fiber temperature data with the multivariate linear regression algorithm, constructing a laser energy intensity output model, and thereby obtaining the corresponding laser energy intensity; calculating the ablation index of each small target tissue ablation area through the target tissue data, analyzing the ablation degree of each small target tissue ablation area, and numbering the ablation degree of each small target tissue ablation area; Based on the uniformity of the distribution of thulium ion concentrations in each small thulium-doped fiber segment, a heat map matrix of each small thulium-doped fiber segment is constructed; based on the degree of ablation of each small target tissue ablation region, an anti-counterfeiting color-coded map of each target tissue ablation region is drawn; and a target tissue damage analysis model is constructed by combining the laser energy intensity, the ablation degree number of each small target tissue ablation region, and a convolutional neural network algorithm;According to the thermal map matrix of each small thulium-doped optical fiber segment and the anti-counterfeiting color coding map of each target tissue ablation area, the distribution uniformity of the thulium ion concentration in each small thulium-doped optical fiber segment and the ablation degree of each small target tissue ablation area are analyzed respectively. According to the number of each small thulium-doped optical fiber segment and the number of each small target tissue ablation area, the small thulium-doped optical fiber segment corresponding to the abnormal uniformity of the thulium ion concentration distribution and the target tissue ablation area corresponding to the abnormal ablation degree are located respectively, and corresponding measures are taken to improve the problems of abnormal uniformity of the thulium ion concentration distribution and abnormal ablation degree. 2.A medical mid-infrared 1720nm pulsed fiber laser according to claim 1, characterized in that: the data acquisition process of the pulsed fiber laser comprises: dividing the thulium-doped fiber into several segments, acquiring each small thulium-doped fiber segment, dividing the target tissue ablation area into several areas of equal area, acquiring each small target tissue ablation area; deploying different types of acquisition equipment to acquire thulium-doped fiber spectral data, pulsed fiber data, laser data and target tissue data, wherein the different types of acquisition equipment include broadband light sources, spectrum analyzers, laser pump sources, tunable lasers, optical power meters, distributed fiber temperature measurement systems, optical time domain reflectometers, wavelength meters, spot analyzers, high-speed photodetectors, oscilloscopes, photoacoustic imaging systems, backscattered light diagnostic systems, optical coherence tomography systems and fiber Bragg grating temperature monitoring systems; The thulium-doped fiber spectral data includes the absorption spectrum, fluorescence spectrum and gain spectrum of trivalent thulium ions in the thulium-doped fiber; the pulse fiber data includes the temperature and length of each section of the pulse fiber; the laser data includes the actual output laser wavelength of the laser, the spot diameter when the laser beam acts on the tissue surface and the pulse frequency of the laser pulse sequence; the target tissue data includes the photoacoustic signal intensity, backscattered light intensity, thermal damage depth and temperature of the ablation area of the target tissue when it is acted upon by the laser;The pulse fiber data, the laser data and the target tissue data are cleaned and normalized, the thulium-doped fiber spectrum data are baseline corrected, the wavelength calibrated and the spectrum smoothed, and timestamps are assigned to the various types of pre-processed pulse fiber laser data, and the assigned timestamps are adjusted to achieve the acquisition time synchronization of the thulium-doped fiber spectrum data, the pulse fiber data, the laser data and the target tissue data, and the thulium-doped fiber spectrum data is normalized. The spectral data, the pulsed fiber data, the laser data and the target tissue data are integrated to generate a pulsed fiber laser data set. 3.A medical mid-infrared 1720nm pulsed fiber laser according to claim 2, characterized in that: the process of obtaining thulium-doped fiber ion concentration distribution evaluation data by feature extraction and normalizing it comprises: the thulium-doped fiber ion concentration distribution evaluation data includes the absorption coefficient, gain coefficient and gain saturation threshold power of each small thulium-doped fiber segment, and the peak wavelength and half-maximum full width of the fluorescence spectrum corresponding to each small thulium-doped fiber segment; using a characteristic peak detection algorithm to identify the characteristic absorption peak wavelength of the trivalent thulium ion from the absorption spectrum of the trivalent thulium ion in the thulium-doped fiber, according to the rare earth ion spectrum theory, the energy level transition corresponding to the characteristic absorption peak of the trivalent thulium ion, the absorbance at the characteristic absorption peak wavelength of the trivalent thulium ion, and according to the Lambert-Beer law, the absorption coefficient of each small thulium-doped fiber segment is calculated in combination with the length of each segment of the pulse fiber; Using spectral analysis technology, the fluorescence intensity and wavelength are extracted from the fluorescence spectrum of trivalent thulium ions in the thulium-doped fiber. Combined with the derivative method, the peak position of the fluorescence spectrum is determined, and then the corresponding peak wavelength is determined. Based on the peak wavelength of the fluorescence spectrum, two fluorescence wavelengths corresponding to half the fluorescence intensity are extracted. The absolute value of the difference between the two extracted fluorescence wavelengths is used to calculate the half-maximum full width of the fluorescence spectrum corresponding to each small thulium-doped fiber segment;Combining the signal gain linear fitting method and the saturation curve nonlinear fitting method, the gain coefficient and gain saturation threshold power of each small thulium-doped fiber segment are extracted from the gain spectrum of trivalent thulium ions in the thulium-doped fiber. Z-score normalization is performed on the absorption coefficient, gain coefficient, and gain saturation threshold power of each small thulium-doped fiber segment extracted from the features, as well as the peak wavelength and half-maximum full width of the corresponding fluorescence spectrum of each small thulium-doped fiber segment. 4.A medical mid-infrared 1720nm pulsed fiber laser according to claim 3, characterized in that: the process of calculating the thulium-doped ion concentration uniformity index and evaluating the distribution uniformity of the thulium-doped ion concentration in each small thulium-doped fiber segment includes: using the weighted average dispersion method, combined with the thulium-doped fiber ion concentration distribution evaluation data after Z-score normalization processing, to calculate the thulium-doped ion concentration uniformity index; when the thulium-doped ion concentration uniformity index is lower than 0.25, the thulium-doped ion concentration in each small thulium-doped fiber segment corresponds to a low distribution uniformity; when the thulium-doped ion concentration uniformity index is between 0.25 and 0.45, the thulium-doped ion concentration in each small thulium-doped fiber segment corresponds to a medium distribution uniformity; when the thulium-doped ion concentration uniformity index is higher than 0.45, the thulium-doped ion concentration in each small thulium-doped fiber segment corresponds to a high distribution uniformity. 5The medical mid-infrared 1720nm pulsed fiber laser according to claim 4 is characterized in that: the process of calculating the laser output wavelength accuracy, combining a multivariate linear regression algorithm, constructing a laser energy intensity output model, and then obtaining the corresponding laser energy intensity comprises: setting the target output laser wavelength of the laser to 1720nm, calculating the laser output wavelength accuracy based on the absolute value of the difference between the actual output laser wavelength of the laser and the target output laser wavelength, performing Z-score normalization on the laser output wavelength accuracy and integrating it into the pulsed fiber laser dataset; extracting the temperature of each pulse fiber segment in the pulsed fiber laser dataset, the spot diameter of the laser beam when it acts on the tissue surface, the pulse frequency of the laser pulse sequence, and the laser output wavelength accuracy, and dividing the extracted data into a first training set and a first test set; The first training set data is used in combination with a multivariate linear regression algorithm. The first training set data is used as input and the laser energy intensity is used as output to learn the temperature of each section of the pulsed optical fiber, the spot diameter when the laser beam acts on the tissue surface, the pulse frequency of the laser pulse sequence, the relationship between the laser output wavelength accuracy and the laser energy intensity. The linear relationship between the laser energy intensity and the laser beam is used to train a laser energy intensity output model; the first test set data is input into the laser energy intensity output model, the regression coefficient and intercept term of the laser energy intensity output model are adjusted, the performance of the laser energy intensity output model is optimized, and the laser energy intensity output model is obtained; the corresponding laser energy intensity is output in combination with the current temperature of each section of the pulsed optical fiber, the spot diameter of the laser beam when acting on the tissue surface, the pulse frequency of the laser pulse sequence, and the output wavelength accuracy of the laser, and the laser energy intensity is integrated into the pulsed fiber laser data set. 6The medical mid-infrared 1720nm pulsed fiber laser according to claim 5 is characterized in that: the process of calculating the ablation index of each small target tissue ablation area, analyzing the ablation degree of each small target tissue ablation area, and numbering the ablation degree of each small target tissue ablation area includes: based on the principle of collaboratively evaluating the ablation degree of each small target tissue ablation area based on the target tissue data of each item, using the preprocessed target tissue data in combination with the nonlinear product model formula to calculate the ablation index of each small target tissue ablation area, and the calculation process is as follows: AQi = SPA + SBS + Dtd + TRE Wherein, AQj is the ablation index of each small target tissue ablation area, SpA, Sb$, Dtd and Tre are the photoacoustic signal intensity, backscattered light intensity, thermal damage depth and temperature of the ablation area of the target tissue when it is subjected to laser action after normalization, respectively, and i is the number of each small target tissue ablation area; when the ablation index of each small target tissue ablation area is lower than 0.35, it corresponds to a low ablation degree; when the ablation index of each small target tissue ablation area is between 0.35-0.7, it corresponds to a medium ablation degree; when the ablation index of each small target tissue ablation area is higher than 0.7, it corresponds to a high ablation degree; the ablation degree of each small target tissue ablation area is numbered, and the ablation degree and number of each small target tissue ablation area are integrated into the pulsed fiber laser data set. 7.A medical mid-infrared 1720nm pulsed fiber laser according to claim 6, characterized in that: the process of constructing the heat map matrix of each small cobalt-doped fiber segment comprises: numbering each small cobalt-doped fiber segment, determining the heat map matrix structure, setting the row dimension to the number of each small cobalt-doped fiber segment, arranging the segments in ascending order of the numbers of each small cobalt-doped fiber segment, and setting the column dimension to the distribution uniformity of the cobalt-doped ion concentration in each small cobalt-doped fiber segment; according to color mapping rules, using continuous color scales, from light to dark, to represent the high distribution uniformity, medium distribution uniformity, and low distribution uniformity of the cobalt-doped ion concentration in each small cobalt-doped fiber segment, thereby constructing the heat map matrix of each small cobalt-doped fiber segment. 8.A medical mid-infrared 1720nm pulsed fiber laser according to claim 7 is characterized in that: the process of drawing the anti-counterfeiting color coding map of each target tissue ablation area includes: using color coding technology to select green, yellow and red, using Pantone color cards to calibrate printing colors, and associating green, yellow and red with the low ablation degree, medium ablation degree and high ablation degree of each small target tissue ablation area respectively; combining numbering and labeling technology, according to the numbering of each small target tissue ablation area, labeling each small target tissue ablation area, and using vector graphics software to arrange each small target tissue ablation area in numbered order; adjusting the printing accuracy to above 300dpi, and then drawing the anti-counterfeiting color coding map of each target tissue ablation area. 9. According to claim 8, a medical mid-infrared 1720nm pulsed fiber laser, characterized in that: The process of constructing the target tissue damage analysis model includes: extracting the laser energy intensity and the ablation degree number of each small target tissue ablation area from the pulse fiber laser data set, and dividing the currently extracted data into a second training set and a second test set; combining the second training set data and the convolutional neural network algorithm, taking the laser energy intensity as input and the ablation degree number of each small target tissue ablation area as output, learning the nonlinear relationship between the laser energy intensity and the ablation degree number of each small target tissue ablation area, and training the target tissue damage analysis model; inputting the second test set data into the target tissue damage analysis model, comparing the output results of the target tissue damage analysis model with the actual ablation degree numbers of each small target tissue ablation area, evaluating the performance of the target tissue damage analysis model, adjusting the parameters of the target tissue damage analysis model, optimizing the target tissue damage analysis model, and then obtaining the final target tissue damage analysis model. 10.A medical mid-infrared 1720nm pulsed fiber laser according to claim 9, characterized in that: the process of evaluating the abnormal thulium-doped ion concentration distribution uniformity and the ablation degree corresponding to each target tissue ablation area, and taking corresponding measures to improve includes: analyzing the thermal map matrix of each small thulium-doped fiber segment, obtaining the distribution uniformity of the thulium-doped ion concentration in each small thulium-doped fiber segment according to the color depth of the continuous color scale, retrieving the corresponding small thulium-doped fiber segment number based on the row dimension, when the thulium-doped ion concentration in the small thulium-doped fiber segment is highly uniformly distributed, retaining the existing fiber manufacturing process of the corresponding small thulium-doped fiber segment, and using a pump source that matches the fiber absorption peak to maximize the energy coupling efficiency; when the thulium-doped ion concentration in the small thulium-doped fiber segment is medium uniformly distributed, The inner cladding gradient refractive index design is adopted to optimize the heat conduction efficiency of the clamping structure of the corresponding small thulium-doped fiber segment, and the pulse width modulation technology is used to shorten the high-power pulse duration of the corresponding small thulium-doped fiber segment; when the thulium ion concentration in the small thulium-doped fiber segment is of low distribution uniformity, the unqualified small thulium-doped fiber segment is eliminated, and the segmented pumping technology is adopted to set an independent pump source in the corresponding small thulium-doped fiber segment. The pump energy of each segment is dynamically adjusted according to the concentration distribution, and the chirped pulse amplification technology is introduced to broaden, amplify and compress the pulse sequence to reduce the peak power density;According to the color of each small target tissue ablation area in the anti-counterfeiting color-coded map of each target tissue ablation area, the ablation degree of each small target tissue ablation area is analyzed. According to the number label, the number corresponding to each small target tissue ablation area is retrieved. When the small target tissue ablation area is green, it corresponds to a low ablation degree, and the ablation plan of the corresponding small target tissue ablation area is maintained; when the small target tissue ablation area is yellow, it corresponds to a medium ablation degree, and intermittent irradiation is used to utilize the thermal relaxation time to reduce damage to the small target tissue ablation area with normal peripheral ablation degree. Single-point irradiation is changed to spiral scanning and fence scanning to evenly cover the corresponding small target tissue ablation area and reduce local excessive ablation; when the small target tissue ablation area is red, it corresponds to a high ablation degree, and physiological saline spray and low-temperature gas are used to cool the corresponding small target tissue ablation area to below 37°C to reduce the expansion of the carbonization range. Laser and other energy forms are combined to complement and adjust the ablation needs. | 2025 | ZIBO KECHUANG MEDICAL INSTR | NA | NA | NA | China | The patent focuses on a medical mid-IR 1720 nm pulsed fiber laser whose data analysis builds heat-map matrices and color-coded charts based on thulium ion concentration uniformity, laser intensity, and an ablation index. It addresses the limitations of bulky, low-efficiency, complex-to-maintain conventional mid-IR lasers in medical settings. The fiber laser plus heat-map analysis supports precise ablation across small target regions with portability and stability. | הפטנט מתמקד בלייזר סיב פולסי רפואי באינפרא-אדום אמצעי 1720 ננומטר שניתוח הנתונים שלו בונה מטריצות מפת חום ותרשימים בקוד צבעים על בסיס אחידות ריכוז יוני תוליום, עוצמת לייזר ומדד אבלציה. הוא מתמודד עם המגבלות של לייזרי אינפרא-אדום אמצעי קונבנציונליים מסורבלים, בעלי יעילות נמוכה ומורכבים לתחזוקה בהגדרות רפואיות. לייזר הסיב בתוספת ניתוח מפת חום תומך באבלציה מדויקת על פני אזורי מטרה קטנים עם ניידות ויציבות. |  | 1 |  | 1 | [9 empty] | 1 | [8 empty] | 1 | [7 empty] | 1 |  | ZIBO KECHUANG MEDICAL INSTR CO LTD | ZIBO KECHUANG MEDICAL INSTR CO LTD | ZIBO KECHUANG MEDICAL INSTR CO LTD | ZIBO KECHUANG MEDICAL INSTR CO LTD | ZIBO KECHUANG MEDICAL INSTR CO LTD | 2025-05-30 | 2025 | 2025-09-16 | 2025 | 2025-05-30 | 2025 | Application | Utility Invention | inactive | CN202510722598 | 323419959 | CN120651489B | CN
376 | JP7587623B2 | SPECTROSCOPY SYSTEM FOR IDENTIFYING LIGHT SOURCE | To provide medical devices, systems and methods that can identify the type of light being used for spectral analysis of a target tissue.SOLUTION: A method of identifying an illumination light source in a surgical system comprises steps of: receiving a signal from a target after the target is illuminated by the illumination light source; performing spectroscopic analysis of the received signal; and, based at least in part on the spectroscopic analysis, determining a characteristic of the illumination light source. A method of treating a target comprises: receiving a signal from the target after the target is illuminated by the illumination light source; performing spectroscopic analysis of the received signal; based at least in part on the spectroscopic analysis, determining a first characteristic of the illumination light source and a second characteristic of the target; and, based at least in part on the determined characteristics of the illumination light source and the target, operating the surgical system to treat the target.SELECTED DRAWING: Figure 1 | 1.Surgical SystemsInThere is a targetShiningreceiving a signal from the target after it is illuminated by a bright light source;The target may include one or both of a test target or an anatomical target.and performing a spectroscopic analysis of the received signal.meansand determining a characteristic of the illumination source based at least in part on the spectroscopic analysis.and the characteristics of said target.Determinemeansandmeans for adjusting at least one parameter or characteristic of the illumination source or a laser light or treatment energy treating the target based on the determined characteristics of the illumination source and the determined characteristics of the target;EquippedSurgical Systems。 2.Comparing the spectroscopic analysis to database information containing spectroscopic information for different light types and/or different materials of the target.means2. The method of claim 1 , further comprising:Surgical Systems。 3.The spectroscopic analysis compares at least one of the intensities or spectra of the received signal to intensities or spectra associated with various types of illumination sources from the target located in the database information.thingThe method according to claim 2,Surgical Systems。 4.3. The method of claim 2, wherein the target comprises a test target configured to reflect illumination light without light absorption.Surgical Systems。 5.3. The method of claim 2, wherein the illumination source comprises at least one of a xenon light source and an LED light source.Surgical Systems。 6.Retrieving database information from an internet-connected servermeansThe method of claim 2 further comprises:Surgical Systems。 7.The database information of claim 6 comprises an artificial intelligence engine.Surgical Systems。 8.Adjusting at least one parameter or characteristic of the illumination source to improve spectroscopic analysis of tissue based on the identified illumination source.meansThe method of claim 1 further comprising:Surgical Systems。 9.9. The method of claim 8, wherein the parameters include at least one of brightness, power, wavelength, and intensity, and the characteristics include a type of the illumination source.Surgical Systems。 10. IlluminationIlluminating anatomy with bright lightmeansand performing a spectroscopic analysis of the reflection of the illumination light.meansand determining a characteristic of the anatomical tissue using the spectroscopic analysis.meansThe method of claim 1 further comprising:Surgical Systems。 11.11. The method of claim 10, wherein the characteristics include one or more of the type, material, composition, composition profile, structure, and/or hardness of the anatomical tissue.Surgical Systems。 12.Adjusting at least one parameter or characteristic of a treatment light source to improve treatment of the anatomical tissue based on the anatomical tissue type.meansThe method of claim 10 further comprising:Surgical Systems。 13. Operation of the Surgical System1. A method comprising:Tareceiving a signal from the target after the target is illuminated by an illumination source; performing a spectroscopic analysis of the received signal; determining a first characteristic of the illumination source and a second characteristic of the target based at least in part on the spectroscopic analysis; and determining a first characteristic of the illumination source and a second characteristic of the target based at least in part on the determined characteristics of the illumination source and the target.the illumination source orTreating said targetAdjust at least one of the parameters or characteristics of the treatment light source.and 14.The aboveAdjust at least one of the parameters or characteristics14. The method of claim 13, wherein generating or adjusting one or more parameters of the surgical system includes an operating mode, power or energy, a pulse shape profile, an emitted pulse regime, and/or a combined output pulse train. 15.The method of claim 14 , wherein adjusting the one or more parameters is performed automatically by a controller. 16.The method of claim 14 , wherein adjusting the one or more parameters comprises prompting a user to accept proposed adjustments to the illumination source. 17.The method of claim 14 , wherein at least one of the spectroscopic analysis, the parameters of the illumination source, or the characteristics of the illumination source are communicated over the Internet. 18.The method of claim 17 , further comprising utilizing an artificial intelligence engine to generate the one or more parameters. 19.14. The method of claim 13, wherein receiving the signal from the target after the target is illuminated by the illumination source comprises receiving a first signal from the test target after the test target is illuminated by the illumination source, and receiving a second signal from the anatomical target after the anatomical target is illuminated by the illumination source. 20.20. The method of claim 19, wherein performing a spectroscopic analysis of the received signal comprises performing a spectroscopic analysis of the first signal to identify a type of light emitted by the illumination source, and performing a spectroscopic analysis of the second signal to identify a tissue characteristic of the anatomical target. 21. twenty one.21. The method of claim 20, wherein the characteristics include one or more of the type, material, composition, composition profile, structure, and/or hardness of the tissue. 22. twenty two.21. The method of claim 20, wherein the test target comprises a reflective surface and the illumination source comprises a xenon light source or an LED light source. 23. 1. A surgical system comprising:means for receiving a signal from the target after the target is illuminated by an illumination source;means for performing a spectroscopic analysis of the received signal;means for determining a characteristic of the illumination source and a characteristic of the target based at least in part on the spectroscopic analysis;means for adjusting at least one parameter or characteristic of the illumination source based on the determined characteristics of the illumination source and the determined characteristics of the target to improve spectroscopic analysis of tissue;Equipped withthe parameters include at least one of brightness, power, wavelength, and intensity, and the characteristics include a type of the illumination source;Surgical systems. 24. 1. A surgical system comprising:means for illuminating the anatomical structure with an illumination light;means for performing a spectroscopic analysis of the reflection of the illumination light from the anatomical tissue;means for determining characteristics of the illumination light and characteristics of the anatomical tissue using said spectroscopic analysis;means for adjusting at least one parameter or characteristic of a treatment light source based on the determined characteristics of the illumination light and the characteristics of the anatomical tissue to improve treatment of the anatomical tissue;A surgical system comprising: | 2022 | Olympus¦GYRUS | Medical and Healthcare | Large-Size Company | http://gyrusacmi.com/ | USA | The patent focuses on a spectroscopic system for identifying the type of light source used for spectral analysis of target tissue during a surgical procedure. It addresses the inability of generic spectroscopy systems to know which light source is producing the signal, which can confuse analysis. Identifying the light source enables proper interpretation of spectroscopic data during surgery. | הפטנט מתמקד במערכת ספקטרוסקופית לזיהוי סוג מקור האור המשמש לניתוח ספקטרלי של רקמת מטרה במהלך פרוצדורה ניתוחית. הוא מתמודד עם חוסר היכולת של מערכות ספקטרוסקופיה גנריות לדעת איזה מקור אור מייצר את האות, מה שיכול לבלבל את הניתוח. זיהוי מקור האור מאפשר פרשנות נכונה של נתונים ספקטרוסקופיים במהלך הניתוח. |  | 1 | [11 empty] | 1 | [8 empty] | 1 | [7 empty] | 1 |  | [NA] | [NA] | [NA] | [NA] | GYRUS ACMI INC DBA OLYMPUS SURGICAL TECH AMERICA | 2023-03-27 | 2023 | 2024-11-20 | 2024 | 2022-03-25 | 2022 | Grant | Utility Invention | active | JP20230050339 | 291096112 | CN116804575A¦DE102023106462A1¦JP2023143888A¦JP2025013557A¦JP7778891B2¦US20230301520A1 | JP
377 | CN121445470A | Cardiac Pulsed Electric Field Ablation Pen and Catheter | This invention relates to the field of smart medical technology, specifically to a cardiac pulsed electric field ablation pen and catheter, comprising: a main body, a lead wire, electrode terminals, a server, and a monitoring module; one end of the main body is connected to one end of the lead wire, and the other end of the lead wire is connected to the electrode terminals; several electrodes are alternately arranged on the electrode terminals; a pulse generator is installed inside the main body and connected to the electrodes via the lead wire for emitting pulsed electric fields; the server is used to acquire lesion information and, based on the lesion information, set and optimize the working parameters of the electrode terminals; the monitoring module is used to perform pulsed electric field ablation on the lesion using the electrode terminals and monitor muscle contraction; the server is also used to, based on muscle contraction, trigger the electrode terminals to reduce voltage and then perform pulsed electric field ablation again if muscle contraction occurs, while the monitoring module monitors muscle contraction; otherwise, pulsed electric field ablation continues. This solution can perform targeted and effective ablation, reduce damage to blood vessels, and ensure treatment effectiveness. | 1 A cardiac pulsed electric field ablation pen, characterized by comprising: a main body, a lead wire, electrode terminals, a server, and a monitoring module; one end of the main body is connected to one end of the lead wire, and the other end of the lead wire is connected to the electrode terminals; several electrodes are alternately arranged on the electrode terminals; a pulse generator is installed inside the main body and connected to the electrodes via the lead wire for emitting pulsed electric fields; the server is used to acquire lesion information and, based on the lesion information, set and optimize the working parameters of the electrode terminals; the monitoring module is used to monitor muscle contraction during pulsed electric field ablation of the lesion by the electrode terminals; the server is also used to determine the next stage of treatment based on muscle contraction: if muscle contraction occurs, the voltage of the electrode terminals is reduced, and pulsed electric field ablation is performed again, while the monitoring module monitors the muscle contraction; otherwise, pulsed electric field ablation continues. 2 According to claim 1, the cardiac pulse electric field ablation pen is characterized in that: the lesion information includes: the location, diameter, and depth of the lesion; the working parameters of the electrode ends include: electrode spacing and voltage; the step of setting and optimizing the working parameters of the electrode ends according to the lesion information includes: setting the corresponding electrode spacing d according to the diameter of the lesion, and adjusting the electrode spacing d according to the location of the lesion; setting the electric field strength E according to the effective ablation threshold of the electric field strength; calculating the voltage according to the electric field strength and the electrode spacing; obtaining the distribution of nerves or blood vessels within a preset distance from the lesion according to the location of the lesion; and adjusting the voltage according to the distribution of nerves or blood vessels and the depth of the lesion. 3 According to claim 2, the cardiac pulse electric field ablation pen is characterized in that: setting the corresponding electrode spacing d according to the diameter of the lesion and adjusting the electrode spacing d according to the location of the lesion includes: if the diameter of the lesion belongs to a first preset diameter range, then setting the electrode distance to the first preset electrode distance; if the diameter of the lesion belongs to a second preset diameter range, then setting the electrode distance according to the mapping relationship between the second preset diameter range and the first preset electrode distance range; if the diameter of the lesion belongs to a third preset diameter range, then setting the electrode distance according to the mapping relationship between the third preset diameter range and the second preset electrode distance range; if the diameter of the lesion belongs to a fourth preset diameter range, then setting the click distance to the second preset electrode distance; if the lesion is located in a thin-walled area, then reducing the electrode distance by the first preset distance; if the lesion is located in a thick-walled area, then increasing the electrode distance by the second preset distance. 4 According to claim 2, the cardiac pulse electric field ablation pen is characterized in that: adjusting the voltage according to the distribution of nerves or blood vessels and the depth of the lesion includes: if there are nerves or blood vessels within a preset distance from the lesion, the voltage is reduced by a first preset proportion; if the depth of the lesion is greater than the preset depth, the voltage is increased by a second preset proportion, thereby compensating for the attenuation of the electric field. 5 According to claim 1, the cardiac pulse electric field ablation pen is characterized in that: the monitoring module is further used to perform impedance monitoring and obtain the impedance change during ablation; the server is further used to analyze whether the impedance drop ratio exceeds a preset ratio based on the impedance change, and if so, determine that it is a valid electroporation; analyze whether the impedance rises more than a preset rise amount per unit time, and if so, generate tissue carbonization prompt information and push it. 6 According to claim 1, the cardiac pulsed electric field ablation pen is characterized in that: the monitoring module is further used to monitor the temperature and obtain the temperature during ablation; the server is further used to determine whether the temperature is greater than a preset temperature threshold, and if so, to trigger the main body to stop the pulsed electric field ablation. melt. 7 According to claim 2, the cardiac pulse electric field ablation pen is characterized in that: the server is further configured to automatically generate a pulse parameter scheme, and use an optimization model to optimize the pulse parameter scheme to obtain the optimal pulse parameter scheme, and set the pulse parameters of the pulse generator according to the optimal pulse parameter scheme, including: S1. Within the corresponding preset range, randomly generate k pulse parameter schemes as individual schemes and combine them into an initial population; wherein the pulse parameters in the pulse parameter schemes include: pulse width type, inter-pulse delay Δin frequency /, and number of pulses N; S2. Construct the evaluation function P; S3. Based on the constructed evaluation function value, calculate the evaluation function value of each individual scheme in the initial population, and sort them from smallest to largest according to the evaluation function value to generate the sorting result; S4. Based on the sorting results, obtain the top r individual schemes, perform iterative optimization, and obtain the individual scheme with the smallest evaluation function value as the optimal impulse parameter scheme. 8 According to claim 7, the cardiac pulse electric field ablation pen is characterized in that: S2 includes: constructing an evaluation function P; wherein P is a neurological risk evaluation function, S is a surgical time evaluation function, S is a pulse total energy evaluation function, and g is an ablation efficiency evaluation function; g1, g2, and g3 correspond to the weights of the neurological risk evaluation function, surgical time evaluation function, pulse total energy evaluation function, and ablation efficiency evaluation function, respectively; wherein the neurological risk evaluation function P is: P = 1 / G + cxp(-2 * (E / /J-0.8)))x(1/(1 + cxp(-0.5 *(/-3))»; wherein P4 = 300x(1 + 2000/Ψ), 300 is the minimum pacing threshold of the myocardium, in units of V/cm, 2000 is the time value of the myocardium, in units of 1%, and Ψ is the stimulation threshold of the electric field; exp is the natural exponential function; wherein the surgical time evaluation function P is: P = -^· 860' where 60 is the conversion factor from seconds to minutes; The pulse total energy evaluation function is: The ablation efficiency evaluation function is: = * £2 + * 0M; Where Gs is the electric field coverage evaluation function, and W is the standardized ablation depth evaluation function; % and w2 are the weight values of the electric field coverage evaluation function and the standardized ablation depth evaluation function, respectively; Where Gi: Where is the predicted ablation radius in mm; is the equivalent radius of the lesion in mm; where /: Among them, such as v's background: Lesion information also includes: lesion volume, where Jiuru represents the lesion volume; Ojelbow: Just likeJhope Where is the maximum effective depth threshold, "2 = 8mm, is the predicted ablation depth in mm; C is the calibration coefficient. 9 According to claim 7, the cardiac pulsed electric field ablation pen is characterized in that: step S4 comprises: S401. Initialize the number of iterations i = l, and the number of invalid optimizations j = 0; S402. Based on the sorting results in the initial population, obtain the first r individual schemes; r S403. Cross and mutate the obtained individual schemes to form several new individual schemes; S404. Based on the constructed evaluation function value, calculate the evaluation function value of the new individual scheme; S405. Based on the evaluation function values, sort the population and new individual schemes from smallest to largest, and generate the sorting results. S406. Obtain the top K individual schemes as the new initial population; S407. Calculate the discriminant degree of individual schemes in the initial population and the new initial population, and determine whether the discriminant degree is greater than the preset discriminant degree. If yes, then execute S409; if no, then j=j+1, and execute S408. S408. Determine if j is greater than the preset threshold for invalid optimization attempts. If yes, execute S4010; otherwise, execute S409. S409. Determine if i is greater than the preset iteration number threshold. If yes, execute S4010; otherwise, execute S402. S4010. Obtain the first individual scheme as the optimal pulse parameter scheme. 10 A cardiac pulsed electric field ablation catheter, characterized in that: it uses a cardiac pulsed electric field ablation pen as described in any one of claims 1-9, and further includes: a catheter; and the length of the lead wire is greater than the length of the catheter. | 2025 | NA | NA | NA | NA | China | The patent focuses on a cardiac PFA ablation pen and catheter with body, wire, electrode end (with alternating electrodes), service end, and a built-in pulse generator and monitoring module. It addresses the difficulty of managing rapid disordered atrial contractions in atrial fibrillation, especially in elderly or comorbid patients, with cumbersome equipment. The pen-style format aims to make cardiac PFA more ergonomic for the operator. | הפטנט מתמקד בעט קטטר אבלציה לבבי PFA עם גוף, חוט, קצה אלקטרודה (עם אלקטרודות מתחלפות), קצה שירות ומחולל פולסים ומודול ניטור מובנים. הוא מתמודד עם הקושי בניהול התכווצויות פרוזדוריות מהירות לא מסודרות בפרפור פרוזדורים, במיוחד אצל מטופלים מבוגרים או עם תחלואה נלווית, עם ציוד מסורבל. פורמט עט שואף להפוך PFA לבבי לארגונומי יותר עבור המפעיל. | 1 | [10 empty] | 1 | [12 empty] | 1 | [7 empty] | [NA] | [NA] | [NA] | [NA] | [NA] | 2025-09-19 | 2025 | 2026-02-03 | 2026 | 2025-09-19 | 2025 | Application | Utility Invention | active | CN202511345579 | 326276193 | [NA] | CN
378 | CN121744407A | A method and system for visualizing the thermal field of microwave ablation based on non-invasive CT thermometry and multiphysics simulation | This invention provides a method and system for visualizing the thermal field of microwave ablation based on non-invasive CT thermography and multiphysics simulation, comprising the following steps: Step 1: Acquiring CT image data during the microwave ablation process; Step 2: Establishing a temperature mapping model based on the relationship between the HU value of the CT image and temperature; Step 3: Visualizing the temperature field in three dimensions; Step 4: Constructing a multiphysics simulation model based on the frequency domain Maxwell's equations and Pennes' equations; Step 5: Predicting the extent of tissue damage using the Arrhenius model; Step 6: Performing bidirectional verification between the CT thermography results and the simulation results. This technical solution can significantly improve the accuracy and visualization effect of microwave ablation boundary determination, and has important clinical application value. | 1 A method for visualizing the thermal field of microwave ablation based on non-invasive CT thermography and multiphysics simulation is characterized by the following steps: Step 1: Acquire CT image data during the microwave ablation process; Step 2: Establish a temperature mapping model based on the relationship between the HU value of the CT image and temperature; Step 3: Perform three-dimensional visualization of the temperature field; Step 4: Construct a multiphysics simulation model based on the frequency domain Maxwell's equations and Pennes' equations; Step 5: Predict the tissue damage range using the Arrhenius model; Step 6: Perform bidirectional verification between the CT thermography results and the simulation results. 2 According to claim 1, a method for visualizing the thermal field of microwave ablation based on non-invasive CT thermography and multiphysics simulation is characterized in that step 1 specifically includes: acquiring CT image data of the target tissue during microwave ablation; using isolated pig liver at 4 °C under refrigeration as the target tissue; heating the isolated pig liver to 37 °C in a water bath before the experiment to meet the initial temperature requirements of the experiment; inserting the microwave ablation needle used in the surgery into the isolated pig liver; setting the surgical ablation power and surgical ablation time; after the isolated pig liver is ablated, placing it in a CT scanner for scanning; simultaneously using the control group isolated pig liver, performing the same ablation scheme on it; after the ablation is completed, cutting along the plane where the ablation needle is located, and using a thermal imager to obtain the temperature value and thermal field distribution of the plane. 3 According to claim 1, a method for visualizing the thermal field of microwave ablation based on non-invasive CT thermography and multiphysics simulation is characterized in that step 2 specifically includes: after obtaining the data through the experiment in step 1, delineating the ROI and creating a new DICOM file in the oblique section state of the CT image; then performing secondary sampling on the CT image to obtain the distribution matrix information of the CT value of the section, which is used for data fitting; establishing a database with [T,HU] pairing information by registering the matrix of CT values with the temperature field matrix of the thermal imager; and using linear fitting, nonlinear fitting, and deep learning algorithms for fitting, comparing the fitting function with its fitting parameters, and selecting the best mathematical model. 4 According to claim 1, a microwave ablation thermal field visualization method based on non-invasive CT thermography and multiphysics simulation is characterized in that step 3 specifically includes: visualizing the temperature distribution results in three-dimensional software and constructing an interactive three-dimensional thermal field visualization system; establishing a mapping framework in a 3D slicer based on the mathematical model established in step 2, and drawing thermal field distribution maps of cross-section, sagittal plane, and coronal plane; generalizing and evolving the thermal field in units of voxels to construct a three-dimensional thermal field model, that is, the user can simultaneously observe the thermal field file and CT image information in the 3D slicer; and judging the thermal field evolution inside the biological tissue at this moment. 5 The microwave ablation thermal field visualization method based on non-invasive CT thermometry imaging and multiphysics simulation according to claim 1 is characterized in that, the construction of the multiphysics model based on the frequency domain Maxwell's equations and Pennes' biological heat transfer equations in step 4 specifically includes: 1) Input conditions: The model input includes the structural parameters of the microwave ablation needle, surgical operation parameters, histological parameters, and boundary conditions; 2) Model Structure: First, the frequency domain Maxwell's equations are solved in COMSOL software to obtain the spatial distribution of the electromagnetic field within the tissue. The dissipated power density generated by the electromagnetic field is used as a heat source term and coupled to the Pennes biological heat transfer equation to calculate the tissue temperature distribution over time. Furthermore, the Arrhenius damage integral model is used to transform the temperature-time integral result into a probability distribution of irreversible tissue damage. The propagation of electromagnetic waves in liver tissue is considered using a transverse magnetic field, described by equation (1). Vx (4J) where j represents the relative permittivity of liver tissue, b represents the conductivity of liver tissue, represents the angular frequency of microwaves, represents the vacuum permittivity, represents the relative permeability, represents the free space wavenumber, j represents the imaginary unit, and represents the angular component of the magnetic field strength; The interaction between electromagnetic waves and liver tissue is defined by = SAR, where SAR represents the specific absorptivity, specifically the electromagnetic radiation energy absorbed per unit mass of liver tissue, P represents the density of liver tissue, and E represents the electric field strength vector within the tissue; When a coaxial slot antenna emits electromagnetic waves, the electromagnetic waves pass through the liver tissue and then propagate throughout the region; the electromagnetic waves are absorbed in the liver tissue and converted into heat generation; The heat Qext generated in the liver tissue by converting SAR is defined as: The heat transfer analysis within liver tissue is described by the Pennes biothermal equation, which is written as: C *5 ΊΈΜ and also-+4, (3) g represents the specific heat capacity of liver tissue, T represents the temperature of liver tissue, t represents time, k represents the thermal conductivity of liver tissue, 0 represents the density of blood, Cp represents the specific heat capacity of blood, then represents the blood perfusion rate, 4 represents the arterial blood temperature, represents the metabolic heat production of liver tissue, and represents the external heat source term generated by microwave energy precipitation; thermal damage is the thermal damage of tissue, and the damage model proposed by Henriques & Moritz, namely the Arrhenius damage equation, is adopted for this purpose; thermal damage in tissue is defined as the denaturation of proteins in tissue, as shown in equation (4): Q = {exp(Spring Divination)4) n represents the degree of thermal damage to liver tissue, A represents the frequency factor, knot represents the activation energy, R represents the universal gas constant; The proportion of necrotic tissue is expressed as: % = 1 - exp(- Ω) Mathematical model of coupled electromagnetic wave propagation and heat transfer in biological tissue; The COMSOL multiphysics finite element method is used to solve the equation set; The calculation scheme first calculates the external heat source by solving the electromagnetic wave propagation equation, and then solves the time-varying temperature of each point under the biological heat equation; 3) Output results: The model output is a three-dimensional temperature field distribution matrix and the corresponding tissue damage probability field, which can automatically generate the predicted ablation range under a specified isothermal boundary or a specific damage threshold. 4) Prediction and application: By inputting the patient's individualized parameters before surgery, the model simulates the diffusion trend of thermal and damage fields during surgery; after comparison with in vitro experiments and clinical validation, it serves as a reference for preoperative pathway planning and risk assessment. 6 According to claim 1, a method for visualizing the thermal field of microwave ablation based on non-invasive CT thermography and multiphysics simulation is characterized in that step 5 specifically includes: processing surgical parameters and experimental objects within the COMSOL software. Consistent modeling was performed to simulate the distribution of thermal and tissue damage fields before surgery. 7 According to claim 1, a method for visualizing the thermal field of microwave ablation based on non-invasive CT thermography and multiphysics simulation is characterized in that step 6 specifically includes: comparing the accuracy of CT thermography and COMSOL simulation results again through in vitro pig liver experiments; further iterating the design of ablation parameters and the dynamic updating of tissue biological parameters in experimental and clinical practice; further reducing the prediction range error of COMSOL; and further reducing the temperature error of CT thermography. 8 A microwave ablation thermal field visualization system based on non-invasive CT thermography and multiphysics simulation includes a processor, a memory, and a bus. The memory stores machine-readable instructions executed by the processor. The system is characterized in that, when running, the processor and the memory communicate via the bus, and the machine-readable instructions, when executed by the processor, constitute a microwave ablation thermal field visualization method based on non-invasive CT thermography and multiphysics simulation as described in any one of claims 1 to 7. | 2025 | NA | NA | NA | NA | China | The patent focuses on a microwave-ablation thermal-field visualization method using non-invasive CT thermometry, where CT HU values map to temperature, and validates with multi-physics simulation. It addresses the inability of operators to directly visualize the microwave thermal field, which limits precise control during solid-tumor MWA. CT becomes a stand-in temperature sensor for microwave ablation. | הפטנט מתמקד בשיטת הדמיית שדה תרמי באבלציית מיקרוגלים באמצעות תרמומטריית CT לא פולשנית, שבה ערכי HU של CT ממופים לטמפרטורה, ומאומתים בסימולציה רב-פיזיקלית. הוא מתמודד עם חוסר היכולת של מפעילים להמחיש ישירות את שדה המיקרוגלים התרמי, המגבילה בקרה מדויקת במהלך MWA של גידולים מוצקים. CT הופך לחיישן טמפרטורה חלופי לאבלציית מיקרוגלים. | [8 empty] | 1 |  |  |  | 1 | [8 empty] | 1 | [6 empty] | 1 |  |  |  | [NA] | [NA] | [NA] | [NA] | [NA] | 2025-12-17 | 2025 | 2026-03-27 | 2026 | 2025-12-17 | 2025 | Application | Utility Invention | active | CN202511906824 | 327059217 | [NA] | CN
379 | CN119732734A | A cardiac radiofrequency ablation steady-state monitoring method and system | The present invention discloses a cardiac radiofrequency ablation steady-state monitoring method and system, which relates to the field of ablation monitoring technology. The method includes: obtaining an ablation catheter, a sensor array interactively connected to the ablation catheter, extracting monitoring data of the sensor array in a preset monitoring window, obtaining a temperature-impedance monitoring data sequence and a current monitoring data sequence; obtaining a temperature-impedance monitoring data feature set; obtaining a temperature-impedance monitoring data deviation coefficient; obtaining a current monitoring stability coefficient; according to the size of the temperature-impedance monitoring data deviation coefficient and the current monitoring stability coefficient, feedback updating the preset monitoring window, obtaining an updated monitoring window, and performing cardiac radiofrequency ablation steady-state monitoring based on the updated monitoring window. The present invention solves the technical problems in the prior art that the cardiac radiofrequency ablation steady-state monitoring has a low fit with the actual situation and a delayed response, and achieves the technical effect of improving the reliability of steady-state monitoring. | 1. L A method for steady-state monitoring of cardiac radiofrequency ablation, characterized in that the method includes: obtaining an ablation catheter, interacting with a sensor array connected to the ablation catheter, extracting monitoring data of the sensor array within a preset monitoring window, and obtaining a temperature-impedance monitoring data sequence and a current monitoring data sequence, wherein the ablation catheter is used for cardiac radiofrequency ablation; performing multiple random data generation based on a receptive field random generator to obtain a receptive field set, and constructing a data feature analyzer set based on the receptive field set, using the data feature analyzer set to extract data from the temperature-impedance monitoring data sequence respectively to obtain a temperature-impedance monitoring data feature set; using preset temperature-impedance monitoring data standard features, performing feature deviation identification on the temperature-impedance monitoring data feature set to obtain a temperature-impedance monitoring data deviation coefficient; traversing the current monitoring data sequence to perform current monitoring stability identification to obtain a current monitoring stability coefficient; feedback updating the preset monitoring window according to the size of the temperature-impedance monitoring data deviation coefficient and the current monitoring stability coefficient to obtain an updated monitoring window, and performing steady-state monitoring of cardiac radiofrequency ablation based on the updated monitoring window. 2.A cardiac radiofrequency ablation steady-state monitoring method as claimed in claim 1, characterized in that it comprises: constructing a receptive field random generator, wherein the receptive field random generator comprises a receptive field random generation function, and the receptive field random generation function is: S = where S is the receptive field, y~/v (. Drama, e obeys the standard normal distribution, now = [Zhilei Ping signed a contract, "(,) is the Gamma function, 0 formula 0, 2). 3.A cardiac radiofrequency ablation steady-state monitoring method as described in claim 1 is characterized in that multiple random data generation is performed based on a receptive field random generator to obtain a receptive field set, and a data feature analyzer set is constructed based on the receptive field set, and the data feature analyzer set is used to extract data from the temperature-impedance monitoring data sequence to obtain a temperature-impedance monitoring data feature set, including: multiple random data generation is performed based on the receptive field random generator to obtain an initial receptive field set; the initial receptive fields are combined in pairs, and similarity recognition is performed on the combination results to obtain a similarity recognition result set; the initial receptive field set is distinguished and authenticated according to the similarity recognition result set, and if the authentication is passed, the initial receptive field set is used as the receptive field set. 4.A cardiac radiofrequency ablation steady-state monitoring method as described in claim 3 is characterized by comprising: obtaining Q sample temperature-impedance monitoring data sequences and corresponding Q sample temperature-impedance monitoring data features, wherein Q is an integer greater than or equal to 1; randomly extracting a first receptive field from the receptive field set, and performing data segmentation on the Q sample temperature-impedance monitoring data sequences based on the size of the first receptive field to obtain a set of Q sample temperature-impedance monitoring data subsequences; determining the number of nodes in the input layer based on the size of the first receptive field and the amount of index data of the temperature-impedance monitoring data, and configuring the first input layer according to the number of nodes in the input layer; using the Q sample temperature-impedance monitoring data subsequence sets and the Q sample temperature-impedance monitoring data features, training a convolutional neural network comprising a first input layer, a fully connected layer, and an output layer, and in the training, performing segmentation on the network according to the output The network parameters are updated until the training converges to obtain a first data feature analyzer that has been trained; and a data feature analyzer is constructed in combination with the receptive field set, the Q sample temperature-impedance monitoring data sequences, and the Q sample temperature-impedance monitoring data features to obtain the data feature analyzer set. 5.A cardiac radiofrequency ablation steady-state monitoring method as described in claim 1 is characterized in that the preset temperature-impedance monitoring data standard features are used to identify the feature deviation of the temperature-impedance monitoring data feature set to obtain the temperature-impedance monitoring data deviation coefficient, including: traversing and identifying the degree of feature deviation between the temperature-impedance monitoring data feature set and the preset temperature-impedance monitoring data standard features to obtain a temperature-impedance monitoring data feature deviation set; performing mean calculation on the temperature-impedance monitoring data feature deviation set to determine the mean of the temperature-impedance monitoring data feature deviation, and taking the mean of the temperature-impedance monitoring data feature deviation as the deviation center; iterating the deviation center based on a preset neighborhood radius until a preset iteration stop condition is met, and taking the temperature-impedance monitoring data feature deviation corresponding to the deviation center updated in the last iteration as the temperature-impedance monitoring data feature deviation coefficient. 6.A cardiac radiofrequency ablation steady-state monitoring method as described in claim 5 is characterized by comprising: constructing a neighborhood for the deviation center according to a preset neighborhood radius to obtain a deviation from the mean neighborhood; updating and iterating the deviation center based on the preset neighborhood radius to obtain an iterative deviation center, and constructing a deviation iterative neighborhood; judging whether the difference in the number of neighborhood deviations between the deviation iterative neighborhood and the deviation from the mean neighborhood is less than or equal to a preset difference, if not, updating the deviation iterative neighborhood to the deviation center, and continuing to update and iterate until a preset iteration stop condition is met to obtain the characteristic deviation coefficient of the temperature-impedance monitoring data, wherein the preset iteration stop condition is that the number of iterations is greater than or equal to the preset number of iterations and/or the difference in the number of neighborhood deviations between two adjacent iterations is less than or equal to the preset difference. 7.A cardiac radiofrequency ablation steady-state monitoring method as described in claim 1, characterized in that it includes: performing weighted calculation on the temperature-impedance monitoring data deviation coefficient and the current monitoring stability coefficient to determine the window feedback adjustment coefficient; performing feedback update on the preset monitoring window based on the window feedback adjustment coefficient to obtain the updated monitoring window. 8.A cardiac radiofrequency ablation steady-state monitoring system, characterized in that the system is used to implement a cardiac radiofrequency ablation steady-state monitoring method as described in any one of claims 1 to 7, and the system comprises: a data sequence acquisition module, used to obtain an ablation catheter, a sensor array interactively connected to the ablation catheter, extract monitoring data of the sensor array within a preset monitoring window, and obtain a temperature-impedance monitoring data sequence and a current monitoring data sequence, wherein the ablation catheter is used for cardiac radiofrequency ablation; a monitoring data feature set acquisition module, used to perform multiple data random generation based on a receptive field random generator to obtain a receptive field set, and to construct a data feature analyzer set based on the receptive field set, and to use the data feature analyzer set to extract data from the temperature-impedance monitoring data sequence respectively to obtain a temperature-impedance monitoring data feature set; a data deviation coefficient acquisition module, used to use a preset temperature-impedance monitoring data standard feature to perform feature deviation identification on the temperature-impedance monitoring data feature set to obtain a temperature-impedance monitoring data deviation coefficient; a current monitoring stability coefficient acquisition module, used to traverse the current monitoring data sequence to perform current monitoring stability Identify and obtain the current monitoring stability coefficient; a steady-state monitoring module is used to feedback update the preset monitoring window according to the temperature-impedance monitoring data deviation coefficient and the current monitoring stability coefficient, obtain an updated monitoring window, and perform cardiac radiofrequency ablation steady-state monitoring based on the updated monitoring window. | 2024 | NA | NA | NA | NA | China | The patent focuses on a cardiac RF ablation steady-state monitor that gathers temperature, impedance, and current data from a sensor array along the catheter and analyzes their time series. It addresses the difficulty of recognizing when an RF ablation has reached a stable, effective state during atrial fibrillation or atrial premature beat treatment. The system detects when ablation has reached a stable, effective state. | הפטנט מתמקד במנטר מצב יציב של אבלציית RF לבבית האוסף נתוני טמפרטורה, אימפדנס וזרם ממערך חיישנים לאורך הקטטר ומנתח את סדרות הזמן שלהם. הוא מתמודד עם הקושי בזיהוי מתי אבלציית RF הגיעה למצב יציב ויעיל במהלך טיפול בפרפור פרוזדורים או דופק פרוזדורי מוקדם. המערכת מזהה מתי האבלציה הגיעה למצב יציב ויעיל. | 1 | [9 empty] | 1 | [11 empty] | 1 |  | 1 | [7 empty] | [NA] | [NA] | [NA] | [NA] | [NA] | 2024-12-26 | 2024 | 2025-04-01 | 2025 | 2024-12-26 | 2024 | Application | Utility Invention | active | CN202411938157 | 320457506 | [NA] | CN
380 | WO2024081158A3 | METHOD AND APPARATUS FOR TREATMENT OF PULMONARY CONDITIONS | Apparatus and methods for deactivating bronchial nerves and ablating smooth muscle extending along a bronchial branch of a mammalian subject to treat asthma and related conditions. An ultrasonic transducer (11) is inserted into the bronchus as, for example, by advancing the distal end of a catheter (10) bearing the transducer into the bronchial section to be treated. The ultrasonic transducer emits ultrasound so as to heat tissues throughout a relatively large impact volume (13) to a temperature sufficient to inactivate nerve conduction and ablate smooth muscle but insufficient to cause rapid ablation or necrosis of the surrounding tissues. The treatment can be performed without locating or focusing on individual bronchial nerves or smooth muscle. | 1. IN THE CLAIMS: 1. System for simultaneously inactivating bronchial nerve conduction and ablating smooth muscle in a mammalian subject, comprising: an elongated member supporting an ultrasound transducer adapted for insertion into a bronchial branch of a bronchial tree of the mammalian subject and for transmitting ultrasound energy; and an actuator or control unit electrically connected to the transducer, the actuator or control unit adapted to control the ultrasound transducer to i) transmit ultrasound energy at a sub-therapeutic level in a diagnostic mode; and ii) transmit ultrasound energy in an impact volume encompassing the bronchial branch in a therapeutic mode so that the ultrasound energy is applied at a therapeutic level sufficient to inactivate conduction of bronchial nerves and ablate smooth muscle throughout the impact volume in a therapeutic mode. 2. The system of claim 1, wherein the actuator or control unit transmits ultrasound energy in the diagnostic mode to generate a signal and integrate ultrasound echoes to measure bronchial diameter. 3. The system of claim 1, wherein the actuator or control unit is configured to analyze the impedance measurements to determine locations of bronchial cartilage rings and the actuator or control unit is additionally configured to activate the ultrasound transducer to transmit ultrasound therapeutic waveform energy between adjacent cartilaginous bronchial tissue. 4. The system of claim 1 , wherein the actuator or control unit is configured to transmit ultrasound energy in the diagnostic mode to generate a signal and analyze ultrasound echoes to ensure circumferential coupling. 5. The system of claim 1 , wherein the actuator or control unit is configured to transmit ultrasound energy in the diagnostic mode to generate a signal and analyze ultrasound echoes to ensure inter-cartilage positioning. The system of claim 1 , wherein analysis of the return signal of the transmitted pulse in the diagnostic mode determines one or both of a type or state of tissue. The system of claim 1, wherein analysis of the return signal of the transmitted in the diagnostic mode provides centering of the transducer in the bronchial branch. The system of claim 1 , wherein the actuator or control unit is configured to transmit ultrasound energy at the therapeutic level interleaved with transmitting in the diagnostic mode. The system of claim 1, wherein the actuator or control unit is configured to transmit pulsed signals at the therapeutic level. The system of claim 1, wherein the actuator or control unit further transmits ultrasound energy at the sub-therapeutic level after transmitting ultrasound energy at the therapeutic level has ablated peri bronchial tissues. The system of claim 1, further comprising a balloon, the balloon containing cooling fluid. The system of claim 11, wherein pulsating fluid within the balloon enables volumetric A mode diameter measurements by analyzing amplitude and signal width fluctuations. The system of claim 11, wherein the cooling fluid is pulsated to provide pulsating flow to measure bronchial compliance and degree of smooth muscle ablation. The system of claim 11, wherein the transducer emits ultrasonic energy at a sub-therapeutic level to detect heating of the fluid and/or efficacy of cooling. The system of claim 1, further comprising at least one electrode for making impedance measurements along a wall of the bronchial branch to locate cartilaginous tissue and spaces or gaps between cartilaginous tissue. The system of claim 1, wherein the transducer transmits ultrasound energy in the therapeutic mode to direct the ultrasound energy between cartilage rings. System for conducting pulmonary treatment in a mammalian subject, comprising: an elongated member having a balloon and an energy transducer within the balloon, the transducer adapted for insertion into a bronchial branch of a bronchial tree of the mammalian subject and for transmitting energy into surrounding tissue, and an actuator or control unit electrically connected to the transducer, the actuator or control unit being adapted to control the transducer to emit ultrasonic energy into an impact volume encompassing the bronchial branch so that the energy is applied at a therapeutic level sufficient to inactivate conduction of bronchial nerves throughout the impact volume; wherein the balloon contains cooling fluid and the cooling fluid is pulsated to measure bronchial compliance. The system of claim 17, wherein the transducer emits ultrasound energy at a sub-therapeutic level in a diagnostic mode to measure bronchial diameter to optimize dosing. The system of claim 17, wherein the cooling fluid is pulsated to generate an A-mode amplitude modulation to enable echo identification for bronchial diameter measurements. The system of claim 19, wherein the pulsating fluid within the balloon enables volumetric A mode diameter measurements by analyzing amplitude and signal width fluctuations. The system of claim 20, wherein the transducer emits ultrasonic energy at a sub-therapeutic level to detect heating of the fluid and/or efficacy of cooling. The system of claim 17, wherein the transducer emits ultrasonic energy at a sub-therapeutic level to detects heating of the fluid or efficacy of cooling. System for simultaneously inactivating bronchial nerve conduction and ablating smooth muscle in a mammalian subject, comprising: an ultrasound transducer adapted for insertion into a bronchial branch of a bronchial tree of the mammalian subject and for transmitting circumferential ultrasound energy; and an actuator or control unit electrically connected to the transducer, the actuator being adapted to control the ultrasound transducer to transmit ultrasound energy into an impact volume encompassing the bronchial branch so that the ultrasound energy is applied at a therapeutic level sufficient to inactivate conduction of bronchial nerves and ablate smooth muscle throughout the impact volume, the actuator or control unit being configured to energize the ultrasound transducer to emit a short pulse at a sub-therapeutic level, the actuator or control unit being further configured to process a volume integrated A- mode signal, which represents an accumulated intensity of the circumferential ultrasound echoes, to enable a precise dosing of ultrasound energy in the impact volume. The system of claim 23, wherein the bronchial branch is a 3rd or 4th generation bronchial branch and the actuator or control unit is adapted to control the ultrasound transducer depending on volumetric A mode bronchial diameter measurement to transmit ultrasound energy at an acoustic power level of approximately 10 to approximately 30 watts for approximately 10 to approximately 30 seconds to provide an absorbed dose of approximately 100 to approximately 900 joules in the impact volume depending on bronchial diameter to optimize simultaneous ablation of both smooth muscle and bronchial nerves in the impact volume surrounding the 3rd or 4th generation bronchial branch. The system of claim 24, further comprising a balloon filled with a coupling fluid, said ultrasound transducer being disposed inside said balloon, said balloon being expandable into contact with a wall of the bronchial branch, wherein the actuator or control unit is adapted to control the temperature of the coupling fluid together with dose parameters of acoustic power level and duration so as to maintain a temperature of the bronchial wall below 40°C while achieving a temperature above 60°C throughout the impact volume including the smooth muscle and or bronchial nerves surrounding 3rd or 4th generation bronchial branch. The system of claim 23, wherein the actuator or control unit is adapted to control the ultrasound transducer depending on volumetric A mode bronchial diameter measurement to transmit ultrasound energy at an acoustic power level of approximately 10 to approximately 30 watts for approximately 10 to approximately 30 seconds to provide an absorbed dose of approximately 100 to approximately 900 joules in the impact volume, the acoustic power level and duration of energy application depending on bronchial diameter to optimize the separate ablation of both smooth muscle and bronchial nerves in impact volumes surrounding bronchial branches of larger diameters. The system of claim 23, wherein the actuator or control unit is adapted to control the ultrasound transducer to transmit the ultrasound energy in a pulsed function interleaved with volumetric A mode diagnostic acquisitions to enable a quasi-simultaneous therapeutic/diagnostic mode. The system of claim 23, further comprising a catheter with a distal end and a proximal end, the transducer being mounted to the catheter adjacent the distal end inside a compliant balloon, filled with circulating fluid in order to cool the ultrasound transducer, the ultrasound transducer being disposed in alignment with and centered on an axis of the bronchial branch. The system of claim 23, further comprising a catheter with a distal end and a proximal end, the transducer being mounted to the catheter adjacent the distal end inside a compliant balloon filled with circulating, pulsating fluid in order to enable volumetric A mode diameter measurements by analyzing amplitude and signal width fluctuations. The system of claim 23, further comprising a catheter with a distal end and a proximal end, the transducer being mounted to the catheter adjacent the distal end inside a compliant balloon filled with circulating, pulsating fluid, the actuator or control unit being configured for measuring bronchial compliance and therewith ASM ablation by analyzing resistance to balloon pressure pulsations. The system of claim 30, wherein the catheter is configured to hold the ultrasound transducer out of contact with the wall of the bronchial branch. The system of claim 30, wherein the ultrasound transducer has an axis, the catheter being constructed and arranged to hold the ultrasound transducer so that the axis of the ultrasound transducer is generally parallel to a longitudinal axis of the bronchial branch, the ultrasound transducer being adapted to transmit the ultrasound energy in a 360° cylindrical pattern surrounding the axis of the ultrasound transducer. The system of claim 23, further comprising a catheter including a compliant balloon surrounding the ultrasound transducer and arranged to hold the ultrasound transducer substantially centered in the bronchial branch, the compliant balloon including a porous membrane engageable with the wall of the bronchial branch and filled with a coupling fluid for improved acoustic coupling, the fluid being exudable through the porous membrane enabling sliding of the balloon axially while in engagement with the wall of the bronchial branch to achieve inter cartilage positioning. The system of claim 33, wherein the compliant balloon is provided with electrodes, said actuator or control unit being operatively connected to said electrodes for making impedance measurements along a wall of the bronchial branch to enable a locating of cartilaginous tissue and spaces or gaps between cartilaginous tissue. The system of claim 34, wherein the actuator or control unit is configured to analyze the impedance measurements to determine locations of bronchial cartilage rings and the actuator or control unit being additionally configured to activate the ultrasound transducer to transmit ultrasound therapeutic waveform energy between adjacent cartilaginous bronchial tissue. The system of claim 23, wherein the ultrasound transducer is further adapted to receive volumetric A mode ultrasound signals representing or encoding bronchial geometry, the actuator or control unit being further adapted to: control the ultrasound transducer to transmit measurement ultrasound energy at a level below therapeutic levels, receive volumetric A mode echo signals from the ultrasound transducer representing reflected measurement ultrasonic energy; analyze the received volumetric A mode echo signals; and determine a size of the bronchial branch based on the received echo volumetric A mode signals. The system of claim 23, further comprising an axially movable circular loop electrode operatively connected to said actuator or control unit, said actuator or control unit being adapted to receive electrical impedance measurements representing bronchial cartilage geometry from said circular loop electrode. The system of claim 23, further comprising a compliant balloon surrounding the ultrasound transducer and provided with multiple circular electrodes on an outer surface, said actuator or control unit being operatively connected to said circular electrodes to receive therefrom electrical impedance measurements representing bronchial cartilage geometry. The system of claim 38, wherein the actuator or control unit is adapted to control the ultrasound transducer to vary the acoustic power used to transmit the ultrasound energy depending on the determined size of the bronchial branch. A method for pulmonary treatment, comprising the steps of: inserting an ultrasound transducer into a bronchial section of the mammalian subject; and actuating the transducer to transmit therapeutically effective ultrasound energy into an impact volume of at least approximately 1 cm3, encompassing the bronchial section so that the therapeutically effective ultrasound energy inactivates all bronchial nerves and smooth muscle in the impact volume. The method of claim 40, wherein the ultrasound energy is transmitted at an acoustic power level of approximately 10 to approximately 30 watts for approximately 10 to approximately 30 seconds to provide an absorbed dose of approximately 100 to approximately 900 joules throughout the impact volume, further comprising conducting a volumetric A mode ultrasound signal scan, analyzing resulting volumetric A mode echo signals, and selecting intensity and duration parameters of the ultrasound energy in accordance with a geometry of the bronchial section and intended ablation of nerves and or smooth muscle. The method of claim 41, wherein the ultrasound transducer is disposed inside a compliant balloon attached proximate a distal end of a catheter and filled with coupling fluid, further comprising controlling a temperature of the coupling fluid in the balloon to prevent or reduce thermal damage to a wall of the bronchial section while enabling therapeutically effective increase in temperature of the bronchial nerves and or smooth muscle in the impact volume. The method of claim 40, wherein the steps of inserting the ultrasound transducer and actuating the transducer to transmit ultrasound energy are performed without determining actual locations of the bronchial nerves and the smooth muscle. The method of claim 40, wherein the bronchial section is a third-generation branch of a bronchial tree of the mammalian subject, a major source of resistance to airflow, the step of actuating the ultrasound transducer being performed so that a single application of the ultrasound energy at a treatment site in the third-generation branch is effective to inactivate nerve conduction distal of the treatment site and to ablate all smooth muscle throughout the impact volume. The method of claim 44, further comprising the steps of: repositioning the ultrasound transducer distal into a next or fourth generation bronchial branch, another major source of airflow resistance, after the step of actuating the transducer; and then repeating the step of actuating the transducer. The method of claim 40, wherein the step of actuating the transducer is performed so that the ultrasound energy is transmitted in a pulsed function. The method of claim 46, wherein the ultrasound transducer is operated in a pulsed mode performing quasi-simultaneous volumetric A mode measurements and treatment. . The method of claim 40, wherein the step of inserting the ultrasound transducer into the bronchial section is performed through the working channel of a bronchoscope under visual guidance. . The method of claim 40, wherein the step of inserting the ultrasound transducer is performed through a steerable sheath. . The method of claim 40, wherein the step of inserting the ultrasound transducer is performed without a sheath or bronchoscope and with a steerable ultrasound catheter. 1. The method of claim 40, wherein the step of inserting the ultrasound treatment catheter is performed over a guide wire which has been placed through a working channel of a bronchoscope. 2. System for conducting pulmonary treatment in a mammalian subject, comprising: an energy transducer adapted for insertion into a bronchial branch of a bronchial tree of the mammalian subject and for transmitting energy into surrounding tissue; an actuator or control unit electrically connected to the transducer, the actuator being adapted to control the transducer to emit energy into an impact volume encompassing the bronchial branch so that the energy is applied at a therapeutic level sufficient to inactivate conduction of bronchial nerves throughout the impact volume; and at least one electrode, the actuator or control unit operatively connected to the at least one electrode for making impedance measurements along a wall of the bronchial branch to enable locating of cartilaginous tissue and spaces or gaps between cartilaginous tissue. 3. The system of claim 52, wherein the at least one electrode comprises a plurality of electrodes and the actuator or control unit is configured to direct RF current through selected ones of said plurality of electrodes or portions thereof positioned between cartilage rings based on the impedance measurements. 54. The system of claim 53, wherein the electrodes are supported by a compliant balloon. 55. The system of claim 55, wherein the transducer is positioned within the balloon. 56. The system of claim 54, wherein the balloon contains cooling fluid. | 2022 | AERWAVE MEDICAL | Medical and Healthcare | Start Up | https://aerwavemed.com/ | USA | The patent focuses on apparatus and methods for treating asthma by deactivating bronchial nerves and ablating bronchial smooth muscle using an ultrasound transducer inserted through a catheter. It addresses the unmet need for a focused, non-pharmacologic asthma treatment that hits the underlying nerves and smooth muscle. Focused ultrasound heats a relatively large volume of bronchial wall tissue at once. | הפטנט מתמקד במכשיר ושיטות לטיפול באסתמה על ידי הפסקת פעילות עצבי הסימפונות ואבלציה של שריר חלק של הסימפונות באמצעות מתמר אולטרסאונד המוחדר דרך קטטר. הוא מתמודד עם הצורך הלא-מסופק בטיפול אסתמה ממוקד ולא תרופתי הפוגע בעצבים ובשריר החלק העומדים מאחורי המחלה. אולטרסאונד ממוקד מחמם נפח גדול יחסית של רקמת קיר הסימפונות בבת אחת. |  |  | 1 | [12 empty] | 1 |  |  |  | 1 | [6 empty] | 1 | [5 empty] | AERWAVE MEDICAL INC | AERWAVE MEDICAL INC | AERWAVE MEDICAL INC | AERWAVE MEDICAL INC | AERWAVE MEDICAL INC | 2023-10-06 | 2023 | 2024-05-23 | 2024 | 2022-10-14 | 2022 | Application | Utility Invention | inactive | WO2023US34606 | 304315509 | EP4601549A2¦WO2024081158A2 | WO
381 | CN119836310A | System and method for subclavian periarterial denervation | A method of treating a heart disease in a subject includes percutaneously introducing a catheter into the vasculature of the subject. The catheter may include a neuromodulation element. A neuromodulation element may be positioned on the expansion element. A neuromodulation element of the catheter may be positioned in a subclavian artery of a subject. The catheter may be used to electrically stimulate at least one of a dorsal subclavian loop or a ventral subclavian loop. The method may include confirming stimulation to a dorsal subclavian loop and/or a ventral subclavian loop by monitoring cardiac parameters. After the stimulation to the dorsal subclavian loop and/or the ventral subclavian loop is confirmed, ablation energy can be provided to the dorsal subclavian loop and/or the ventral subclavian loop. | 1.A method for treating a cardiac disease in a subject, the method comprising: percutaneously introducing a catheter into the vasculature of the subject, wherein the catheter includes a neuromodulatory element; positioning the neuromodulatory element in the left subclavian artery of the subject; electrically stimulating the left dorsal subclavian loop; confirming the stimulation of the left dorsal subclavian loop by monitoring a first cardiac parameter; providing ablative energy to the left dorsal subclavian loop after confirming the stimulation of the left dorsal subclavian loop; electrically stimulating the left ventral subclavian loop; confirming the stimulation of the left ventral subclavian loop by monitoring a second cardiac parameter; providing ablative energy to the left ventral subclavian loop after confirming the stimulation of the left ventral subclavian loop; positioning the neuromodulatory element in the right subclavian artery of the subject; electrically stimulating the right dorsal subclavian loop; confirming the stimulation of the right dorsal subclavian loop by monitoring a third cardiac parameter; After confirming stimulation of the right dorsal subclavian loop, providing ablative energy to the right dorsal subclavian loop; electrically stimulating the right ventral subclavian loop; confirming stimulation of the right ventral subclavian loop by monitoring a fourth cardiac parameter; and after confirming stimulation of the right ventral subclavian loop, providing ablative energy to the right ventral subclavian loop. 2.The method of claim 1, wherein the first cardiac parameter is different from at least one of the second cardiac parameter, the third cardiac parameter, or the fourth cardiac parameter. 3.The method of claim 2, wherein the second cardiac parameter is different from at least one of the first cardiac parameter, the third cardiac parameter, or the fourth cardiac parameter. 4.The method of claim 3, wherein the third cardiac parameter is different from at least one of the first cardiac parameter, the second cardiac parameter, or the fourth cardiac parameter. 5.The method of claim 4, wherein the fourth cardiac parameter is different from at least one of the first cardiac parameter, the second cardiac parameter, or the third cardiac parameter. 6.The method according to claim 1, wherein the first cardiac parameter is the same as at least one of the second cardiac parameter, the third cardiac parameter or the fourth cardiac parameter. 7.The method according to claim 1, wherein the second cardiac parameter is the same as at least one of the first cardiac parameter, the third cardiac parameter or the fourth cardiac parameter. 8.The method of claim 1, wherein the third cardiac parameter is the same as at least one of the first cardiac parameter, the second cardiac parameter, or the fourth cardiac parameter. 9.The method of claim 1, wherein the fourth cardiac parameter is the same as at least one of the first cardiac parameter, the second cardiac parameter, or the third cardiac parameter. 10.The method according to any one of claims 1 to 9, wherein at least one of the first cardiac parameter, the second cardiac parameter, the third cardiac parameter or the fourth cardiac parameter includes at least one of arterial micro-blood pressure changes, whether AF is induced, whether arrhythmia is induced, whether cardiac cycle length changes are induced, repetitive recovery curves, right atrial ERP, left atrial ERP, dERP, heart rate, IACT, or a multipolar catheter unipolar electrogram of short DI under or after steady-state RV pacing. 11. The method according to any one of claims 1 to 9 further comprises re-stimulating the subclavian loop after providing the ablation energy, and further providing ablation energy to the subclavian loop if the cardiac parameters confirm the stimulation. 12.A method according to claim 11, comprising repeating said restimulation and providing further ablation energy until said cardiac parameters no longer confirm stimulation. 13. The method according to any one of claims 1 to 9, wherein the catheter includes an expansion element and the neuromodulation element is positioned on the expansion element. 14. The method of claim 13, comprising expanding the expansion element to dilate the vasculature. 15.A method for treating heart disease in a subject, the method comprising: positioning a neuromodulatory element in the subclavian artery of the subject; electrically stimulating at least one of the dorsal subclavian loop or the ventral subclavian loop; confirming the stimulation of the dorsal subclavian loop and/or the ventral subclavian loop by monitoring cardiac parameters; and providing ablation energy to the dorsal subclavian loop and/or the ventral subclavian loop after confirming the stimulation of the dorsal subclavian loop and/or the ventral subclavian loop. 16.The method of claim 15, wherein the ablation energy comprises radiofrequency ablation energy. 17.The method according to claim 15, wherein the ablation energy includes cryogenic ablation energy. 18. The method of claim 15, wherein the electrical stimulation comprises using a combination of electrodes of the neuromodulatory element, and wherein providing ablation energy comprises using the same combination of electrodes of the neuromodulatory element. 19.The method of claim 15, further comprising assessing whether sympathetic nerve drive drives the refractory arrhythmia. 20. A method according to any one of claims 15 to 19, wherein the heart disease comprises atrial fibrillation. 21. twenty one. The method according to any one of claims 15 to 19, wherein the heart disease comprises a refractory arrhythmia or a ventricular arrhythmia. 22. twenty two.A catheter for treating heart disease in a subject, the catheter comprising: a slender element configured to be positioned in at least one of a left subclavian artery or a right subclavian artery, wherein a distal portion of the slender element comprises: an expansion element, a neuromodulatory element positioned on the expansion element and configured to stimulate at least one of a dorsal subclavian loop or a ventral subclavian loop and configured to ablate at least one of the dorsal subclavian loop or the ventral subclavian loop, the neuromodulatory element of the catheter being in the subclavian artery of the subject, the neuromodulatory element comprising an array of electrodes spaced circumferentially about the circumference of the catheter and longitudinally along the length of the catheter. 23. twenty three.A catheter according to claim 22, wherein the neuromodulatory element comprises a plurality of electrodes connected to a plurality of struts. 24. twenty four.A catheter according to claim 22, wherein the neuromodulatory element comprises a plurality of electrodes connected to a ring. 25.The catheter according to claim 22 includes a distal protection device located distal to the neuromodulatory element. 26. The catheter of any one of claims 23 to 25, wherein at least one of the plurality of electrodes comprises a triangular base comprising a tip configured to press into a blood vessel wall. 27. A catheter for treating a heart disease in a subject, the catheter comprising: an elongated element configured to be positioned in at least one of a left subclavian artery or a right subclavian artery, wherein a distal portion of the elongated element comprises: a neuromodulatory element configured to stimulate at least one of a dorsal subclavian haptic or a ventral subclavian haptic and configured to ablate at least one of the dorsal subclavian haptic or the ventral subclavian haptic, the neuromodulatory element comprising: An array of electrodes is spaced circumferentially about the circumference of the catheter and longitudinally along the length of the catheter. 28.A catheter according to claim 27, wherein the electrode array comprises a plurality of electrodes connected to a plurality of struts. 29.A catheter according to claim 27, wherein the electrode array comprises a plurality of electrodes connected to a ring. 30.A catheter according to claim 27, wherein the electrode array comprises a plurality of electrodes connected to a helix. 31.A catheter according to any one of claims 27 to 30, wherein at least one electrode of the electrode array comprises a triangular base comprising a tip configured to press into a blood vessel wall. 32.A catheter according to any one of claims 27 to 30, further comprising an expansion element on which the electrode array is located. 33.A catheter according to claim 32, wherein the expansion element comprises a balloon. 34.A method for treating heart disease in a subject, the method comprising: percutaneously introducing a catheter into the vascular system of the subject, wherein the catheter comprises a neuromodulatory element; positioning the neuromodulatory element in the subclavian artery of the subject; electrically stimulating at least one of the dorsal subclavian loop or the ventral subclavian loop; confirming the stimulation of the dorsal subclavian loop and/or the ventral subclavian loop by monitoring cardiac parameters; and after confirming the stimulation of the dorsal subclavian loop and/or the ventral subclavian loop, providing ablative energy to the dorsal subclavian loop and/or the ventral subclavian loop. 35.The method of claim 34, further comprising re-stimulating the subclavian loop after providing the ablative energy, and further providing ablative energy to the subclavian loop if the cardiac parameters confirm the stimulation. 36.A method according to claim 35, comprising repeating said restimulation and providing further ablation energy until said cardiac parameters no longer confirm stimulation. 37. The method of any one of claims 34 to 36, wherein percutaneously introducing the catheter into the vasculature comprises inserting the catheter into the femoral artery of the subject. 38. The method of any one of claims 34 to 36, wherein percutaneously introducing the catheter into the vasculature comprises inserting the catheter into a radial artery of the subject. 39. The method of any one of claims 34 to 36, wherein percutaneously introducing the catheter into the vasculature comprises inserting the catheter into the subject's carotid artery. 40.The method of any one of claims 34 to 36, wherein percutaneously introducing the catheter into the vasculature comprises inserting the catheter into a femoral vein of the subject, and at least one of positioning the neuromodulatory element in the left subclavian artery or positioning the neuromodulatory element in the right left subclavian artery comprises traversing from a venous vasculature to an arterial vasculature. 41.A method for treating heart disease in a subject, the method comprising: percutaneously introducing a catheter into the vascular system of the subject, wherein the catheter includes a neuromodulatory element; positioning the neuromodulatory element in the subclavian artery of the subject; electrically stimulating the dorsal subclavian loop; confirming the stimulation of the dorsal subclavian loop by monitoring a first cardiac parameter; after confirming the stimulation of the dorsal subclavian loop, providing ablative energy to the dorsal subclavian loop; electrically stimulating the ventral subclavian loop; confirming the stimulation of the ventral subclavian loop by monitoring a second cardiac parameter; and after confirming the stimulation of the ventral subclavian loop, providing ablative energy to the ventral subclavian loop. 42.The method of claim 41, wherein the first cardiac parameter is different from the second cardiac parameter. 43.The method of claim 41, wherein the first cardiac parameter is the same as the second cardiac parameter. 44.The method according to any one of claims 41 to 43, wherein at least one of the first cardiac parameter or the second cardiac parameter includes at least one of arterial micro-blood pressure changes, whether AF is induced, whether arrhythmia is induced, whether cardiac cycle length changes are induced, repetitive recovery curves, right atrial ERP, left atrial ERP, dERP, heart rate, IACT, or a multipolar catheter unipolar electrogram of short DI under or after steady-state RV pacing. 45. A method according to any one of claims 41 to 43, further comprising re-stimulating the subclavian haptic after providing the ablation energy, and further providing ablation energy to the subclavian haptic if the cardiac parameters confirm the stimulation. 46. A method according to claim 45, comprising repeating said restimulation and providing further ablation energy until said cardiac parameters no longer confirm stimulation. 47. A method according to any one of claims 41 to 43, wherein the catheter includes an expansion element, the neuromodulation element is positioned thereon, and the expansion element is optionally expanded an amount to dilate the vasculature. 48.A catheter for treating heart disease in a subject, the catheter comprising: a slender element configured to be positioned in at least one of a left subclavian artery or a right subclavian artery, wherein a distal portion of the slender element comprises: a first neuromodulatory element configured to stimulate at least one of a dorsal subclavian haptic or a ventral subclavian haptic; a second neuromodulatory element configured to ablate at least one of the dorsal subclavian haptic or the ventral subclavian haptic; and a distal protection device distal to the first neuromodulatory element and the second neuromodulatory element. 49.A catheter according to claim 48, wherein the first neuromodulatory element comprises a first plurality of electrodes directly or indirectly coupled to one or more struts, rings and/or spirals. 50.A catheter according to claim 49, wherein at least one of the plurality of electrodes comprises a triangular base comprising a tip configured to press into a blood vessel wall, wherein, optionally, the at least one electrode further comprises a plurality of electrode protrusions located on different sides of the tip of the triangular base. 51.A catheter according to any one of claims 48 to 50, wherein the second neuromodulatory element comprises the first plurality of electrodes and a second plurality of electrodes directly or indirectly coupled to a second plurality of struts, rings and/or spirals or other elements of similar functionality. 52.The catheter of claim 48, further comprising an expansion element, and wherein the first neuromodulation element and the second neuromodulation element are positioned on the expansion element, wherein the expansion element optionally comprises a balloon. 53.A catheter for treating heart disease in a subject, the catheter comprising: a slender element configured to be positioned in at least one of a left subclavian artery or a right subclavian artery, wherein a distal portion of the slender element comprises: a neuromodulatory element configured to stimulate at least one of a dorsal subclavian loop or a ventral subclavian loop, and configured to ablate at least one of the dorsal subclavian loop or the ventral subclavian loop; and a distal protection device at the distal end of the neuromodulatory element. 54.A catheter according to claim 53, wherein the neuromodulatory element includes a plurality of electrodes directly or indirectly connected to a plurality of struts, rings and/or spirals or other elements of similar functionality. 55.A catheter according to any one of claims 53 to 54, wherein at least one of the plurality of electrodes comprises a triangular base comprising a tip configured to press into a blood vessel wall, and optionally comprising a plurality of electrode protrusions located on different sides of the tip of the triangular base. 56.The catheter of claim 55, further comprising an expansion element, and wherein the neuromodulation element is positioned on the expansion element, wherein the expansion element optionally comprises a balloon. 57.A method for treating a subject, the method comprising: stimulating specific electrodes of a neuromodulatory device; providing the results to an interface computer, the computer being configured to: perform real-time data analysis; and superimpose the response to the stimulation on an anatomical image, including marking the location where the response is greatest; and performing ablation at said location. 58.The method according to claim 57 further comprises one or more of the following: after ablation, re-stimulation to verify the effect of the ablation; if the ablation is ineffective, further ablation, and optionally repeating the verification and further ablation until the ablation is effective. 59.A method for treating a subject experiencing a refractory arrhythmia, the method comprising: stabilizing the subject; assessing the cause of the refractory arrhythmia; assessing whether sympathetic nerve drive is driving the refractory arrhythmia; coordinating an assessment of an arrhythmogenic lesion if sympathetic nerve drive is not driving the refractory arrhythmia; and performing subclavian loop ablation if sympathetic nerve drive is driving the refractory arrhythmia. 60.The method of claim 59, wherein assessing whether sympathetic nerve drive is driving the refractory arrhythmia comprises infusing epidural anesthesia and/or wherein assessing whether sympathetic nerve drive is driving the refractory arrhythmia comprises shutting down an implantable cardioverter-defibrillator of the subject. 61.A method for treating heart disease in a subject, the method comprising: positioning a neuromodulatory element in the subclavian artery of the subject; and providing ablative energy to the dorsal subclavian loop and/or the ventral subclavian loop. 62.The method of claim 61, wherein providing the ablation energy comprises selectively providing the ablation energy. 63.A method according to claim 62, wherein selectively providing the ablation energy includes one or more of: non-therapeutic stimulation to determine the location of the dorsal subclavian loop and/or the ventral subclavian loop; imaging the subclavian artery to determine the possible location of the dorsal subclavian loop and/or the ventral subclavian loop. 64.The method of claim 63, wherein providing the ablation energy comprises providing the ablation energy around the subclavian artery. 65A method of treating a subject, the method comprising: positioning a neuromodulatory element of a catheter in a subclavian artery of the subject, the neuromodulatory element comprising an array of electrodes spaced circumferentially about the circumference of the catheter and longitudinally along the length of the catheter; stimulating specific electrodes in the electrode array; providing the results to an interface computer configured to: perform real-time data analysis; and superimpose the response to the stimulation on an anatomical image, including marking the location where the response is greatest; and perform ablation at said location. 66.The method according to claim 65 further includes, after ablation, restimulation to verify the effectiveness of the ablation, and if the ablation is ineffective, further ablation, and optionally includes repeating restimulation and further ablation until it is verified that the ablation is effective. 67.A method according to claim 65 or 66, wherein the neuromodulatory element is positioned on an expansion element. 68. The method of claim 67, comprising expanding the expansion element to dilate the subclavian artery. 69.The method of claim 68, comprising performing impedance measurements when the expansion element is expanded. 70.A method for treating heart disease in a subject, the method comprising: percutaneously introducing a catheter into the vascular system of the subject, wherein the catheter comprises a neuromodulatory element; positioning the neuromodulatory element in the subclavian artery of the subject; electrically stimulating at least one of a dorsal subclavian loop or a ventral subclavian loop; confirming the stimulation of the dorsal subclavian loop and/or the ventral subclavian loop by monitoring biomarkers, wherein the biomarkers comprise at least one of an inotropic biomarker, a chronotropic biomarker, a conduction biomarker, a comfort biomarker, or an inflammatory biomarker; after confirming the stimulation of the dorsal subclavian loop and/or the ventral subclavian loop, providing ablative energy to the dorsal subclavian loop and/or the ventral subclavian loop. 71.The method of claim 70, further comprising re-stimulating the subclavian loop after providing the ablative energy, and further providing ablative energy to the subclavian loop if the biomarker confirms the stimulation. 72.A method according to claim 71, comprising repeating said restimulation and providing further ablation energy until said biomarker no longer confirms stimulation. 73.The method of claim 71, wherein the biomarker comprises at least one of increased dP/dt, arterial blood pressure (systolic, diastolic, mean), pulse pressure variation, increased systemic blood pressure, or increased left ventricular dP/dt. 74.The method of claim 71, wherein the biomarker comprises at least one of a change in heart rate (a change in the R_R or Q_Q interval on a surface ECG), an increase in heart rate (shortening of the RR interval), or a change in the length of the basal cycle on an intracardiac electrogram. 75.The method of claim 71, wherein the biomarker comprises a rheotropic biomarker. 76.A method according to any one of claims 70 to 75, wherein the catheter comprises a balloon, and wherein the neuromodulatory element is positioned on the balloon, and optionally the balloon catheter is inflated an amount to dilate the vasculature of the subject. 77. A method of treating a heart condition in a subject, the method comprising: positioning a neuromodulatory element on an expansion element in a subclavian artery of the subject; inflating the expansion element; electrically stimulating at least one of a dorsal subclavian haptic or a ventral subclavian haptic; Confirming stimulation of the dorsal subclavian loop and/or the ventral subclavian loop by monitoring cardiac parameters; and providing ablation energy to the dorsal subclavian loop and/or the ventral subclavian loop after confirming stimulation of the dorsal subclavian loop and/or the ventral subclavian loop. 78.A method according to any of the preceding claims, wherein the ablation or other neuromodulation is performed using radiofrequency, ultrasound, microwaves, steam, laser, heat, alcohol or other chemicals, cryogenics, or a combination thereof, and wherein the ablation is reversible or irreversible. 79.A method according to claim 78, wherein the electrical stimulation includes using an electrode combination of the neuromodulatory element, and wherein providing ablation energy includes using the same electrode combination of the neuromodulatory element. 80.A method according to any one of claims 77 to 79, wherein the heart disease comprises atrial fibrillation and/or refractory arrhythmias. 81.A signal generator for generating ablation energy according to any one of the preceding claims, having one or more of the features described in the preceding description. 82.A method for regulating and/or ablating the subclavian haptic, having one or more of the features described in the preceding description. 83.A therapeutic system having one or more of the features described in the preceding description. 84.A tissue treatment system having one or more of the features described in the preceding description. | 2022 | ARTHA PARTNERS B V | NA | Start Up | https://www.creditsafe.com/business-index/en-ie/company/artha-partners-bv-nl06033541 | Netherlands | The patent focuses on a system to treat cardiac disease by introducing a catheter with a neuromodulatory element on an expansion element into the subclavian artery and stimulating the dorsal or ventral subclavian loop. It addresses the limitation of conventional cardiac neuromodulation that doesn't target the subclavian periarterial loops. The targeted subclavian neuromodulation aims to influence cardiac autonomic tone. | הפטנט מתמקד במערכת לטיפול במחלת לב על ידי החדרת קטטר עם אלמנט נוירומודולציה על אלמנט התרחבות לתוך עורק תת-בריחי וגירוי הלולאה התת-בריחית הגבית או הבטנית. הוא מתמודד עם המגבלה של נוירומודולציה לבבית קונבנציונלית שאינה ממקדת את הלולאות הפרי-עורקיות התת-בריחיות. הנוירומודולציה התת-בריחית הממוקדת שואפת להשפיע על הטונוס האוטונומי הלבבי. | 1 | [9 empty] | 1 |  |  |  | 1 | [14 empty] | 1 |  |  | ATHER PARTNERS PLC | ATHER PARTNERS PLC | ATHER PARTNERS PLC | ATHER PARTNERS PLC | ARTHA PARTNERS B V¦ATHER PARTNERS PLC | 2023-06-29 | 2023 | 2025-04-15 | 2025 | 2022-07-01 | 2022 | Application | Utility Invention | active | CN20238061603 | 321388172 | AU2023297473A1¦CA3260367A1¦EP4547321A1¦JP2025521829A¦US12484962B2¦US20250160940A1¦WO2024006466A1 | CN
382 | WO2025222080A1 | MEDICAL DEVICES CONFIGURED FOR TEMPERATURE-CONTROLLED THERAPEUTIC ELECTROPORATION | A medical device includes an elongate member and one or more expandable elements positioned on the elongate member, the one or more expandable elements being configured to expand radially outward from the elongate member. The one or more expandable elements have inner surfaces facing the elongate member and outer surfaces opposite the inner surfaces. At least a portion of the one or more expandable elements provide a radiofrequency ablation electrode. A set of temperature sensors is attached to the inner surfaces of at least one of the one or more expandable elements, or on a flexible circuit surrounding at least a portion of the outer surfaces of the one or more expandable elements. | 1. CLAIMSWhat is claimed is: 1. A medical device, comprising: an elongate member; one or more expandable elements positioned on the elongate member, the one or more expandable elements being configured to expand radially outward from the elongate member, the one or more expandable elements having inner surfaces facing the elongate member and outer surfaces opposite the inner surfaces, at least a portion of the one or more expandable elements providing a radiofrequency ablation electrode; and a set of temperature sensors attached to the inner surfaces of at least one of the one or more expandable elements. 2. The medical device of claim 1, wherein the elongate member comprises a delivery catheter. 3. The medical device of claim 1, wherein the one or more expandable elements comprise a wire cage. 4. The medical device of claim 3, wherein the wire cage comprises a plurality of struts, the set of temperature sensors being attached to inner surfaces of at least two different ones of the plurality of struts. 5. The medical device of claim 4 wherein the at least two different ones of the plurality of struts on which the set of temperature sensors are attached are arranged such that when the one or more expandable elements are expanded radially outward from the elongate member, the set of temperature sensors are at two or more different positions of a circumference of the wire cage. 6. The medical device of claim 1, wherein the set of temperature sensors comprise at least four temperature sensors. 7. The medical device of claim 1, wherein the set of temperature sensors comprise thermocouples. 8. The medical device of claim 1, further comprising one or more return electrodes, wherein the radiofrequency ablation electrode and at least one of the one or more return electrodes are configured to perform bipolar radiofrequency ablation. 9. The medical device of claim 8, wherein said at least one of the one or more return electrodes is disposed on the elongate member. 10. The medical device of claim 8, wherein said at least one of the one or more return electrodes is disposed on a guide sheath in which the elongate member is inserted. 11. The medical device of claim 8, wherein said at least one of the one or more return electrodes is disposed on a flexible circuit surrounding at least a portion of the one or more expandable elements. 12. The medical device of claim 1, further comprising a pull wire coupled to a distal tip of the elongate member, the pull wire being configured for retracting the distal tip of the elongate member to expand the one or more expandable elements radially outward from the elongate member. 13. The medical device of claim 1, further comprising a controller configured: to controllably expand the one or more expandable elements to bring the radiofrequency ablation electrode into contact with a wall of a lumen; to monitor a temperature of target tissues in a vicinity of the wall of the lumen utilizing the set of temperature sensors; and to control, based at least in part on the monitored temperature of the target tissues in the vicinity of the wall of the lumen, a current applied to the radiofrequency ablation electrode. 14. The medical device of claim 13, wherein the controller is further configured to monitor a tissue impedance of the target tissues in the vicinity of the wall of the lumen, and wherein the current applied to the radiofrequency ablation electrode is further controlled based at least in part on the monitored tissue impedance. 15. The medical device of claim 1, further comprising a flexible circuit surrounding at least a portion of the one or more expandable elements. 16. The medical device of claim 15, wherein the flexible circuit comprises a plurality of struts, each of at least a subset of the plurality of struts comprising one or more sensing electrodes. 17. The medical device of claim 16, further comprising an isolation layer disposed between the plurality of struts and the one or more expandable elements. 18. The medical device of claim 16, further comprising one or more return electrodes disposed on the flexible circuit. 19. The medical device of claim 16, further comprising a controller configured: to controllably expand the one or more expandable elements to bring the radiofrequency ablation electrode into contact with a wall of a lumen; and to utilize the one or more sensing electrodes of the subset of the plurality of struts of the flexible circuit to determine an extent of ablation of different portions of target tissues in a vicinity of the wall of the lumen, the different portions being positioned at different locations of a circumference of the wall of the lumen. 20. The medical device of claim 19, wherein the controller is further configured to control, based at least in part on the determined extent of the ablation of the different portions of the target tissues, a current applied to the radiofrequency ablation electrode. 21. The medical device of claim 19, wherein the one or more sensing electrodes of the subset of the plurality of struts are configured for selective operation in a therapeutic mode, and wherein the controller is further configured to operate at least one of the one or more sensing electrodes of at least one of the plurality of struts in the therapeutic mode to apply current for radiofrequency ablation of at least one location of the circumference of the wall of the lumen based at least in part on the determined extent of the ablation of the different portions of the target tissues. 22. A medical device comprising: an elongate member; one or more expandable elements positioned on the elongate member, the one or more expandable elements being configured to expand radially outward from the elongate member, the one or more expandable elements having inner surfaces facing the elongate member and outer surfaces opposite the inner surfaces, at least a portion of the one or more expandable elements providing a radiofrequency ablation electrode; and a flexible circuit surrounding at least a portion of the outer surfaces of the one or more expandable elements, the flexible circuit comprising a set of temperature sensors. 23. The medical device of claim 22, wherein the elongate member comprises a delivery catheter. 24. The medical device of claim 22, wherein the one or more expandable elements comprise a wire cage. 25. The medical device of claim 22, wherein the set of temperature sensors comprise at least four temperature sensors. 26. The medical device of claim 22, wherein the set of temperature sensors comprise thermocouples. 27. The medical device of claim 22, further comprising one or more return electrodes, wherein the radiofrequency ablation electrode and at least one of the one or more return electrodes are configured to perform bipolar radiofrequency ablation. 28. The medical device of claim 27, wherein said at least one of the one or more return electrodes is disposed on the elongate member. 29. The medical device of claim 27, wherein said at least one of the one or more return electrodes is disposed on a guide sheath in which the elongate member is inserted. 30. The medical device of claim 27, wherein said at least one of the one or more return electrodes is disposed on the flexible circuit. 31. The medical device of claim 22, further comprising a pull wire coupled to a distal tip of the elongate member, the pull wire being configured for retracting the distal tip of the elongate member to expand the one or more expandable elements radially outward from the elongate member. 32. The medical device of claim 22, further comprising a controller configured: to controllably expand the one or more expandable elements to bring the radiofrequency ablation electrode into contact with a wall of a lumen; to monitor a temperature of target tissues in a vicinity of the wall of the lumen utilizing the set of temperature sensors; and to control, based at least in part on the monitored temperature of the target tissues in the vicinity of the wall of the lumen, a current applied to the radiofrequency ablation electrode. 33. The medical device of claim 32, wherein the controller is further configured to monitor a tissue impedance of the target tissues in the vicinity of the wall of the lumen, and wherein the current applied to the radiofrequency ablation electrode is further controlled based at least in part on the monitored tissue impedance. 34. The medical device of claim 22, wherein the flexible circuit comprises a plurality of struts, each of at least a subset of the plurality of struts comprising one or more sensing electrodes. 35. The medical device of claim 34, further comprising an isolation layer disposed between the plurality of struts and the one or more expandable elements. -SO- 36. The medical device of claim 35 wherein the set of temperature sensors are attached to at least two different ones of the plurality of struts of the flexible circuit such that when the one or more expandable elements are expanded radially outward from the elongate member, the set of temperature sensors are at two or more different positions of a circumference of a wall of a lumen in which the medical device is deployed. 37. The medical device of claim 34, further comprising a controller configured: to controllably expand the one or more expandable elements to bring the radiofrequency ablation electrode into contact with a wall of a lumen; and to utilize the one or more sensing electrodes of the subset of the plurality of struts of the flexible circuit to determine an extent of ablation of different portions of target tissues in a vicinity of the wall of the lumen, the different portions being positioned at different locations of a circumference of the wall of the lumen. 38. The medical device of claim 37, wherein the controller is further configured to control, based at least in part on the determined extent of the ablation of the different portions of the target tissues, a current applied to the radiofrequency ablation electrode. 39. The medical device of claim 37, wherein the one or more sensing electrodes of the subset of the plurality of struts are configured for selective operation in a therapeutic mode, and wherein the controller is further configured to operate at least one of the one or more sensing electrodes of at least one of the plurality of struts in the therapeutic mode to apply current for radiofrequency ablation of at least one location of the circumference of the wall of the lumen based at least in part on the determined extent of the ablation of the different portions of the target tissues. 40. A method, comprising: delivering an elongate member of a medical device to a target site within a lumen; expanding one or more expandable elements positioned on the elongate member radially outward from the elongate member such that outer surfaces of the one or more expandable elements contact a wall of the lumen, at least a portion of the one or more expandable elements providing a radiofrequency ablation electrode; monitoring a temperature of target tissues in a vicinity of the wall of the lumen utilizing a set of temperature sensors attached to inner surfaces of at least one of the one or more expandable elements; and controlling a current applied to the radiofrequency ablation electrode based at least in part on the monitored temperature of the target tissues in the vicinity of the wall of the lumen. 41. The method of claim 40, wherein the one or more expandable elements comprise a wire cage, wherein the wire cage comprises a plurality of struts, and wherein the set of temperature sensors being attached to inner surfaces of at least two different ones of the plurality of struts. 42. The method of claim 41, wherein the at least two different ones of the plurality of struts on which the set of temperature sensors are attached are arranged such that when the one or more expandable elements are expanded radially outward from the elongate member, the set of temperature sensors are at two or more different positions of a circumference of the wire cage. 43. The method of claim 40, wherein expanding the one or more expandable elements radially outward from the elongate member comprises retracting a pull wire coupled to a distal tip of the elongate member. 44. The method of claim 40, further comprising determining an extent of ablation of different potions of target tissues in a vicinity of the wall of the lumen utilizing sensing electrodes disposed on at least a subset of a plurality of struts of a flexible circuit surrounding at least a portion of the one or more expandable elements, the different portions being positioned at different locations of a circumference of the wall of the lumen. 45. The method of claim 44, wherein controlling the current applied to the radiofrequency ablation electrode is further based at least in part on the determined extent of the ablation of the different portions of the target tissues. 46. The method of claim 44, wherein the one or more sensing electrodes of the subset of the plurality of struts are configured for selective operation in a therapeutic mode, and further comprising operating at least one of the one or more sensing electrodes of at least one of the plurality of struts in the therapeutic mode to apply current for radiofrequency ablation of at least one location of the circumference of the wall of the lumen based at least in part on the determined extent of the ablation of the different portions of the target tissues. 47. The method of claim 40, further comprising monitoring a tissue impedance of the target tissues in the vicinity of the wall of the lumen, and wherein the current applied to the radiofrequency ablation electrode is further controlled based at least in part on the monitored tissue impedance. 48. A method comprising: delivering an elongate member of a medical device to a target site within a lumen; expanding one or more expandable elements positioned on the elongate member radially outward from the elongate member such that outer surfaces of the one or more expandable elements contact a wall of the lumen, at least a portion of the one or more expandable elements providing a radiofrequency ablation electrode; monitoring a temperature of target tissues in a vicinity of the wall of the lumen utilizing a set of temperature sensors attached to a flexible circuit surrounding at least a portion of the outer surfaces of the one or more expandable elements; and controlling a current applied to the radiofrequency ablation electrode based at least in part on the monitored temperature of the target tissues in the vicinity of the wall of the lumen. 49. The method of claim 48, wherein the one or more expandable elements comprise a wire cage. 50. The method of claim 48, wherein the flexible circuit comprises a plurality of struts, and wherein the set of temperature sensors are attached to at least two different ones of the plurality of struts of the flexible circuit such that when the one or more expandable elements are expanded radially outward from the elongate member, the set of temperature sensors are at two or more different positions of a circumference of a wall of a lumen in which the medical device is deployed. 51. The method of claim 48, further comprising determining an extent of ablation of different potions of target tissues in a vicinity of the wall of the lumen utilizing sensing electrodes disposed on at least a subset of a plurality of struts of the flexible circuit, the different portions being positioned at different locations of a circumference of the wall of the lumen. 52. The method of claim 51, wherein controlling the current applied to the radiofrequency ablation electrode is further based at least in part on the determined extent of the ablation of the different portions of the target tissues. 53. The method of claim 51, wherein the one or more sensing electrodes of the subset of the plurality of struts are configured for selective operation in a therapeutic mode, and further comprising operating at least one of the one or more sensing electrodes of at least one of the plurality of struts in the therapeutic mode to apply current for radiofrequency ablation of at least one location of the circumference of the wall of the lumen based at least in part on the determined extent of the ablation of the different portions of the target tissues. 54. The method of claim 48, further comprising monitoring a tissue impedance of the target tissues in the vicinity of the wall of the lumen, and wherein the current applied to the radiofrequency ablation electrode is further controlled based at least in part on the monitored tissue impedance. | 2024 | AUTONOMIX MEDICAL | Medical and Healthcare | Small-size company | https://www.autonomix.com/ | USA | The patent focuses on a medical electroporation device with expandable elements that radially deploy and provide an RF ablation electrode surface, plus temperature sensors. It addresses the difficulty of delivering electroporation while controlling unwanted heating in target tissue. Temperature data lets the device deliver electroporation while controlling unwanted heating. | הפטנט מתמקד במכשיר אלקטרופורציה רפואי עם אלמנטים מתרחבים הנפרשים רדיאלית ומספקים משטח אלקטרודת אבלציית RF, בתוספת חיישני טמפרטורה. הוא מתמודד עם הקושי במסירת אלקטרופורציה תוך בקרת חימום בלתי רצוי ברקמת המטרה. נתוני טמפרטורה מאפשרים למכשיר לספק אלקטרופורציה תוך בקרת חימום בלתי רצוי. | 1 |  |  |  | 1 | [5 empty] | 1 | [13 empty] | 1 | [7 empty] | AUTONOMIX MEDICAL INC | AUTONOMIX MEDICAL INC | AUTONOMIX MEDICAL INC | AUTONOMIX MEDICAL INC | AUTONOMIX MEDICAL INC | 2025-04-18 | 2025 | 2025-10-23 | 2025 | 2024-04-19 | 2024 | Application | Utility Invention | active | WO2025US25299 | 324041880 | [NA] | WO
383 | CN118177977A | Intelligent minimally invasive diagnosis and treatment integrated device and method for common-scanning-path wide-field imaging | According to the intelligent minimally invasive diagnosis and treatment integrated device and method for common-scanning-path wide-field imaging, a mechanical arm controls an optical interference tomography module to obtain a real-time optical tomography image in a biological tissue operation, pathological boundary and volume capacity information of a lesion is recognized, and structural information of the lesion tissue is obtained; calculating tissue optical characteristics and diagnosis information; the multi-dimensional controllable intelligent optical minimally invasive diagnosis and treatment module obtains diagnosis and treatment parameters in minimally invasive diagnosis and treatment and constructs a laser tissue ablation biological effect model; the integrated common scanning path diagnosis and treatment mapping module constructs an intelligent diagnosis and treatment mapping relation according to the position information of the lesion tissue to control the corresponding light beams to perform multi-dimensional scanning; the intelligent optical detection and probe control module controls the mechanical arm to control the minimally invasive diagnosis and treatment time sequence of the optical interference tomography module and the photo-thermal diagnosis and treatment and conduct real-time diagnosis and treatment guidance. The device is simple in structure, accurate in precision, easy to operate, high in imaging speed, small in size, portable and high in intelligent degree. | 1· An intelligent minimally invasive diagnosis and treatment integrated device with common scanning path wide-field imaging, characterized in that it includes: a robotic arm controlled optical interference tomography module, used to obtain real-time optical tomography images of biological tissues during surgery, and identify the pathological boundaries and volume capacity information of the lesions from the real-time optical tomography images, obtain the structural information of the lesion tissue, and calculate the optical properties and diagnostic information of the tissue; a multi-dimensional controllable intelligent optical minimally invasive diagnosis and treatment module, used to obtain the diagnosis and treatment parameters in minimally invasive diagnosis and treatment, and pre-construct a laser tissue ablation biological effect model with the structural information of the lesion tissue and laser parameters; an integrated common scanning path diagnosis and treatment mapping module, used to control the corresponding light beam to perform multi-dimensional scanning according to the constructed intelligent minimally invasive diagnosis and treatment mapping relationship based on the position information of the lesion tissue; and an intelligent optical detection and probe control module, used to control the robotic arm to control the optical interference tomography module and the minimally invasive diagnosis and treatment timing of photothermal diagnosis and treatment, and perform real-time guidance of diagnosis and treatment. 2.The device according to claim 1 is characterized in that the robotic arm controls the optical interference tomography module to obtain high-resolution optical diagnostic images for the physical structure of the tissue under a wide-field imaging structure, and the high-resolution optical diagnostic images include depth structure images and processed tissue elasticity images. 3.The device according to claim 1 or 2 is characterized in that the robotic arm-controlled optical interference tomography module includes a robotic arm-controlled two-dimensional galvanometer or MEMS device with a controllable scanning range. 4.The device according to claim 3 is characterized in that the multi-dimensional controllable intelligent optical minimally invasive diagnosis and treatment module is specifically used to control the scanning laser beam through the oblique incident MEMS so that the scanning area coincides with the scanning range of the two-dimensional galvanometer of the optical interference tomography. 5. The device according to claim 1 is characterized in that the calculation formula of the coordinate mapping relationship of the intelligent minimally invasive diagnosis and treatment wide-field imaging is: endu * camera μ * object μ - object μ base11 * endn * camera*1 - base11 end u , oct μ . object μ - object μ base^ ' md ' octn - basen Among them, b is the coordinate transformation matrix from the base of the manipulator to the end of the manipulator, which is obtained by the 5-point calibration of the manipulator, b is the coordinate transformation matrix from the end of the manipulator to the camera, h is the transformation relationship between the calibration plate and the camera, which is obtained by Zhang's calibration, mh is the coordinate transformation matrix from the base of the manipulator to the calibration plate, m is the coordinate transformation matrix from the end of the manipulator to the OCT probe, and a2 is the coordinate transformation matrix from the OCT probe to the calibration plate; The coordinate system relationship between the two systems before and after changing the posture of the manipulator can be written as: endlu * cameralu * object μ _ end2u * earner a2 μ * object only base*1 * endln * cameraln - base*1 * end2n * camera2n endlu * octlu * object μ - end2u * oct2u * object μ basen endln octln - basen end2n oct2n Among them, H is the coordinate conversion matrix from the base of the manipulator to the end of the manipulator before changing the posture of the manipulator, θ is the coordinate conversion matrix from the end of the manipulator to the binocular camera before changing the posture of the manipulator, θ is the coordinate conversion matrix from the calibration plate to the camera before changing the posture of the manipulator, θ is the coordinate conversion matrix from the base of the manipulator to the end of the manipulator after changing the posture of the manipulator, θ is the coordinate conversion matrix from the end of the manipulator to the binocular camera after changing the posture of the manipulator, and θ is the coordinate conversion matrix from the calibration plate to the camera after changing the posture of the manipulator. is the coordinate transformation matrix from the end of the manipulator to the OCT probe before the manipulator posture is changed, is the coordinate transformation matrix from the OCT probe to the calibration plate before the manipulator posture is changed, is the coordinate transformation matrix from the end of the manipulator to the OCT probe after the manipulator posture is changed, is the coordinate transformation matrix from the OCT probe to the calibration plate after the manipulator posture is changed; Transform the equations to obtain the following equations: end2u-l . endlu . cameralu - camera2μ * object object base*1 ' base*1 ' endln - end2n * camera2n cameraln end2 zj-1 * color strong 1 ear * octi u oct2 u * object μ * object μτ-± basen base11 endln - end2n od2n octln The equation simplifies to: A·X=X·B Μ·Y=Y·N Solve two equations, the unknown x=cam or NH, Y=mengH is the hand-eye relationship between the camera and the OCT probe and the end of the robotic arm. To plan the path and control the scanning area of the OCT probe through the information collected by the camera, it is necessary to find the rotation displacement relationship between the camera and the OCT probe, which can be obtained through X and y: camera_ χ . yT 6·An intelligent minimally invasive diagnosis and treatment integrated method with common scanning path wide-field imaging, characterized in that it includes the following steps: obtaining real-time optical tomography images of biological tissues during surgery, and identifying the pathological boundaries and volume capacity information of the lesions from the real-time optical tomography images, obtaining the structural information of the lesion tissues, and calculating the optical properties and diagnostic information of the tissues; obtaining the diagnosis and treatment parameters in minimally invasive diagnosis and treatment, and pre-constructing a laser tissue ablation biological effect model with the structural information of the lesion tissues and laser parameters; controlling the corresponding light beams to perform multi-dimensional scanning based on the constructed intelligent minimally invasive diagnosis and treatment mapping relationship based on the position information of the lesion tissues; and controlling the robotic arm to control the minimally invasive diagnosis and treatment timing of optical interference tomography and photothermal diagnosis and treatment, and performing real-time guidance of diagnosis and treatment. 7.The method according to claim 6 is characterized in that the acquisition of real-time optical tomographic images of biological tissues during surgery includes: acquiring high-resolution optical diagnostic images for the physical structure of the tissue under a wide-field imaging structure, wherein the high-resolution optical diagnostic images include depth structure images and processed tissue elasticity images. 8.The method according to claim 6 or 7 is characterized in that a two-dimensional galvanometer or MEMS device is controlled by a robotic arm with a controllable scanning range to obtain real-time optical tomographic images of biological tissue during surgery. 9.The method according to claim 8 is characterized in that it also includes: controlling the scanning laser beam through an obliquely incident MEMS so that the scanning area coincides with the scanning range of the two-dimensional galvanometer of the optical interference tomography. 10. The method according to claim 6 is characterized in that the calculation formula of the coordinate mapping relationship of the intelligent minimally invasive diagnosis and treatment wide-field imaging is: end μ , camera μ . object Dan_ object basen endn ' cameran ~ base11 end ft,oct ft,obj&ct if - handsome% base ' endn 'octn - basen Among them, a is the coordinate transformation matrix from the base of the manipulator to the end of the manipulator, which is obtained through the 5-point calibration of the manipulator; m is the coordinate transformation matrix from the end of the manipulator to the camera, which is obtained through Zhang's calibration; a is the coordinate transformation matrix from the base of the manipulator to the calibration plate; h is the coordinate transformation matrix from the end of the manipulator to the OCT probe, which is obtained through Zhang's calibration; h is the coordinate transformation matrix from the end of the manipulator to the OCT probe; the coordinate system relationship between the two systems before and after changing the posture of the manipulator can be written as: endlu . earner alii . object μ _ end2 u . camera2 u . object μ basen ' ' cameraln - base11 *2nd2 port*camera2n endl jj octi it object Tj - end2 rj oct2 jj object Tj basen * endln 'octi-base11 * end2n *oct2n Among them, H is the coordinate transformation matrix from the base of the manipulator to the end of the manipulator before the manipulator posture is changed, h is the coordinate transformation matrix from the end of the manipulator to the binocular camera before the manipulator posture is changed, h is the coordinate transformation matrix from the calibration plate to the camera before the manipulator posture is changed, s is the coordinate transformation matrix from the base of the manipulator to the end of the manipulator after the manipulator posture is changed, s is the coordinate transformation matrix from the end of the manipulator to the binocular camera after the manipulator posture is changed, s is the coordinate transformation matrix from the end of the manipulator to the OCT probe before the manipulator posture is changed, s is the coordinate transformation matrix from the OCT probe to the calibration plate before the manipulator posture is changed, s is the coordinate transformation matrix from the end of the manipulator to the OCT probe after the manipulator posture is changed, s is the coordinate transformation matrix from the OCT probe to the calibration plate after the manipulator posture is changed; Transform the equations to obtain the following equations: end2u-i * endl u * camerai u = camera! u * objectobject basen base11 endln endzn camera2n cameraln end2 u-1 * endl u * octi u - oct2 u object μ object rr-± basen basen endln - end2n oct2n octln The equation is simplified to A·X=X·B Μ·Y=Y·N. Solving the two equations, the unknowns x= aminorthH and y=mengH are the hand-eye relationship between the camera and the OCT probe and the end of the robotic arm. To plan the path and control the scanning area of the OCT probe through the information collected by the camera, it is necessary to find the rotational displacement relationship between the camera and the OCT probe. Through X and y, we can get: camerai^ _ χ . yT | 2024 | Beijing Institute of Technology | University | NA | https://www.bit.edu.cn/ | China | The patent focuses on an intelligent minimally invasive diagnosis-and-treatment device using a robotic arm controlling an OCT module that scans tissue across a wide field with a common scanning path. It addresses the need for wide-field, high-resolution tissue imaging during minimally invasive procedures, where OCT typically has narrow field of view. OCT identifies pathological boundaries in real time during procedures. | הפטנט מתמקד במכשיר אבחון וטיפול זעיר-פולשני חכם המשתמש בזרוע רובוטית השולטת במודול OCT הסורק רקמה על פני שדה רחב עם נתיב סריקה משותף. הוא מתמודד עם הצורך בהדמיית רקמה ברזולוציה גבוהה ובשדה רחב במהלך פרוצדורות זעיר-פולשניות, שבהן ל-OCT בדרך כלל יש שדה ראייה צר. OCT מזהה גבולות פתולוגיים בזמן אמת במהלך הפרוצדורות. |  | 1 | [11 empty] | 1 | [6 empty] | 1 | 1 | [8 empty] | 1 |  | BEIJING INSTITUTE TECH | Beijing Institute of Technology | BEIJING INSTITUTE TECH | Beijing Institute of Technology | Beijing Institute of Technology | 2024-02-05 | 2024 | 2024-06-14 | 2024 | 2024-02-05 | 2024 | Application | Utility Invention | active | CN202410166876 | 313174154 | [NA] | CN
384 | CN118177950A | Tumor electrical impedance change analysis and electric pulse parameter optimization method | The invention discloses a tumor electrical impedance change analysis and electric pulse parameter optimization method, which comprises the following steps: S100, applying irreversible electroporation pulse stimulation to a to-be-detected tumor tissue to form nanoscale micropores on the cell surface of the to-be-detected tumor tissue; s200, collecting electrical impedance information of the tumor tissue to be detected; s300, reconstructing an electrical impedance image for the electrical impedance information by using a time domain difference or frequency domain difference algorithm, and quantitatively describing the tumor tissue to be detected; s400, selecting parameters to be optimized from the pulse parameters and the electrode parameters, performing statistical model modeling by adopting a polynomial regression method, and establishing a target function; s500, selecting a target function needing to be optimized according to the optimization target, and establishing an optimization function; and S600, solving the optimization function to obtain an optimal pulse parameter and an optimal electrode parameter. According to the invention, the tumor can be eliminated in the maximum area, and the surrounding normal tissues can be minimally damaged at the same time. | 1A method for analyzing tumor electrical impedance changes and optimizing electrical pulse parameters, comprising: S100, applying irreversible electroporation pulse stimulation to the tumor tissue to be tested, so that nanoscale micropores are formed on the cell surface of the tumor tissue to be tested; S200, collecting electrical impedance information of the tumor tissue to be tested; S300, using a time domain difference or frequency domain difference algorithm on the electrical impedance information to reconstruct an electrical impedance image and quantitatively describe the tumor tissue to be detected; S400, selecting parameters to be optimized from the pulse parameters and the electrode parameters, using a polynomial regression method to perform statistical model building, and establishing an objective function; S500, according to the optimization target, select the objective function to be optimized and establish the optimization function; S600, solving the optimization function to obtain the optimal pulse parameters and electrode parameters. 2.The method for analyzing tumor electrical impedance changes and optimizing electric pulse parameters guided by nanoscale micropores according to claim 1 is characterized in that, in S100, the irreversible electroporation pulse stimulation includes at least one of single pulse stimulation, continuous stimulation, and specific frequency stimulation. 3The method for analyzing electrical impedance changes of tumors and optimizing electrical pulse parameters guided by nanoscale micropores according to claim 1 is characterized in that, in S200, collecting electrical impedance information of the tumor tissue to be tested comprises: S201, prepare the electrical impedance measurement equipment and ensure that the equipment is in normal working condition; S202, setting initial values of pulse parameters according to the size, shape and conductivity distribution of the tumor tissue to be measured; S203, applying irreversible electroporation pulse stimulation to the tumor tissue to be measured by the electrical impedance measuring device, so that current passes through the tumor tissue to be measured, and recording electrical impedance information. 4The method for analyzing tumor electrical impedance changes and optimizing electrical pulse parameters guided by nanoscale micropores according to claim 1 is characterized in that, in S300, the electrical impedance information is subjected to a time domain difference algorithm to reconstruct an electrical impedance image, and the quantitative description of the tumor tissue to be measured includes: S3011, converting the electrical impedance information into an electrical impedance digital signal; S3012, using the collected electrical impedance digital signal as a time series signal V (t); S3013, sampling the sampling points of the time series signal V(t), the number of sampling points is n, and obtaining a discrete time point electrical impedance sampling value sequence V(nΤ), where T is a sampling interval; S3014, calculating a leading differential sequence AV(n) =V((n+1)Τ) -V(nt) for the electrical impedance sampling value sequence V(nT); S3015, performing weighted summation on the leading differential sequence AV (n) to obtain a weighted summation value S (n) = a·Δv(n) + (1-a)·S(nl), where a is a smoothing coefficient used for weighting, and has a value range of 0 to 1; S3016, repeatedly executing the calculation of the leading difference sequence and performing weighted summation on the leading difference sequence, and performing iterative summation until the last sampling point is processed; S3017, reconstruct according to the S (η) values of all sampling points to obtain the electrical impedance image reconstruction result. 5The method for analyzing tumor electrical impedance changes and optimizing electric pulse parameters guided by nanoscale micropores according to claim 1 is characterized in that, in S300, the electrical impedance information is subjected to a frequency domain difference algorithm to reconstruct an electrical impedance image, and the quantitative description of the tumor tissue to be measured includes: S3021, converting the electrical impedance information into an electrical impedance digital signal; S3022, convert the collected electrical impedance digital signal V (n) from the time domain to the frequency domain, perform discrete Fourier transform, and obtain the frequency domain signal V(n) = Σ2/n, where n is the number of sampling points, and the value is repeatedly taken with n1. The value corresponds to each discrete sampling point in the time domain signal, k is the number of frequency units, which are repeated from 0 to N-1, j is the imaginary unit, and e is the Napier constant; S3023, taking the modulus of the frequency domain signal V(k) to obtain the amplitude spectrum A(k) = ¦V(k) I of each frequency unit; S3024, performing a differential operation on the amplitude spectrum to obtain a differential value of the frequency domain signal between adjacent frequency points AA (k) = A(k+1) -A(k); S3025, repeatedly perform the differential operation until deep differential analysis of the frequency domain signal is achieved, and obtain the frequency domain differential result S (k) = aX AA(k) + (la) XS (kl), where a is a smoothing parameter; S30 26, perform inverse Fourier transform on the difference result S (k) to obtain the reconstructed time domain signal S (n) = (1 / M L: F (k) m), double, this group of sampling points is the electrical impedance image reconstruction result. 6.The method for analyzing tumor electrical impedance changes and optimizing electric pulse parameters guided by nanoscale micropores according to claim 3 is characterized in that, in S400, the pulse parameters include at least one of pulse voltage, pulse width, number of pulses and pulse frequency. 7.The method for analyzing tumor electrical impedance changes and optimizing electric pulse parameters guided by nanoscale micropores according to claim 6 is characterized in that, in S400, the electrode parameters include at least one of the electrode needle spacing, the number of electrode needles, the electrode needle length and the electrode needle diameter. 8The method for analyzing the change of tumor electrical impedance and optimizing electric pulse parameters guided by nanoscale micropores according to claim 7 is characterized in that, in S400, the parameters to be optimized are selected from the pulse parameters and the electrode parameters, and the statistical model is built by the polynomial regression method, and the objective function is established, including: S401, selecting a parameter to be optimized from pulse parameters and electrode parameters, wherein the parameter to be optimized includes at least one of pulse voltage U/V, number of electrode needles n, electrode needle spacing D/mm, and electrode needle length L/mm; S402, using the polynomial regression method, according to the form of the multivariate quadratic regression equation, establish the objective function y = ko+k]X1+··· +3[+Jinnai2+·.·+1</;+1<6;+benevolence2kj2+...+One tile high one high, where y is the response value, Xj is the variable of the parameter to be optimized, i is the number of parameters to be optimized, and k is the regression coefficient. 9.The method for analyzing tumor electrical impedance changes and optimizing electric pulse parameters guided by nanoscale micropores according to claim 8 is characterized in that, in S500, according to the optimization target, the objective function to be optimized is selected and the optimization function is established; S501, determining an optimization goal of achieving maximum ablation area with minimum damage to surrounding normal tissues; S502, establish the objective function of the ablation area f(x) =&0+&ιυ+&2N+&3Ο+&4ί+&5υ2+&6Ν2+&7Ο2+&8ί2+&9υϋ+&10α+&1ιυΝ+&12ΝΏ+&13 NL, where a. to ai4 are regression coefficients; S503, establish the objective function of damage degreef2(X)=b0+b1U+b2N+b3D+b4L+b5U2+b6N2+b7D2+b8L2+b9UD+b10UL+b11UN+b12ND+b13 NL, where b0 to s are regression coefficients; S504, establish optimization function Max f 1(X), Min f2 (x). 10. The method for analyzing tumor electrical impedance changes and optimizing electric pulse parameters guided by nanoscale micropores according to claim 9, characterized in that, 5600φ, solving the optimization function to obtain the optimal pulse parameters and electrode parameters includes: S601, limiting the pulse voltage U/V to a range of 1000WUW2500, and the number of electrode needles N to a range of 1WN W5, the electrode needle spacing D/mm ranges from 4WDW13, and the electrode needle length L/mm ranges from 4WLW13; S602, solving the optimization function according to the limited range to obtain the optimal pulse parameters and electrode parameters. | 2024 | BEIJING ZHONGLIAN GUOCHENG TECH | NA | NA | NA | China | The patent focuses on a method that applies IRE pulses to a tumor, measures impedance, reconstructs an electrical impedance image, and analyzes the response to optimize pulse parameters. It addresses the lack of feedback-driven IRE pulse optimization that uses the tumor's own impedance response. Optimization of pulse parameters is then driven by these tumor impedance changes. | הפטנט מתמקד בשיטה המיישמת פולסי IRE על גידול, מודדת אימפדנס, משחזרת תמונת אימפדנס חשמלי ומנתחת את התגובה לאופטימיזציה של פרמטרי הפולס. הוא מתמודד עם החוסר באופטימיזציית פולס IRE מונעת היזון חוזר המשתמשת בתגובת האימפדנס של הגידול עצמו. אופטימיזציה של פרמטרי הפולס ואז מונעת על ידי שינויי אימפדנס גידול אלו. | 1 | [10 empty] | 1 | [7 empty] | 1 |  | 1 | [6 empty] | 1 |  |  |  | BEIJING ZHONGLIAN GUOCHENG TECH CO LTD | BEIJING ZHONGLIAN GUOCHENG TECH CO LTD | BEIJING ZHONGLIAN GUOCHENG TECH CO LTD | BEIJING ZHONGLIAN GUOCHENG TECH CO LTD | BEIJING ZHONGLIAN GUOCHENG TECH CO LTD | 2024-03-14 | 2024 | 2024-06-14 | 2024 | 2024-03-14 | 2024 | Application | Utility Invention | active | CN202410295411 | 313174656 | [NA] | CN
385 | CN119131797A | Cardiac ultrasonic intelligent recognition system and method for myocardial ablation | The invention discloses a cardiac ultrasound intelligent recognition system for myocardial ablation, which relates to the technical field of ultrasound medicine, and comprises a cardiac echocardiogram acquisition module, a neural network construction module, a neural network training module, a neural network recognition module and an image processing module, and is used for assisting in setting ablation parameters in an ablation process. Comprise output power, time, electrode position and the like. Wherein the cardiac echocardiogram acquisition module comprises an ultrasonic diagnosis device and an ablation instrument, compared with a traditional separated device, the cardiac echocardiogram intelligent identification system is higher in integration degree and more convenient to use, in addition, the invention further discloses a cardiac ultrasound intelligent identification method for myocardial ablation, an echocardiogram two-dimensional dynamic image of a detected object is acquired, and the detection accuracy is improved. The two-dimensional dynamic image of the echocardiogram of the detected object is input into the convolutional neural network model, the features of each section of the ultrasound image are automatically recognized and positioned, and the problem that the score deviation is large due to too strong subjective factors can be avoided by adopting the method. | 1.A cardiac ultrasound intelligent recognition system for myocardial ablation, characterized by: comprising a cardiac ultrasound acquisition module, for collecting and acquiring a two-dimensional dynamic image of the ultrasound cardiogram of the object to be detected; a neural network construction module, for constructing a convolutional neural network model; a neural network training module, the neural network training module is used to train the convolutional neural network model; a neural network recognition module, using the trained convolutional neural network model to process and recognize the two-dimensional dynamic image of the ultrasound cardiogram, and automatically screen out the end-diastolic frame image of each standard section; an image processing module, for performing image processing on each recognized frame image, automatically identifying, locating, and marking the mitral valve, tricuspid valve, aortic valve, and left and right atria and chambers on the image, and measuring the inner diameter of the left and right atria and chambers and the thickness of each chamber wall on the standard section; an image quality evaluation module, for evaluating and analyzing the input two-dimensional dynamic image of the ultrasound cardiogram, and performing quantitative scoring according to whether it contains the end-diastolic frame image of the standard section required for cardiac ultrasound examination, the clarity of the image, and the fit of the measured data; The cardiac ultrasound cardiogram acquisition module includes an ultrasonic diagnostic device and an ablation device. The lower end of the ultrasonic diagnostic device is fixedly connected to a plurality of vertical poles, and the upper ends of the vertical poles are fixedly connected to a circular track. The ablation device is arranged on a mounting mechanism; the mounting mechanism includes a placement plate for mounting the ablation device, a first drive component for driving the placement plate to rotate, and a second drive component for driving the first drive component to slide on the circular track. 2.According to claim 1, a cardiac ultrasound intelligent recognition system for myocardial ablation is characterized in that: the neural network training module specifically includes the following units: a training sample acquisition unit, used to acquire training samples; a neural network training unit, used to input the training set in the training sample into the initial convolutional neural network to obtain an initial convolutional neural network model; an output result acquisition unit, used to input the verification set in the training sample into the initial convolutional neural network model for processing to obtain an output result; a result comparison unit, used to compare the output result with the verification set image label to obtain a comparison result; a parameter optimization unit, used to optimize the parameters of the initial convolutional neural network model using the obtained comparison result to obtain a trained convolutional neural network model. 3.According to claim 1, a cardiac ultrasound intelligent recognition system for myocardial ablation is characterized in that: the neural network construction module specifically includes the following units: a convolution layer construction unit, which applies a preset convolution kernel according to the received echocardiographic data, performs convolution calculation according to the convolution step size, and obtains a convolution result; a pooling layer construction unit, which performs aggregation and statistical processing on the received convolution result to obtain a pooling result; a fully connected layer construction unit, which integrates the pooling results of the received local information to obtain global information. 4.According to claim 1, a cardiac ultrasound intelligent recognition system for myocardial ablation is characterized in that: the image processing module specifically includes the following units: an image preprocessing unit, used to preprocess the ultrasound cardiogram obtained from the object to be detected to obtain a preprocessed image; an image recognition unit, used to identify and extract corresponding cardiac standard section frames from the preprocessed image according to the image features of each standard section of the heart; an image structure positioning unit, used to determine the positions of the mitral valve, tricuspid valve, aorta, left and right atrioventricular cavities of various cardiac structures in the preprocessed image according to the image features of each standard section of the heart, and calibrate the corresponding structures; The image structure measurement unit is used to measure the inner diameters of the left and right atria and ventricles and the thickness of the ventricular septum of each structure of the heart in the preprocessed image according to the image features of each standard section of the heart. 5.A cardiac ultrasound intelligent recognition system for myocardial ablation according to claim 1 is characterized in that: the second driving component includes an arc-shaped slider, a mounting plate, a plug-in plate plugged into the side of the arc-shaped slider, and a second driving module for driving the arc-shaped slider to slide, the lower end of the arc-shaped slider is provided with a slot that can be plugged into the circular track, the side of the circular track is provided with a slide groove for plugging the plug-in plate, the upper half of the arc-shaped slider is provided with a through hole, the mounting plate is slidably arranged in the through hole along the up and down directions, and the driving module is arranged at the lower end of the mounting plate located in the circular track. 6.According to claim 5, a cardiac ultrasound intelligent recognition system for myocardial ablation is characterized in that: a fixing rod is fixedly connected to the lower end of the mounting plate outside the annular track, a through groove for plugging the fixing rod is provided at the end of the plug plate away from the slide groove, the lower end of the fixing rod is an inverted right-angled trapezoid, the upper end of the mounting plate is fixedly connected to a tension spring, and the upper end of the tension spring is fixedly connected to the upper end of the arc-shaped slider. 7.A cardiac ultrasound intelligent recognition system for myocardial ablation according to claim 6, characterized in that: the second driving module includes a first motor fixedly connected to the lower end of the mounting plate, a driving gear fixedly connected to the output shaft of the first motor, and an inner gear ring fixedly connected to the inner side of the annular track, and the inner gear ring and the driving gear are meshed and connected; when the fixing rod is temporarily fixed to the mounting plate, the inner gear ring and the driving gear are meshed and connected, and when not temporarily fixed, the inner gear ring and the driving gear are not meshed. 8.A cardiac ultrasound intelligent recognition system for myocardial ablation according to claim 7, characterized in that: the first driving assembly includes a circular plate, a rectangular groove provided on the circular plate, a telescopic rod fixed in the rectangular groove, a second motor fixed in the arc-shaped slider, an inner polygonal sleeve rotatably connected to the middle of the circular plate, and a fan-shaped gear fixedly connected to the outer side of the inner polygonal sleeve; the placement plate is provided with a circular hole for matching the installation of the circular plate, and the placement plate at the upper end of the circular hole is provided with an annular tooth groove for engaging the fan-shaped gear; the upper end of the mounting plate is fixedly connected with an insertion rod that can be plugged into the rectangular groove, and the insertion rod is provided with a limiting groove for the insertion of the telescopic rod; the output shaft of the second motor is fixedly connected with a polygonal shaft for matching the installation of the inner polygonal sleeve; the upper end of the arc-shaped slider is provided with a mounting groove for mounting the second motor; the placement plate at the lower end of the circular hole is fixedly connected with a lower ring plate, and the upper end of the outer surface of the circular plate is fixedly connected with an upper ring plate, and the upper ring plate is used to limit the lower ring plate on the arc-shaped slider; the placement plate is fixed with a bandage for mounting the ablation device. 9. A cardiac ultrasound intelligent recognition method for myocardial ablation, comprising an intelligent recognition system for cardiac ultrasound examination according to any one of claims 1 to 8, characterized in that it comprises the following steps: Step S1, move the ultrasonic diagnostic device to the treatment position, use the ultrasonic diagnostic device to determine the position of the lesion, and then control the first motor and the second motor to adjust the position of the ablation device, and plan the path of the ablation needle or catheter; Step S2, disinfect and locally anesthetize the puncture site; Step S3, under ultrasonic guidance, accurately insert the ablation needle or catheter into the lesion area, start the ablation device, set the ablation parameters according to the treatment plan, start the ablation process, and continuously use the ultrasonic diagnostic device for monitoring during the entire ablation process to ensure that the position of the ablation needle or catheter is correct, while monitoring changes in tissue temperature; Step S4, obtain a two-dimensional dynamic image of the ultrasonic cardiogram of the detection object; Step S5, input the two-dimensional dynamic image of the ultrasonic cardiogram of the detection object into a convolutional neural network model, and automatically recognize The ultrasonic image features of each segment are located separately, the end-diastolic image frames of each standard section are captured, the mitral valve, tricuspid valve, aortic valve, left and right atrioventricular cavities are marked, and the thickness of each atrioventricular cavity and ventricular wall is measured and automatically marked. | 2024 | Zhengzhou University | University | NA | https://www.zzu.edu.cn | China | The patent focuses on a cardiac-ultrasound intelligent recognition system for myocardial ablation, with modules to acquire echocardiograms, build/train/run a neural network, and process images. It addresses the gap that operators rely on manual interpretation of echocardiograms to set ablation parameters during myocardial ablation. AI-recognized cardiac structures help set ablation parameters more accurately. | הפטנט מתמקד במערכת זיהוי חכמה של אולטרסאונד לבבי לאבלציה של שריר הלב, עם מודולים לרכישת אקוקרדיוגרמות, בנייה/אימון/הרצת רשת עצבית ועיבוד תמונות. הוא מתמודד עם הפער שמפעילים מסתמכים על פרשנות ידנית של אקוקרדיוגרמות להגדרת פרמטרי אבלציה במהלך אבלציה של שריר הלב. מבני לב מזוהים על ידי AI עוזרים לקבוע פרמטרי אבלציה מדויקים יותר. |  |  | 1 |  | 1 | [13 empty] | 1 |  |  |  | 1 |  | 1 | [7 empty] | FIRST AFFILIATED HOSPITAL ZHENGZHOU UNIV | Zhengzhou University | FIRST AFFILIATED HOSPITAL ZHENGZHOU UNIV | Zhengzhou University | Zhengzhou University | 2024-09-03 | 2024 | 2024-12-13 | 2024 | 2024-09-03 | 2024 | Application | Utility Invention | active | CN202411226495 | 317380352 | [NA] | CN
386 | CN121129419A | Radio frequency ablation electrode control method | The invention discloses a radio frequency ablation electrode control method which comprises the following steps: S1, initializing an ablation temperature target value, a safety limit value and a power adjusting range, and loading a pre-stored anal fistula ablation three-stage feature library and impedance-temperature correlation degree thresholds of all stages; s2, collecting the electrode temperature, the electrode-tissue impedance and the ablation duration in real time, and judging the current ablation stage according to the matching relationship between the ablation duration and the three-stage feature library; s3, formulating a state judgment rule according to the current stage; s4, a fuzzy PID algorithm is adopted to calculate a power basic adjusting amount, a power dynamic compensation amount is calculated, and final output power is determined; and S5, repeating the steps S2-S4 to realize closed-loop control, triggering first-level early warning when the parameters are close to the boundary, triggering second-level alarm when the parameters exceed the safety limit value, and recording abnormal data. The stage adaptability is high, the problem of temperature adjustment lag can be solved, the adjustment precision is accurate enough, and the ablation effect can be improved. | 1 A method for controlling a radiofrequency ablation electrode, characterized by comprising the following steps: S1. Initialize the target ablation temperature value, safety limit value, and power adjustment range; load the pre-stored anal ablation three-stage feature library and impedance-temperature correlation thresholds for each stage. S2. Real-time acquisition of electrode temperature, electrode-tissue impedance and ablation time; determination of the current ablation stage based on the matching relationship between ablation time and the three-stage feature library. S3. Develop state determination rules according to the current stage: determine the impedance range and temperature response matching during the initial contact period, determine the impedance-temperature correlation and parameter fluctuation during the stable period, predict the impedance rise rate and temperature rise risk during the drying period, and determine the energy output state. S4. Use the fuzzy PID algorithm to calculate the basic power adjustment amount and the dynamic power compensation amount to determine the final output power. S5. Repeat S2-S4 to achieve closed-loop control. When the parameter approaches the boundary, a first-level warning is triggered. When the parameter exceeds the safety limit, a second-level alarm is triggered, and abnormal data is recorded. 2 According to claim 1, a radiofrequency ablation electrode control method is characterized in that: in S1, the three-stage feature library for anal ablation includes: initial contact period: duration 0-5s, impedance decrease rate > 30 Ω/s, temperature rise rate v > 5°C/s; stable ablation period: duration 5s to 10s before the end of the procedure, impedance fluctuation amplitude v > 20 Ω/s, temperature fluctuation amplitude v > 0.5°C/s; tissue drying period: duration greater than 10s before the end of the procedure, impedance rise rate > 10 Ω/s, temperature rise rate > 0.2°C/s; the impedance-temperature correlation threshold for the stable ablation period is set to 0.6 and 0.8, and the impedance-temperature correlation threshold for the tissue drying period is set to 0.4-0.6, and the correlation is calculated using the Pearson correlation coefficient. 3 According to claim 2, a radiofrequency ablation electrode control method is characterized in that: in S2, the electrode temperature data is acquired in real time as temperature T by a thermocouple temperature sensor attached to the tip of the electrode; Electrode-tissue impedance data is acquired in real-time using a 0.01Ω high-precision sampling resistor connected in series with a 24-bit ADC chip. The real-time impedance Z is calculated as: U<sub>sample</sub> = U<sub>sample</sub> × R<sub>sample</sub>, where U<sub>sample</sub> is the voltage across the sampling resistor acquired by the ADC chip, R<sub>sample</sub> = 0.01Ω, which is the high-precision sampling resistor of the RF circuit, and Ur is the total voltage of the RF circuit. Duration data is recorded by a built-in timer in the sampling system to record the ablation duration. 4 According to claim 3, the radiofrequency ablation electrode control method is characterized in that: in S3, the judgment logic for each stage is as follows: Initial contact period: The core judgment is whether the impedance change trend matches the temperature response. If the impedance has dropped to the range of 100Ω~2000Ω but the temperature is still >64°C, it is judged as "insufficient energy"; if the impedance has not dropped to the target range but the temperature is >66°C, it is judged as "local overheating"; if both are within the initial adaptation range, with impedance of 100Ω~2000Ω and temperature of 64°C~66°C, it is judged as "normal"; Stable ablation period: The focus is on monitoring the impedance-temperature correlation and parameter fluctuation. If the correlation is in the range of 0.6~0.8 and the real-time temperature T is in the range of 64.5°C~65.5°C and the real-time impedance Z is in the range of 100Ω~1000Ω. If the correlation coefficient v is greater than 0.6 and the real-time temperature Tv is 64.5°C, it is judged as "normal"; if the correlation coefficient v is greater than 0.8 and the real-time temperature Tv is 65.5°C, it is judged as "energy excess"; During the tissue drying period: Prioritize the risk of a sudden temperature rise; if the real-time impedance rise rate is > 10 Q/s or the real-time temperature is above 7... If the rise rate is >0.5°C/s, it is judged as "predicted overheating"; if the real-time impedance Z rises but the real-time temperature TV is 64°C, it is judged as "insufficient energy"; if the real-time temperature T is between 64.5°C and 65.5°C and the real-time impedance Z rise rate is W80Q/s, it is judged as "normal". 5 According to claim 3, the radiofrequency ablation electrode control method is characterized in that: in S4, the fuzzy PID algorithm is used to calculate the power base adjustment specifically as follows: Input variable quantization: the actual temperature deviation AT = TT target and the impedance deviation AZ = ZZ stage target are converted into fuzzy universes of discourse, and the quantization formulas are: X = AT/ATmax * n, y = AZ/AZmax xn, where ATmax=5°C, Zmax = 100Q /=3, the calculation result is rounded to the nearest integer; Fuzzy set partitioning: AZ and the power base adjustment quantity API are both divided into 7 fuzzy sets: "negative large, negative medium, negative small, zero, positive small, positive medium, and positive large". The membership function of each fuzzy set adopts a triangular distribution; Fuzzy inference: The smaller value method is used to calculate the output membership degree of each rule. The rule base contains 49 core rules. The formula for calculating the output membership degree of a single rule is: Rule i(APl) = min[gAT(Ai),UAZ(Bi)]; where, i and Bi are the fuzzy sets of AT and AZ, respectively; Defuzzification: The output membership degrees of all rules are combined to calculate the precise value of P1. The formula is:Lri - £Lu1 μruleQPl) X PL d = :ΣChangμruleQP1); where Pli is the center value of the API corresponding to the i-th rule. 6 According to claim 5, a radiofrequency ablation electrode control method is characterized in that: power dynamic compensation is calculated: the compensation amount AP2 is used to offset the effects of electrode wear and tissue thermal conductivity differences, and the calculation formula is: AP2 = Ktxt + KWithx ¦ΔΖ¦; Where Kt=0.02W/s is the time compensation coefficient, Kz=0.01W/Q is the impedance deviation compensation coefficient, and t is the ablation time in seconds. 7 According to claim 6, a radiofrequency ablation electrode control method is characterized in that: determining the final output power specifically involves: Preliminary calculation: P<sub>preliminary</sub> = P<sub>0</sub> + API + AP<sub>2</sub>, where P<sub>0</sub> is the current output power; Single adjustment amplitude constraint: the upper limit of single power adjustment P<sub>upper limit</sub> = P<sub>0</sub> x 20%, and the lower limit p<sub>lower limit</sub> = -P<sub>0</sub> x 20%; if API + Δp<sub>2</sub> > p<sub>upper limit</sub>, then p<sub>constraint</sub> = P<sub>0</sub> + P<sub>upper limit</sub>; if ΔP<sub>1</sub> + ΔP<sub>2</sub> < P<sub>upper limit</sub>, then P<sub>constraint</sub> = P<sub>0</sub> + P<sub>lower limit</sub>; otherwise, P<sub>constraint</sub> = P<sub>preliminary</sub>; Stage power upper limit constraint: the maximum allowable power for each stage is 20W for the initial contact period, 25W for the stabilization period, and 15W for the drying period. If P<sub>constraint</sub> > P<sub>stage max</sub>, then the final output power P = P<sub>stage max</sub>; if P<sub>constraint</sub> < 5W, then the final output power P = 5W; otherwise, the final output power P = P<sub>constraint</sub>. 8 According to claim 5, a radiofrequency ablation electrode control method is characterized in that: in S5, the first-level warning is triggered when the real-time temperature T approaches the stage temperature boundary or the real-time impedance Z approaches the stage impedance boundary, prompting the doctor to pay attention; the second-level alarm is triggered immediately when the real-time temperature T exceeds the safety limit or the real-time impedance Z exceeds the safety limit, and the power is automatically reduced to the safety value, and the real-time temperature T, real-time impedance Z, and final output power P data at the abnormal moment are recorded for postoperative analysis. | 2025 | Jilin University | University | NA | https://www.jlu.edu.cn | China | The patent focuses on an RF ablation electrode controller specialized for anal-fistula ablation: it loads a three-stage feature library and impedance–temperature thresholds, then monitors temperature, impedance, and time live. It addresses the dual risk in anal-fistula RF ablation that low temperature leaves residual disease and high temperature burns healthy tissue. The system advances through three stages based on measured signals to keep ablation in the right window. | הפטנט מתמקד בבקר אלקטרודת אבלציית RF המתמחה באבלציית פיסטולה אנאלית: הוא טוען ספריית תכונות בשלוש שלבים וספי אימפדנס-טמפרטורה, ואז מנטר טמפרטורה, אימפדנס וזמן בזמן אמת. הוא מתמודד עם הסיכון הכפול באבלציית RF של פיסטולה אנאלית שטמפרטורה נמוכה משאירה מחלה שיורית וטמפרטורה גבוהה שורפת רקמה בריאה. המערכת מתקדמת דרך שלושה שלבים על בסיס אותות נמדדים כדי לשמור על האבלציה בחלון הנכון. | 1 |  |  |  | 1 | [5 empty] | 1 | [10 empty] | 1 | [5 empty] | 1 |  |  |  |  | FIRST HOSPITAL JILIN UNIV | Jilin University | FIRST HOSPITAL JILIN UNIV | Jilin University | Jilin University | 2025-11-17 | 2025 | 2025-12-16 | 2025 | 2025-11-17 | 2025 | Application | Utility Invention | active | CN202511677524 | 325136967 | [NA] | CN
387 | WO2025046567A1 | A SYSTEM AND METHOD FOR ELECTROPORATION | A system and method are disclosed to position a 2D array active elements on a 2D array of transperineal stylets in a tissue containing a tumor under guidance by imaging technologies, for evaluating, monitoring, and treating tumors by electroporation. | 1. CLAIMSWhat is claimed is: 1. A system for insertion of an array of stylets into tissue of a subject, comprising: a cartridge housing a plurality of stylets each of said stylets including a pointed distal end and a proximal portion connected to the cartridge; and a first template with a plurality of channels through which a distal portion of the stylets pass into the tissue. 2. The system of claim 1, further comprising an actuator configured to move said stylets longitudinally with respect to said first template thereby moving said distal portion of said plurality of stylets longitudinally through said channels of said first template. 3. The system of claim 2, wherein the actuator is further configured to move said stylets longitudinally with respect to said cartridge. 4. The system according to claim 2, further comprising an x-y actuator for positioning a pusher against a selected stylet of said plurality of stylets and wherein said actuator pushes the selected stylet to a desired depth. 5. The system according to claim 4, further wherein at least one of said x-y actuator and a second actuator is configured for detaching the pusher from the selected stylet. 6. The system according to claim 1, further comprising a second template, said first template in contact with the subject and the second template positioned between the first template and the cartridge, wherein the second template is connected to at least a portion of said stylets and wherein said second template is moved longitudinally by an actuator between said first template thereby moving said portion of the stylets longitudinally. 7. The system of claim 6 wherein said second template is configured for retracting said stylets for a predetermined distance. 8. The system according to claim 7, further comprising a processor running an algorithm configured for said retracting and for positioning active electrodes to a next more superficial depth of a procedure. 9. The system according to claim 7, wherein said retracting is for a distance between 5 mm to 30 mm. 10. The system according to claim 1, wherein each stylet is configured to be extended and retracted individually through the first template. 11. The system according to claim 1, comprising an operator interface providing a real time image of the tissue, wherein the operator interface is configured for the operator marking on the real time image a region of a procedure. 12. The system of claim 11, configured to perform a procedure in two disconnected regions simultaneously. 13. The system according to claim 11, wherein the marking is performed at various depths. 14. The system according to claim 13, further comprising a processor running an algorithm to activate to activate only stylets of said plurality of stylets in the region of a procedure at each of said various depths. 15. The system according to claim 11, wherein an electronic grid is superimposed on the real time image and wherein of the grid, coincide with expected positions of the stylets after being pushed into tissue. 16. The system according to claim 11, further comprising a processor running an algorithm to activate a remote actuator that will push each stylet of the plurality of stylets into the tissue at a desired depth and within boundaries of the region of a procedure. 17. The system according to claim 1, further comprising a processor running an algorithm to pass current from a first subset of said stylets to a second subset of said stylets that are intermingled with each other. 18. The system according to claim 17, wherein the first subset of said stylets are dispersed at equal distance from stylets of the second subset. 19. The system according to claim 17, wherein the first subset of said stylets and the second subset of said stylets are distributed and intercalated according to a pattern chosen from parallel rows, concentric pattern and other such equidistant patterns. 20. The system according to claim 19, wherein the pattern is predetermined to facilitate uniformity of current distribution within an area of treatment. 21. The system according to claim 19, wherein the first subset of said stylets and the second subset of said stylets are intercalated along a long axis of a shape to be ablated. 22. The system according to claim 1, further comprising a bioimpedance module that is connected to the stylets fortesting impedance between electrodes on the stylets. 23. The system according to claim 22, wherein the impedance testing is for differentiating between different types of tissue. 24. The system of claim 23, wherein said different types of tissue includes benign and malignant tissue. 25. The system according to claim 22, wherein the impedance testing is for evaluation of a performance and effect of a treatment treated tissue. 26. The system according to claim 1, further comprising an imaging modality and wherein the imaging modality and the system are configured for independent positioning. 27. The system according to claim 1, further comprising an imaging modality and wherein the imaging modality and the system are configured for synchronized positioning. 28. The system according to claim 14, further comprising an imaging modality and wherein an actuator is configured to facilitate rotation of the imaging modality for imaging stylet positioning. 29. The system according to claim 28, wherein an angle of rotation is determined by a predetermined position of the stylet in the cartridge relative to the imaging modality. 30. The system according to claim 29, wherein the imaging modality and the system are configured to be held in pre -determined relative positions. 31. The system of claim 1, wherein the at least one of the cartridge and the stylet includes an individual actuator for longitudinal movement of an individual stylet. 32. The system according to claim 1, wherein each stylet includes a lumen and wherein the lumen is configured to permit insertion of active electrodes. 33. The system according to claim 1, wherein each stylet is at least partially covered by a covering including at least one of a sheath and a coating. 34. The system according to claim 33, wherein the covering is at least one of hydrophobic, hydrophilic, electrically insulating, thermally insulating, and friction reducing. 35. The system according to claim 1, wherein each of said stylets includes an active electrode of /i cm to 3 cm length. 36. The system according to claim 33, wherein the sheath is removable or retractable. 37. The system according to claim 1, wherein each of said stylets is provided with at least two sensing electrodes for bioimpedance measurements. 38. The system according to claim 1, wherein further comprising activatable visual markers on the cartridge. 39. The system according to claim 38, where the activable markers are include one or more of LEDs, projected bright dots, and bright project lines. 40. The system according to claim 39, wherein the activatable markers are produced by a projector situated behind the cartridge. 41. The system according to claim 39, further comprising a processor running an algorithm for activating the activatable markings corresponding to a proximal end of the stylets that were chosen to be inserted into tissue for easy identification. 42. The system of claim 1, wherein the stylets are configured for a procedure selected from the group consisting of: irreversible electroporation (IRE), reversible electroporation (RE), instillation of chemotherapeutics, calcium solution or gene therapy, radiofrequency ablation, microwave ablation, laser ablation, photodynamic therapy, bioimpedance measurement and a combination thereof. 43. The system according to claim 1, wherein the template is configured to provide a variety of densities and patterns of the plurality of stylets. 44. A method for electro-stimulation of tissue of a subject, the method comprising: attaching a cartridge with an array of stylets to a cradle of an imaging modality, wherein the cartridge is configured for parallel insertion of stylets into the tissue; introducing the stylets into the tissue to a desired depth, under monitoring by the imaging modality; selecting a desired area; executing a procedure on the selected area; retracting the stylets to a more superficial transverse plane and repeating the previous steps, as needed. 45. The method according to claim 44, wherein the retracting is accomplished manually, automatically, robotically, remotely or any combination thereof. 46. The method according to claim 45, wherein the retracting is in a longitudinal direction and wherein retracting stops periodically for a procedure on a new layer in the transverse and lateral directions. 47. The method according to claim 44, wherein the area is selected manually or automatically. 48. The method according to claim 44, wherein the area is pre-designated. 49. The method according to claim 44, wherein the procedure is selected from a group including: contact radiofrequency energy, non-contact radiofrequency energy, electroporation, ultrasonic energy, laser energy, gamma radiation, beta radiation, alpha radiation, immunotherapy, or a combination thereof. 50. The method according to claim 44, wherein multiple stylets are configured for acting simultaneously. | 2023 | FOCAL MEDICAL TECH | Medical and Healthcare | Start Up | https://www.focalmd.com/ | Israel | The patent focuses on a 2D-array electroporation system using transperineal stylets carrying active elements, guided by imaging to position them in tumor-bearing tissue. It addresses the limitation of single-needle electroporation, which can't shape the field across a tumor volume. The 2D array allows mapping, monitoring, and electroporation treatment of solid tumors with greater spatial control. | הפטנט מתמקד במערכת אלקטרופורציה במערך דו-ממדי המשתמשת בסטיילטים טרנס-פרינאליים הנושאים אלמנטים פעילים, מונחים על ידי הדמיה למיקום ברקמה נושאת גידולים. הוא מתמודד עם המגבלה של אלקטרופורציה במחט יחידה, שאינה יכולה לעצב את השדה על פני נפח גידול. המערך הדו-ממדי מאפשר מיפוי, ניטור וטיפול אלקטרופורציה של גידולים מוצקים עם בקרה מרחבית רבה יותר. | 1 | [10 empty] | 1 | [16 empty] | 1 |  |  |  | FOCAL MEDICAL TECH LTD | FOCAL MEDICAL TECH LTD | FOCAL MEDICAL TECH LTD | FOCAL MEDICAL TECH LTD | FOCAL MEDICAL TECH LTD | 2024-08-21 | 2024 | 2025-03-06 | 2025 | 2023-08-29 | 2023 | Application | Utility Invention | inactive | WO2024IL50840 | 319651356 | CN121712553A | WO
388 | WO2022093848A1 | SYSTEMS AND METHODS FOR IMAGE GUIDED TISSUE ABLATION | A real-time, intra-uterine operating field image is displayed to an operator. A processor identifies anatomical features and marks anatomical boundaries on the image with the aid of a trained classifier model. Labels for the identified anatomical features and the anatomical boundaries are displayed on the image. The trained classifier model is trained based on a set of prior surgical field images that have anatomical features and anatomical boundaries therein identified and marked, respectively, by previous operators. The trained classifier model is typically a convolutional deep neural network. The operator is allowed to modify the anatomical feature labels and anatomical boundaries, as well as update the set of training images with the modified image. The labeling of the anatomical features and establishment of the anatomical features facilitate performing treatments in the uterus, such as the ablation of uterine fibroids. | 1. CLAIMSWHAT IS CLAIMED IS: 1. A computer-implemented method of planning and implementing a treatment procedure, the method comprising: (i) displaying a real-time operating field image to an operator; (ii) identifying one or more of (a) at least one anatomical feature or (b) at least one anatomical boundary on the real-time operating field image, with the aid of the trained classifier model; (iii) displaying one or more of (a) at least one label for the identified at least one anatomical feature or (b) at least one marking for the identified at least one anatomical boundary on the displayed real-time operating field image; (iv) allowing the operator to modify the real-time operating field image by one or more of (a) re-labelling the identified at least one anatomical feature or (b) re-positioning the at least one marking for the at least one anatomical boundary; and(v) provide an update to the trained classifier model based on the modifications by the operator. 2. The method of claim 1, wherein the trained classifier model is generated by: (vi) displaying an operating field image to a user; (vii) establishing at least one anatomical boundary in the operating field image, the at least one anatomical boundary being marked by the user; (viii) labelling at least one anatomical feature separated by the at least one anatomical boundary in the operating field image, the at least one anatomical feature being labeled by user; and(ix) repeating steps (vi) to (viii) to generate a training set of operating field images each with at least one marked anatomical boundary and at least one labeled anatomical feature; and(x) training a classifier model, based on the training set of operating field images, to identify (a) at least one anatomical feature or (b) at least one anatomical boundary on the realtime operating image. 3. The method of claim 2, wherein the trained classifier model is updated at step (v) by adding the modified real-time operating field image from step (iv) to the training set of operating field images and updating the classifier model based on the set of operating field images with the added modified real-time operating field image. -40-RECTIFIED SHEET (RULE 91) 4. The method of claim 1, further comprising confirming with the operator the one or more of (a) the re-labelled identified at least one anatomical feature or (b) the re-positioned at least one marking for the at least one anatomical boundary before providing the update to the trained classifier model. 5. The method of claim 4, wherein providing the update to the trained classifier model comprises adding the modified real-time operating field image to a training set of operating field images. 6. The method of claim 1, wherein the operating field image at step (i) is a surgical field image. 7. The method of claim 1, further comprising treating an anatomical structure based on one or more of (a) the displayed at least one label for the at least one identified anatomical feature or (b) the displayed at least one marking for the at least one anatomical boundary on the real-time operating field image. 8. The method of claim 7, wherein treating the anatomical structure comprises ablating the anatomical structure. 9. The method of claim 8, wherein the anatomical structure is ablated using one or more of radiofrequency (RF) energy, plasma energy, therapeutic ultrasound energy, microwave energy, heat, or cold. 10. The method of claim 7, wherein the treated anatomical structure is a uterine fibroid. 11. The method of claim 7, further comprising repeating steps (ii) to (iv) in real-time during the treatment procedure. 12. The method of claim 7, further comprising repeating one or more of steps (ii) or (iii) after the treatment procedure to identify one or more changes to the anatomical structure from the treatment procedure. 13. The method of claim 1, wherein the operating field image at step (i) comprises an ultrasound image. 14. The method of claim 13, wherein the ultrasound image is an intra-uterine ultrasound image. 15. The method of claim 1, wherein the operating field image at step (i) comprises an image of a uterus of a patient. -41-RECTIFIED SHEET (RULE 91) 16 The method of claim 15, wherein at least one marking for the at least one anatomical boundary separates a plurality of anatomical features of the uterus, the plurality of anatomical features including one or more of a serosa, a myometrium, a fibroid, a uterine wall, a bladder wall, or a bladder. 17. The method of claim 15, wherein the at least one label for the at least one anatomical feature includes one or more of a serosa, a myometrium, a fibroid, a uterine wall, a bladder wall, or a bladder. 18. The method of claim 1, wherein the trained classifier model comprises a machine learning algorithm. 19. The method of claim 18, wherein the machine learning algorithm is a convolutional deep learning network. 20. The method of claim 1, wherein the trained classifier model is one or more of a binary tree, a random forest, a decision tree, a plurality of decision trees, a plurality of decision trees with controlled variance, a multinomial logistic regression, a naive Bayes classifier, a linear classifier, an ensemble of linear classifiers, a boosting algorithm, a boosting algorithm trained with stochastic gradient descent, a boosting algorithm trained with stochastic gradient descent, a boosting algorithm comprising training data weighting, a boosting algorithm comprising updating training data weighting, or a boosting algorithm comprising updating misclassified training data with higher weights. 21. The method of claim 1, further comprising providing a pre-treatment image to the operator as a guide for identifying the one or more of (a) the at least one anatomical feature or (b) the at least one anatomical boundary on the real-time operating field image. 22. The method of claim 21, wherein the pre-treatment image comprises one or more of an MRI image, an X-ray image, a CT image, or an ultrasound image. 23. An apparatus for aiding in planning and implementing a treatment procedure, the apparatus comprising: a display; an interface; a processor; and a non-transitory computer readable storage medium including instructions configured to cause the processor to: (i) cause the display to display a real-time operating field image to an operator,-42-RECTIFIED SHEET (RULE 91) (ii) identify one or more of (a) at least one anatomical feature or (b) at least one anatomical boundary on the real-time operating field image, with the aid of the trained classifier model,(iii) display one or more of (a) at least one label for the identified at least one anatomical feature or (b) at least one marking for the identified at least one anatomical boundary on the displayed real-time operating field image,(iv) allow the operator to modify the real-time operating field image by one or more of (a) re-labelling the at least one identified anatomical features or (b) re-positioning the at least one marking for the at least one identified anatomical boundary, and(v) provide an update to the trained classifier model based on the modifications by the operator. 24. The apparatus of claim 23, wherein the instructions are further configured to cause the processor to: (vi) display an operating field image to a user; (vii) establish at least one anatomical boundary in the operating field image, the at least one anatomical boundary being marked by the user via the interface; (viii) label at least one anatomical feature separated by the at least one anatomical boundary, the at least one anatomical feature being labeled by the user via the interface; (ix) repeat steps (vi) to (viii) to generate a training set of operating field images each with at least one marked anatomical boundary and at least one labeled anatomical feature; and(x) train a classifier model, based on the training set of operating field images, to at least one of (a) identify at least one anatomical feature or (b) establish at least one anatomical boundary on the real-time operating image. 25. The apparatus of claim 24, wherein the trained classifier model is updated at step (v) by adding the modified real-time operating field image from step (iv) to the training set of operating field images and updating the classifier model based on the set of operating field images with the added modified real-time operating field image. 26. The apparatus of claim 23, wherein the instructions are further configured to cause the processor to confirm with the operator the one or more of (a) the re-labelled at least one anatomical feature or (b) the re-positioned at least one marking for the at least one anatomical boundary before providing the update to the trained classifier model. 27. The apparatus of claim 26, wherein the update is provided to the trained classifier model by adding the modified real-time operating field image to a training set of operating field images. -43-RECTIFIED SHEET (RULE 91) 28. The apparatus of claim 23, wherein the operating field image at step (i) is an operating field image. 29. The apparatus of claim 23, wherein the instructions are further configured to cause the processor to further allow the operator to use the interface to treat an anatomical structure based on one or more of (a) the displayed at least one label for the at least one identified anatomical feature or (b) the displayed at least one marking for the at least one anatomical boundary on the real-time operating field image. 30. The apparatus of claim 29, wherein treating the anatomical structure comprises ablating the anatomical structure with an ablation element, and wherein the apparatus further comprises the ablation element. 31. The apparatus of claim 30, wherein the ablation element is configured to delivery one or more of radiofrequency (RF) energy, plasma energy, therapeutic ultrasound energy, microwave energy, heat, or cold to the anatomical structure. 32. The apparatus of claim 29, wherein the treated anatomical structure is a uterine fibroid. 33. The apparatus of claim 29, wherein the instructions are further configured to cause the processor to repeat steps (ii) to (iv) in real-time during the treatment procedure. 34. The apparatus of claim 29, wherein the instructions are further configured to cause the processor to repeat one or more of steps (ii) or (iii) after the treatment procedure to identify one or more changes to the anatomical structure from the treatment procedure. 35. The apparatus of claim 23, wherein the operating field image at step (i) comprises an ultrasound image. 36. The apparatus of claim 35, wherein the ultrasound image is an intra-uterine ultrasound image. 37. The apparatus of claim 23, wherein the operating field image at step (i) comprises an image of a uterus of a patient. 38. The apparatus of claim 37, wherein the at least one marking for the at least one anatomical boundary separates a plurality of anatomical features of the uterus, the plurality of anatomical features including one or more of a serosa, a myometrium, a fibroid, a uterine wall, a bladder wall, or a bladder. -44-RECTIFIED SHEET (RULE 91) 39. The apparatus of claim 37, wherein the at least one label for the at least one anatomical feature includes one or more of a serosa, a myometrium, a fibroid, a uterine wall, a bladder wall, or a bladder. 40. The apparatus of claim 23, wherein the classifier model comprises a machine learning algorithm. 41. The apparatus of claim 40, wherein the machine learning algorithm is a convolutional deep learning network. 42. The apparatus of claim 23, wherein the classifier model is one or more of a binary tree, a random forest, a decision tree, a plurality of decision trees, a plurality of decision trees with controlled variance, a multinomial logistic regression, a naive Bayes classifier, a linear classifier, an ensemble of linear classifiers, a boosting algorithm, a boosting algorithm trained with stochastic gradient descent, a boosting algorithm trained with stochastic gradient descent, a boosting algorithm comprising training data weighting, a boosting algorithm comprising updating training data weighting, or a boosting algorithm comprising updating misclassified training data with higher weights. 43. The apparatus of claim 23, wherein the instructions are further configured to cause the processor to further provide a pre-treatment image to the operator as a guide for identifying the one or more of (a) the at least one anatomical feature or (b) the at least one anatomical boundary on the real-time operating field image. 44. The apparatus of claim 43, wherein the pre-treatment image comprises one or more of an MRI image, an X-ray image, a CT image, or an ultrasound image. 45. A computer-implemented method of planning and implementing a treatment procedure, the method comprising: (i) displaying a real-time operating field image to an operator; (ii) identifying one or more of (a) at least one anatomical feature or (b) at least one anatomical boundary on the real-time operating field image, with the aid of a trained classifier model; and(iii) displaying one or more of (a) at least one label for the identified at least one anatomical feature or (b) at least one marking for the at least one anatomical boundary on the real-time operating field image, wherein the trained classifier model is trained based on a set of operating field images each with at least one marked anatomical boundary and at least one labeled anatomical feature,-45-RECTIFIED SHEET (RULE 91) wherein the at least one marked anatomical boundary is identified by a previous operator, and wherein the at least one labeled anatomical feature is identified by the previous operator. 46. The method of claim 45, further comprising allowing the operator to modify the real-time operating field image by one or more of (a) re-labelling the identified at least one anatomical feature or (b) re-positioning the identified at least one anatomical boundary. 47. The method of claim 46, further comprising updating the set of operating field images by adding the modified real-time operating field image to the set. 48. The method of claim 47, further comprising updating the classifier model based on the updated set of operating field images. 49. The method of claim 45, further comprising treating an anatomical structure based on one or more of (a) the displayed at least one label for the identified at least one anatomical feature or (b) the displayed at least one marking for the at least one anatomical boundary on the real-time surgical field image. 50. The method of claim 49, wherein treating the anatomical structure comprises ablating the anatomical structure. 51. The method of claim 49, wherein the anatomical structure is ablated using one or more of radiofrequency (RF) energy, plasma energy, therapeutic ultrasound energy, microwave energy, heat, or cold. 52. The method of claim 49, wherein the treated anatomical structure is a uterine fibroid. 53. The method of claim 49, further comprising repeating steps (ii) and (iii) in realtime during the treatment procedure. 54. The method of claim 49, further comprising repeating one or more of steps (ii) or (iii) after the treatment procedure to identify one or more changes to the anatomical structure from the treatment procedure. 55. The method of claim 45, wherein the real-time operating field image comprises an ultrasound image. 56. The method of claim 55, wherein the ultrasound image is an intra-uterine ultrasound image. 57. The method of claim 45, wherein the identified at least one anatomical boundary separates a plurality of anatomical features of a uterus, the plurality of anatomical features-46-RECTIFIED SHEET (RULE 91) including one or more of a serosa, a myometrium, a fibroid, a uterine wall, a bladder wall, or a bladder. 58. The method of claim 45, wherein the at least one labeled anatomical feature includes one or more of a serosa, a myometrium, a fibroid, a uterine wall, a bladder wall, or a bladder. 59. The method of claim 45, wherein the classifier model comprises a machine learning algorithm. 60. The method of claim 59, wherein the machine learning algorithm is a convolutional deep learning network. 61. The method of claim 59, wherein the classifier model is one or more of a binary tree, a random forest, a decision tree, a plurality of decision trees, a plurality of decision trees with controlled variance, a multinomial logistic regression, a naive Bayes classifier, a linear classifier, an ensemble of linear classifiers, a boosting algorithm, a boosting algorithm trained with stochastic gradient descent, a boosting algorithm trained with stochastic gradient descent, a boosting algorithm comprising training data weighting, a boosting algorithm comprising updating training data weighting, or a boosting algorithm comprising updating misclassified training data with higher weights. 62. The method of claim 45, further providing a pre-treatment image to the operator as a guide for identifying the one or more of (a) the at least one anatomical feature or (b) the at least one anatomical boundary on the real-time operating field image. 63. The method of claim 62, wherein the pre-treatment image comprises one or more of an MRI image, an X-ray image, a CT image, or an ultrasound image. 64. An apparatus for aiding in performing implementing a treatment procedure, the apparatus comprising: a display; an interface; a processor; and a non-transitory computer readable storage medium including instructions configured to cause the processor to: (i) display a real-time operating field image to an operator; (ii) identify one or more of (a) at least one anatomical feature or (b) at least one anatomical boundary on the real-time operating field image, with the aid of a trained classifier model; and-47-RECTIFIED SHEET (RULE 91) (iii) display one or more of (a) at least one label for the identified at least one anatomical feature or (b) at least one marking for the identified at least one anatomical boundary on the real-time operating field image, wherein the trained classifier model is trained based on a set of operating field images each with at least one marked anatomical boundary and at least one labeled anatomical feature, wherein the at least one marked anatomical boundary is identified by a previous operator, and wherein the at least one labeled anatomical feature is identified by the previous operator. 65. The apparatus of claim 64, wherein the instructions are further configured to cause the processor to allow the operator to modify the real-time operating field image by one or more of (a) re-labelling the identified at least one anatomical feature or (b) re-positioning the identified at least one anatomical boundary. 66. The apparatus of claim 64, wherein the instructions are further configured to cause the processor to update the set of operating field images by adding the modified real-time operating field image to the set. 67. The apparatus of claim 66, wherein the instructions are further configured to cause the processor to further update the classifier model based on the updated set of operating field images. 68. The apparatus of claim 64, wherein the instructions are further configured to cause the processor to further allow the operator to use the interface to treat an anatomical structure based on one or more of (a) the displayed at least one label for the identified at least one anatomical feature or (b) the displayed at least one marking for the at least one anatomical boundary on the real-time surgical field image. 69. The apparatus of claim 68, wherein treating the anatomical structure comprises ablating the anatomical structure with an ablation element, and wherein the apparatus further comprises the ablation element. 70. The apparatus of claim 69, wherein the ablation element is configured to delivery one or more of radiofrequency (RF) energy, plasma energy, therapeutic ultrasound energy, microwave energy, heat, or cold to the anatomical structure. 71. The apparatus of claim 68, wherein the treated anatomical structure is a uterine fibroid. 72. The apparatus of claim 68, wherein the instructions are further configured to cause the processor to repeat steps (ii) and (iii) during the treatment procedure. -48-RECTIFIED SHEET (RULE 91) 73. The apparatus of claim 68, wherein the instructions are further configured to cause the processor to repeat one or more of steps (ii) or (iii) after the treatment procedure to identify one or more changes to the anatomical structure from the treatment procedure. 74. The apparatus of claim 64, wherein the real-time operating field image comprises an ultrasound image. 75. The apparatus of claim 73, wherein the ultrasound image is an intra-uterine ultrasound image. 76. The apparatus of claim 64, wherein the real-time operating field image comprises an image of a uterus of a patient. 77. The apparatus of claim 64, wherein the identified at least one anatomical boundary separates a plurality of anatomical features of the uterus, the plurality of anatomical features including one or more of a serosa, a myometrium, a fibroid, a uterine wall, a bladder wall, or a bladder. 78. The apparatus of claim 64, wherein the at least one labeled anatomical feature includes one or more of a serosa, a myometrium, a fibroid, a uterine wall, a bladder wall, or a bladder. 79. The apparatus of claim 64, wherein the classifier model comprises a machine learning algorithm. 80. The apparatus of claim 79, wherein the machine learning algorithm is a convolutional deep learning network. 81. The apparatus of claim 79, wherein the classifier model is one or more of a binary tree, a random forest, a decision tree, a plurality of decision trees, a plurality of decision trees with controlled variance, a multinomial logistic regression, a naive Bayes classifier, a linear classifier, an ensemble of linear classifiers, a boosting algorithm, a boosting algorithm trained with stochastic gradient descent, a boosting algorithm trained with stochastic gradient descent, a boosting algorithm comprising training data weighting, a boosting algorithm comprising updating training data weighting, or a boosting algorithm comprising updating misclassified training data with higher weights. 82. The apparatus of claim 64, wherein the instructions are further configured to cause the processor to further provide a pre-treatment image to the operator as a guide for-49-RECTIFIED SHEET (RULE 91) identifying the one or more of (a) the at least one anatomical feature or (b) the at least one anatomical boundary on the real-time operating field image. 83. The apparatus of claim 82, wherein the pre-treatment image comprises one or more of an MRI image, an X-ray image, a CT image, or an ultrasound image. 84. A computer-implemented method of planning and implementing a treatment procedure, the method comprising: (i) displaying a real-time operating field image to an operator; (ii) identifying one or more of (a) at least one anatomical feature or (b) at least one anatomical boundary on the real-time operating field image, with the aid of the trained classifier model; (iii) displaying one or more of (a) at least one label for the identified at least one anatomical feature or (b) at least one marking for the identified at least one anatomical boundary on the displayed real-time operating field image; and(iv) treating an anatomical structure based on one or more of (a) the displayed at least one label for the identified at least one anatomical feature or (b) the displayed at least one marking for the identified at least one anatomical boundary on the real-time operating field image. 85. The method of claim 84, wherein the real-time operating field image is a real-time surgical field image. 86. The method of claim 84, wherein the real-time operating field image is an ultrasound image. 87. The method of claim 86, wherein the ultrasound image is an intra-uterine ultrasound image. 88. The method of claim 84, wherein the operating field image is an image of a uterus of a patient. 89. The method of claim 88, wherein the identified at least one anatomical boundary separate a plurality of anatomical features of the uterus, the plurality of anatomical features including one or more of a serosa, a myometrium, a fibroid, a uterine wall, a bladder wall, or a bladder. 90. The method of claim 88, wherein the identified at least one anatomical feature includes one or more of a serosa, a myometrium, a fibroid, a uterine wall, a bladder wall, or a bladder. -50-RECTIFIED SHEET (RULE 91) 91. The method of claim 84, wherein treating the anatomical structure comprises ablating the anatomical structure. 92. The method of claim 91, wherein the anatomical structure is ablated using one or more of radiofrequency (RF) energy, plasma energy, therapeutic ultrasound energy, microwave energy, heat, or cold. 93. The method of claim 84, wherein the treated anatomical structure is a uterine fibroid. 94. The method of claim 84, wherein the trained classifier model comprises a machine learning algorithm. 95. The method of claim 84, further comprising: allowing the operator to modify the real-time operating field image by one or more of (a) re-labelling the identified at least one anatomical feature or (b) re-positioning the marked at least one anatomical boundary; and providing an update to the trained classifier model based on the modifications by the operator. 96. The method of claim 84, further comprising repeating steps (ii) and (iii) during the treatment procedure. 97. The method of claim 84, further comprising repeating one or more of steps (ii) or (iii) after the treatment procedure to identify one or more changes to the anatomical structure from the treatment procedure. 98. The method of claim 84, further comprising providing a pre-treatment image to the operator as a guide for identifying the one or more of (a) the at least one anatomical feature or (b) the at least one anatomical boundary on the real-time operating field image. 99. The method of claim 98, wherein the pre-treatment image comprises one or more of an MRI image, an X-ray image, a CT image, or an ultrasound image. 100. An apparatus for aiding in implementing a treatment procedure, the apparatus comprising: a display; an interface; a processor; and a non-transitory computer readable storage medium including instructions configured to cause the processor to: -51-RECTIFIED SHEET (RULE 91) (i) display a real-time operating field image to an operator; (ii) identify one or more of (a) at least one anatomical feature or (b) at least one anatomical boundary on the real-time operating field image, with the aid of the trained classifier model; (iii) display one or more of (a) at least one label for the identified at least one anatomical feature or (b) at least one marking for the identified at least one anatomical boundary on the displayed real-time operating field image; and(iv) treat an anatomical structure based on one or more of (a) the displayed at least one label for the identified at least one anatomical feature or (b) the displayed at least one marking for the identified at least one anatomical boundary on the real-time operating field image. 101. The apparatus of claim 100, wherein the real-time operating field image is a realtime surgical field image. 102. The apparatus of claim 100, wherein the real-time operating field image is an ultrasound image. 103. The apparatus of claim 102, wherein the ultrasound image is an intra-uterine ultrasound image. 104. The apparatus of claim 100, wherein the operating field image is an image of a uterus of a patient. 105. The apparatus of claim 104, wherein the identified at least one anatomical boundary separates a plurality of anatomical features of the uterus, the plurality of anatomical features including one or more of a serosa, a myometrium, a fibroid, a uterine wall, a bladder wall, or a bladder. 106. The apparatus of claim 104, wherein the identified at least one anatomical feature includes one or more of a serosa, a myometrium, a fibroid, a uterine wall, a bladder wall, or a bladder. 107. The apparatus of claim 100, wherein treating the anatomical structure comprises ablating the anatomical structure, and wherein the apparatus further comprises an ablation element. 108. The apparatus of claim 107, wherein the ablation element is configured to ablate the anatomical structure using one or more of radiofrequency (RF) energy, plasma energy, therapeutic ultrasound energy, microwave energy, heat, or cold. -52-RECTIFIED SHEET (RULE 91) 109. The apparatus of claim 100, wherein the treated anatomical structure is a uterine fibroid. 110. The apparatus of claim 100, wherein the trained classifier model comprises a machine learning algorithm. 111. The apparatus of claim 100, wherein the instructions are further configured to cause the processor to: allow the operator to modify the real-time operating field image by one or more of (a) relabelling the identified at least one anatomical feature or (b) re-positioning the marked at least one anatomical boundary; and provide an update to the trained classifier model based on the modifications by the operator. 112. The apparatus of claim 100, wherein the instructions are further configured to cause the processor to repeat steps (ii) and (iii) during the treatment procedure. 113. The apparatus of claim 100, wherein the instructions are further configured to cause the processor to repeat one or more of steps (ii) or (iii) after the treatment procedure to identify one or more changes to the anatomical structure from the treatment procedure. 114. The apparatus of claim 100, wherein the instructions are further configured to cause the processor to further provide a pre-treatment image to the operator as a guide for identifying the one or more of (a) the at least one anatomical feature or (b) the at least one anatomical boundary on the real-time operating field image. 115. The apparatus of claim 114, wherein the pre-treatment image comprises one or more of an MRI image, an X-ray image, a CT image, or an ultrasound image. -53-RECTIFIED SHEET (RULE 91) | 2020 | GYNESONICS | Medical and Healthcare | Small-size company | https://gynesonics.com | USA | The patent focuses on a real-time intra-uterine image-guided ablation tool that uses a trained classifier to identify and label anatomical features and boundaries on the live image. It addresses the difficulty operators face in interpreting intra-uterine views, where visual landmarks can be subtle and inconsistent. Operators see automatic anatomical labels overlaid on the operating field. | הפטנט מתמקד בכלי אבלציה מונחה תמונה תוך-רחמית בזמן אמת המשתמש במסווג מאומן לזיהוי וסימון תכונות וגבולות אנטומיים על התמונה החיה. הוא מתמודד עם הקושי שיש למפעילים בפרשנות תצוגות תוך-רחמיות, שבהן ציוני דרך ויזואליים יכולים להיות עדינים ולא עקביים. מפעילים רואים תוויות אנטומיות אוטומטיות מצופות על שדה הניתוח. |  |  | 1 | [5 empty] | 1 | [13 empty] | 1 | [8 empty] | 1 | GYNESONICS INC | GYNESONICS INC | GYNESONICS INC | GYNESONICS INC | GYNESONICS INC | 2021-10-26 | 2021 | 2022-05-05 | 2022 | 2020-10-27 | 2020 | Application | Utility Invention | inactive | WO2021US56692 | 316062879 | CN116669635A¦EP4238106A1¦EP4238106A4¦US20230260121A1 | WO
389 | US20250380980A1 | SYSTEMS AND METHODS FOR UTERINE FIBROID ABLATION | Various systems, devices, and methods for tissue penetration are disclosed. The system includes an ablation element that penetrates the tissue and a radiofrequency generator that delivers energy to the ablation element. The radiofrequency generator may have two modes: a cutting or insertion mode and an ablation or coagulation mode. The cutting or insertion mode is used to cut or penetrate the tissue by providing a voltage or power modulation that acts to cut or soften tissue contacted by the ablation element. The ablation or coagulation mode is used to ablate or coagulate the tissue. | 1. A system for penetrating target tissue, the system comprising: an ablation instrument; and a controller configured to control the delivery of energy to the ablation instrument, wherein the controller is configured to monitor a temperature measured by the ablation instrument at an interface with the target tissue and maintain the temperature between about 80° C. and about 115° C. as the ablation instrument penetrates into the target tissue. 2. The system of claim 1, wherein the ablation instrument comprises an introducer and a plurality of electrode ablation needles extendible from a retracted position within the introducer. 3. The system of claim 2, wherein the temperature measured by the ablation instrument is measured by a thermocouple positioned on one of the electrode ablation needles when the plurality of electrode ablation needles is retracted within the introducer. 4. The system of claim 3, wherein the plurality of electrode ablation needles comprises a central electrode extendible from a retracted position within the introducer, and wherein the thermocouple is positioned on the central electrode. 1. ystem of claim 1, wherein the controller is configured to modulate power delivered to the ablation instrument to maintain the temperature between about 80° C. and about 115° C. as the ablation instrument penetrates into the target tissue. 6. The system of claim 1, wherein the controller is configured to modulate power delivered to the ablation instrument as the ablation instrument penetrates into the target tissue to provide an oscillation in the temperature at the interface with the target tissue. 7. The system of claim 1, wherein the controller is configured to provide an alert when the temperature at the interface is between about 80° C. and about 115° C. to provide an indication to begin penetration by the ablation instrument into the target tissue. 8. The system of claim 1, wherein the system further comprises an ultrasonic imaging device configured to provide for visualization of the ablation instrument as the ablation instrument penetrates the target tissue. 9. The system of claim 8, wherein the ultrasonic imaging device is coupled to the ablation instrument. 10. The system of claim 1, wherein the controller is configured to control the delivery of energy to the ablation instrument to ablate the target tissue after the ablation instrument has penetrated into the target tissue. 11. The system of claim 10, wherein the controller is configured to control the delivery of energy to the ablation instrument to maintain a substantially constant temperature at the target tissue to ablate the target tissue. 12. The system of claim 1, further comprising a radiofrequency generator configured to deliver energy to the ablation instrument while the ablation instrument penetrates into the target tissue. 13. A method of penetrating target tissue, the method comprising: delivering an ablation instrument into contact with a target tissue; delivering energy to the ablation instrument to cause a temperature measured by the ablation instrument at an interface with the target tissue to rise to between about 80° C. and about 115° C.; and advancing the ablation instrument to penetrate into the target tissue while modulating the power to maintain the temperature at the interface with the target tissue between about 80° C. and about 115° C. 14. The method of claim 13, wherein the ablation instrument penetrates into the target tissue while the temperature maintained at the interface with the target tissue oscillates. 15. The method of claim 13, further comprising ablating the target tissue by delivering energy to the ablation instrument to maintain the temperature at the interface with the target tissue at a substantially constant temperature. 16. The method of claim 13, wherein the ablation instrument comprises an introducer and a plurality of electrode ablation needles extendible from a retracted position within the introducer, wherein the plurality of electrode ablation needles is at least partially extended from the introducer while maintaining the temperature at the interface with the target tissue between about 80° C. and about 115° C. 17. A system for penetrating target tissue, the system comprising: an ablation element configured to penetrate the target tissue; and a radiofrequency generator configured to deliver energy to the ablation element, wherein the radiofrequency generator comprises: a cutting or insertion mode configured to cut through target tissue, wherein the cutting or insertion mode provides for a power oscillation configured to cut tissue contacted by the ablation element; and an ablation or coagulation mode configured to ablate and/or coagulate the target tissue, wherein the ablation or coagulation mode provides for an increase in power and then a decrease in power after the target tissue reaches a target temperature. 18. The system of claim 17, wherein the cutting or insertion mode is configured to maintain a substantially constant surface temperature of the ablation element. 19. The system of claim 17, wherein the cutting or insertion mode is configured to cause a rapid increase in temperature in tissue in contact with the ablation element. 20. The system of claim 17, wherein the radiofrequency generator is configured to be controlled to provide a limit on temperature during the cutting or insertion mode and during the ablation or coagulation mode, wherein the limit on temperature during the cutting or insertion mode is lower than the limit on temperature during the ablation or coagulation mode. 21. The system of claim 17, further configured to provide a mechanical vibration or cutting force during the cutting or insertion mode. 22. The system of claim 17, wherein the ablation element comprises an introducer. 23. The system of claim 17, wherein the ablation element comprises a plurality of electrode ablation needles. 24. The system of claim 17, further comprising a controller configured to control the delivery of energy to the ablation element, wherein the controller is configured to monitor a temperature measured by the ablation element at an interface with the target tissue and maintain the temperature between about 80° C. and about 115° C. as the ablation element penetrates into the target tissue. 25. The system of claim 17, wherein the target tissue is a uterine fibroid. 26. A method of uterine fibroid ablation, comprising: delivering an ablation element into contact with a uterine fibroid; delivering radiofrequency energy according to a cutting or insertion mode to the ablation element while the ablation element is in contact with the uterine fibroid, wherein in the cutting or insertion mode a voltage or power is modulated based on a temperature measured at an interface between the ablation element and the uterine fibroid to assist the ablation element in penetrating into the uterine fibroid; and after the ablation element penetrates into the uterine fibroid, delivering radiofrequency energy according to a coagulation mode, wherein in the coagulation mode a voltage or power applied to the ablation element is sufficient to ablate the uterine fibroid. 27. The method of claim 26, wherein the voltage or power is oscillated in the cutting or insertion mode. 28. The method of claim 26, wherein the voltage or power is modulated in the cutting or insertion mode to maintain a temperature at the interface between about 80° C. and about 115° C. 29. The method of claim 26, wherein the voltage or power in the coagulation mode increases and then decreases after tissue in contact with the ablation element reaches a target temperature. 30. The method of claim 26, wherein delivering radiofrequency energy according to the cutting or insertion mode preheats the uterine fibroid. 31. The method of claim 26, wherein delivering radiofrequency energy according to the cutting or insertion mode softens the uterine fibroid thereby allowing the ablation element to penetrate the uterine fibroid without substantially deforming the uterine fibroid upon penetration. 32. The method of claim 26, wherein delivering radiofrequency energy according to the cutting or insertion mode maintains a substantially constant surface temperature of the ablation element. 33. The method of claim 32, wherein the substantially constant surface temperature of the ablation element is between about 80° C. and about 115° C. 34. The method of claim 26, wherein delivering radiofrequency energy according to the cutting or insertion mode is controlled to a maximum output not to exceed 70 watts. 35. The method of claim 26, wherein delivering radiofrequency energy according to the cutting or insertion mode is controlled not to exceed 30 seconds. 36. The method of claim 26, further comprising detecting a force on the ablation element when in contact with the uterine fibroid, and wherein delivering radiofrequency energy according to the cutting or insertion mode is controlled based on detection of the force. | 2022 | GYNESONICS | Medical and Healthcare | Small-size company | https://gynesonics.com | USA | The patent focuses on a uterine-fibroid ablation device whose RF generator has two modes — cutting/insertion (modulated voltage that softens tissue) and ablation/coagulation. It addresses the inefficiency of using separate tools to penetrate and then ablate fibroids. The same probe penetrates the fibroid and then ablates it without switching tools. | הפטנט מתמקד במכשיר אבלציה למיומה רחמית שמחולל ה-RF שלו בעל שני מצבים — חיתוך/החדרה (מתח מודולרי המרכך רקמה) ואבלציה/קרישה. הוא מתמודד עם חוסר היעילות של שימוש בכלים נפרדים לחדירה ואז אבלציה של מיומה. אותו גשוש חודר למיומה ואז מבצע בה אבלציה ללא החלפת כלים. |  |  | 1 |  | 1 | [5 empty] | 1 | [20 empty] | 1 | GYNESONICS INC | GYNESONICS INC | GYNESONICS INC | GYNESONICS INC | GYNESONICS INC | 2023-06-27 | 2023 | 2025-12-18 | 2025 | 2022-06-28 | 2022 | Application | Utility Invention | active | US202318879178 | 324988831 | CN119630358A¦EP4547135A1¦JP2025521812A¦WO2024006789A1 | US
390 | CN116983075A | Temperature monitoring compensation algorithm and system for ablation needle | The invention relates to a temperature monitoring compensation algorithm and system for an ablation needle, and relates to the technical field of ablation. According to the temperature monitoring compensation algorithm of the ablation needle, the temperature of the target area is obtained by calibrating the measured temperature data x and considering the relative temperature change of the cooling medium, and the actual temperature of the target area can be indicated, so that precise surgical treatment of tumors can be realized; and the risk of relapse caused by incomplete ablation inactivation and serious complications caused by excessive ablation can be greatly reduced. | 1. A temperature monitoring and compensation algorithm for an ablation needle, characterized in that it includes the following steps: Obtaining the measured temperature data X of the proximal end of the needle shaft of the ablation needle; Obtaining the temperature data % of the cooling medium in the needle shaft; Calibrate the above-mentioned measured temperature data Calculate the temperature F(Xk) of the target area based on the calibration value f(X) and the relative temperature change value fω of the cooling medium, where, F (ζ, t) =f (ζ) + f (t) ο 2. The temperature monitoring and compensation algorithm of the ablation needle according to claim 1, characterized in that calibrating the measured temperature data X to obtain the calibration value f ω includes the following sub-steps: The standard temperature of the environment is divided into multiple standard temperature intervals; Obtain the starting measurement value and end measurement value corresponding to the starting standard temperature and end standard temperature in any standard temperature interval; According to the starting standard temperature, end standard temperature, The starting measurement value and the end measurement value are used to fit the measured temperature data X to obtain the calibration value f CO. 3. The temperature monitoring and compensation algorithm of the ablation needle according to claim 2, characterized in that the calibration value f (z) satisfies the following relationship: f fx) = - ~ (yi - y.) + y. ; Sun 1 3 Among them, is the starting standard temperature in any standard temperature interval; is the end standard temperature in the standard temperature interval; X. It is the starting measurement value corresponding to the starting standard temperature within the standard temperature range; X1 is the end measurement value corresponding to the end standard temperature within the standard temperature range. 4. The temperature monitoring and compensation algorithm of the ablation needle according to any one of claims 1 to 3, characterized in that when the measured temperature data X is calibrated to obtain the calibration value f ω, the temperature of the same cooling medium is Reference; or when calibrating the measured temperature data X to obtain the calibration value, the temperature of the cooling medium is the initial temperature. 5. The temperature monitoring and compensation algorithm of the ablation needle according to claim 3, characterized in that, obtaining the relative temperature change value f&) of the cooling medium according to the temperature data t of the cooling medium includes the following sub-steps: Obtaining the initial temperature of the cooling medium; Obtaining The first measured value corresponding when the temperature of the cooling medium is the first cooling temperature, and the second measured value corresponding when the temperature of the cooling medium is the second cooling temperature is obtained; For the initial temperature, the first cooling temperature, the second cooling temperature, The first measured value and the second measured value are subjected to a fitting operation to obtain the relative temperature change value of the cooling medium fl). 6. The temperature monitoring and compensation algorithm of the ablation needle according to claim 5, characterized in that the relative temperature change value f of the cooling medium satisfies the following relationship: where is the initial cooling water temperature; J is the first cooling temperature; G is the second cooling temperature; Xi is the first measured value corresponding to the first cooling temperature J; X2 is the second measured value corresponding to the second cooling temperature t2. 7. The temperature monitoring and compensation algorithm of the ablation needle according to claim 5, characterized in that the relative temperature change value f of the cooling medium satisfies the following relationship: f (t) = kx (t - to); wherein 4 is the coefficient. Water=0.2-0.4; Product is the initial cooling water temperature. 8. The temperature monitoring and compensation algorithm of the ablation needle according to any one of claims 1 to 3, characterized in that the needle shaft of the ablation needle is obtained through the first temperature measurement element at the proximal end of the needle shaft of the ablation needle. Proximal measured temperature data The second temperature measuring element in the needle bar obtains the temperature data of the cooling medium in the needle bar, wherein the second temperature measuring element is located on the distal outer surface of the needle bar or in the cooling medium. 9. The temperature monitoring and compensation algorithm of the ablation needle according to claim 6, characterized in that, when the relative temperature change value f(t) of the cooling medium is obtained according to the temperature data t of the cooling medium, the same needle of the ablation needle is used. The standard temperature of the environment near the rod's proximal end is used as a reference. 10. A temperature monitoring and compensation system for an ablation needle, characterized in that it includes an ablation needle and an ablation instrument connected to the ablation needle. The ablation instrument includes: a temperature acquisition module, which is used to obtain the needle temperature of the ablation needle respectively. The measured temperature data t obtains the relative temperature change value f S) of the cooling medium, and calculates the temperature of the target area based on the calibration value f 0O and the relative temperature change value f1) of the cooling medium, where, F(X, t) =f(X)+ f (t) ο 11. The temperature monitoring and compensation system of the ablation needle according to claim 10, wherein the temperature acquisition module includes a first temperature measurement element for acquiring the measured temperature data X of the proximal end of the needle shaft of the ablation needle and A second temperature measuring element used to obtain the temperature data t of the cooling medium in the needle bar. The ablation needle includes a needle bar and a multi-layer isolation structure provided on the needle bar. The multi-layer isolation structure includes an internal Isolation structure and external isolation structure, the internal isolation structure is arranged outside the needle bar, the external isolation structure includes an external isolation film covering the outer wall of the internal isolation structure, wherein the first temperature measurement The element is located on the proximal outer surface of the needle bar or between the internal isolation structure and the outer isolation film; the second temperature measuring element is located on the distal outer surface of the needle bar or is located on the in the cooling medium. | 2023 | HYGEA MEDICAL TECH | Medical and Healthcare | NA | https://www.hygeamed.com/ | China | The patent focuses on an ablation-needle temperature compensation algorithm that calibrates measured temperature using the relative temperature change of the cooling medium. It addresses the inaccuracy of microwave/RF needle temperature readings caused by the cooling medium itself, which biases the sensor. The compensation gives a more accurate read of the actual target tissue temperature. | הפטנט מתמקד באלגוריתם פיצוי טמפרטורה למחט אבלציה המכייל את הטמפרטורה הנמדדת באמצעות שינוי הטמפרטורה היחסי של חומר הקירור. הוא מתמודד עם חוסר הדיוק של קריאות טמפרטורה במחטי מיקרוגלים/RF הנגרם מחומר הקירור עצמו, המטה את החיישן. הפיצוי נותן קריאה מדויקת יותר של טמפרטורת רקמת המטרה האמיתית. |  |  |  |  | 1 | [16 empty] | 1 | [6 empty] | 1 |  |  |  | HYGEA MEDICAL TECH CO LTD | HYGEA MEDICAL TECH CO LTD | HYGEA MEDICAL TECH CO LTD | HYGEA MEDICAL TECH CO LTD | HYGEA MEDICAL TECH CO LTD | 2023-09-28 | 2023 | 2023-11-03 | 2023 | 2023-09-28 | 2023 | Application | Utility Invention | inactive | CN202311267891 | 291945325 | CN116983075B | CN
391 | CN120616749A | Temperature data acquisition and report generation method for cervical cryotherapy | The invention belongs to the technical field of cervical cryotherapy, and discloses a cervical cryotherapy temperature data acquisition and report generation method, which comprises the following steps: capturing a temperature gradient based on a multi-point thermosensitive probe array, and generating temperature distribution original data; performing thermal diffusion path analysis to generate freezing front migration characteristic data; positioning a thermal diffusion critical point, and generating freezing depth critical value data; analyzing the interface temperature field to generate a three-dimensional distribution characteristic spectrum; representing an ice crystal formation dynamic state, and generating ice crystallization process data; quantifying the heat conduction effect to generate freezing response data; evaluating tissue damage, and generating depth prediction data; constructing a decision model, and generating regulation and control guidance data; the treatment effect is comprehensively evaluated, and a standardized electronic report is automatically generated; according to the invention, temperature monitoring and individualized scheme optimization in the whole process of cervical cryotherapy are realized, and the therapeutic accuracy and safety are improved. | 1A method for temperature data collection and report generation for cervical cryotherapy, characterized in that it includes: step S1: capturing the freezing interface temperature gradient based on a multi-point thermal probe array to generate raw data on cervical tissue temperature distribution; performing heat diffusion path tracing analysis based on the raw data on cervical tissue temperature distribution to generate freezing front migration characteristic data; step S2: locating and identifying the heat diffusion critical point based on the freezing front migration characteristic data to generate freezing depth critical value data; performing interface temperature field analysis based on the freezing depth critical value data to generate a three-dimensional distribution characteristic map of the freezing zone temperature field; step S3: dynamically characterizing tissue ice crystal formation based on the three-dimensional distribution characteristic map of the freezing zone temperature field to generate tissue ice crystallization process data; performing multi-level heat conduction effect quantification based on the tissue ice crystallization process data to generate tissue freezing response channel characteristic data; Step S4: Map the tissue damage assessment under different freezing cycles for the tissue freezing response channel characteristic data to generate treatment effect depth prediction data; construct an individualized freezing plan decision model based on the treatment effect depth prediction data to generate freezing parameter dynamic control guidance data; conduct a comprehensive evaluation of the treatment effect based on the freezing parameter dynamic control guidance data and the three-dimensional distribution characteristic map of the freezing zone temperature field, and automatically generate a standardized electronic report for cervical cryotherapy. 2.The temperature data collection and report generation method for cervical cryotherapy according to claim 1 is characterized in that, step S1 includes: using a ring-distributed thermal probe array to capture the freezing interface temperature gradient to generate the original data of cervical tissue temperature distribution; performing temperature-time curve deconstruction processing based on the original data of cervical tissue temperature distribution to generate non-uniform cooling characteristic analysis data; calculating the tissue thermal diffusion coefficient based on the non-uniform cooling characteristic analysis data to generate patient-specific thermal conduction characteristic data; performing heat diffusion path tracking analysis based on the patient-specific thermal conduction characteristic data to generate freezing front migration characteristic data. 3.The temperature data collection and report generation method for cervical cryotherapy according to claim 2 is characterized in that, step S2 includes: calculating the temperature contour change rate based on the freezing front migration characteristic data to generate freezing rate gradient change data; accurately locating the tissue freezing transition point based on the freezing rate gradient change data to generate freezing depth critical value data; using high-precision thermal imaging technology to perform temperature field scanning on the freezing depth critical value data to generate deep temperature distribution high-density sampling data; accurately depicting and quantifying the freezing zone boundary based on the deep temperature distribution high-density sampling data to generate a freezing range boundary map; performing interface temperature field analysis based on the freezing range boundary map to generate a three-dimensional distribution characteristic map of the freezing zone temperature field. 4The temperature data collection and report generation method for cervical cryotherapy according to claim 3 is characterized in that the method accurately depicts and quantifies the freezing zone boundary based on the deep temperature distribution high-density sampling data to generate a freezing range boundary map, including: calculating the temperature gradient vector based on the deep temperature distribution high-density sampling data to generate a temperature gradient vector field number. According to the temperature gradient vector field data, the temperature gradient vector field data includes freezing peak point distribution data and temperature reduction rate quantification data; based on the freezing peak point distribution data, low-temperature area clustering analysis is performed to generate freezing cluster area data; based on the freezing cluster area data and the temperature reduction rate quantification data, potential unfrozen areas are identified to generate freezing dead corner warning data; based on the freezing dead corner warning data, high-density sampling data of deep temperature distribution are quantitatively graded to generate a freezing range boundary map. 5.The temperature data collection and report generation method for cervical cryotherapy according to claim 4 is characterized in that step S3 includes: dynamically characterizing tissue ice crystal formation based on the three-dimensional distribution characteristic map of the freezing zone temperature field, and generating tissue ice crystal process data; analyzing the growth pattern of ice crystals inside and outside the cells based on the tissue ice crystal process data, and generating microscopic ice crystal morphological characteristic data; quantifying multi-level heat conduction effects based on the microscopic ice crystal morphological characteristic data, and generating tissue freezing response channel characteristic data. 6.The temperature data collection and report generation method for cervical cryotherapy according to claim 5 is characterized in that, the analysis of ice crystal growth patterns inside and outside the cells based on tissue ice crystal process data to generate microscopic ice crystal morphological characteristic data includes: performing microscopic freezing injury mechanism analysis based on tissue ice crystal process data to generate cell freezing mechanism characteristic data; performing tissue dehydration effect analysis based on tissue ice crystal process data to generate cell osmotic pressure change characteristic data; constructing an ice crystal growth pattern classification matrix based on cell freezing mechanism characteristic data and cell osmotic pressure change characteristic data to generate an ice crystal morphological evolution matrix; performing analysis of ice crystal growth patterns inside and outside the cells based on the ice crystal morphological evolution matrix to generate microscopic ice crystal morphological characteristic data. 7.The temperature data collection and report generation method for cervical cryotherapy according to claim 6 is characterized in that, the multi-level heat conduction effect quantification based on the microscopic ice crystal morphological characteristic data to generate tissue freezing response channel characteristic data includes: performing micro-macro freezing effect correlation analysis under the influence of tissue blood supply factors on the microscopic ice crystal morphological characteristic data to generate cross-scale freezing effect linkage characteristic data; performing multi-level heat conduction effect quantification based on the cross-scale freezing effect linkage characteristic data to generate tissue freezing response channel characteristic data. 8.The temperature data collection and report generation method for cervical cryotherapy according to claim 7 is characterized in that, step S4 includes: performing tissue damage assessment mapping under different freezing cycles on the tissue freezing response channel characteristic data to generate treatment effect depth prediction data; extracting the freezing-thawing cycle influencing factors based on the treatment effect depth prediction data to generate freezing cycle impact quantitative data, and establishing tissue necrosis prediction distribution nodes through the freezing cycle impact quantitative data; constructing an individualized freezing scheme decision model for the tissue necrosis prediction distribution nodes and the treatment effect depth prediction data to generate freezing parameter dynamic control guidance data; performing matching analysis on the three-dimensional distribution characteristic map of the freezing zone temperature field and the treatment effect depth prediction data to generate disease Lesion-treatment range fit assessment data; Based on the lesion-treatment range fit assessment data and the dynamic control guidance data of freezing parameters, a comprehensive evaluation of the cryotherapy effect is conducted to generate a treatment integrity evaluation result; Based on the treatment integrity evaluation results, basic patient information, treatment parameter records, temperature collection data and efficacy predictions are integrated to automatically generate a standardized electronic report for cervical cryotherapy. 9. A temperature data collection and report generation device for cervical cryotherapy, characterized in that the wiring harness terminal intelligent assembly device includes: a memory and at least one processor, the memory storing instructions; the at least one processor calls the instructions in the memory to enable the wiring harness terminal intelligent assembly device to execute the temperature data collection and report generation method for cervical cryotherapy as described in any one of claims 1-8. 10. A computer-readable storage medium having instructions stored thereon, wherein when the instructions are executed by a processor, the temperature data collection and report generation method for cervical cryotherapy as described in any one of claim 187 is implemented. | 2025 | JIANGSU HAIHUI DATONG MEDICAL TECH | NA | NA | NA | China | The patent focuses on a cervical cryotherapy temperature acquisition and report-generation method using a multi-point thermosensitive probe array to map temperature gradients, analyze heat-diffusion paths, and locate critical points. It addresses the limited temperature data and reporting in conventional cervical cryotherapy. The output includes freezing front migration data and depth criteria for clinical reporting. | הפטנט מתמקד בשיטת רכישת טמפרטורה ויצירת דוח לקריו-תרפיה של צוואר הרחם המשתמשת במערך גשוש תרמו-רגיש רב-נקודתי למיפוי שיפועי טמפרטורה, ניתוח נתיבי דיפוזיית חום ואיתור נקודות קריטיות. הוא מתמודד עם נתוני הטמפרטורה והדיווח המוגבלים בקריו-תרפיה צווארית קונבנציונלית. הפלט כולל נתוני הגירת חזית ההקפאה וקריטריוני עומק לדיווח קליני. |  |  |  |  | 1 | [9 empty] | 1 | [5 empty] | 1 | 1 | [9 empty] | 1 | JIANGSU HAIHUI DATONG MEDICAL TECH CO LTD | JIANGSU HAIHUI DATONG MEDICAL TECH CO LTD | JIANGSU HAIHUI DATONG MEDICAL TECH CO LTD | JIANGSU HAIHUI DATONG MEDICAL TECH CO LTD | JIANGSU HAIHUI DATONG MEDICAL TECH CO LTD | 2025-06-06 | 2025 | 2025-09-12 | 2025 | 2025-06-06 | 2025 | Application | Utility Invention | active | CN202510751552 | 323327199 | [NA] | CN
392 | CN119564320A | Ablation parameter determination method and device for high-frequency pulse electric field ablation | The invention discloses an ablation parameter determination method and device for high-frequency pulsed electric field ablation, and the method comprises the steps: obtaining collected medical image data containing depth information, and determining a corresponding lesion box containing the three-dimensional information of a lesion according to the medical image data; analyzing and segmenting the lesion box to obtain the contour and depth of the lesion and the contour of an electrode for performing ablation on the lesion; taking a plurality of equidistant points on the boundary of the electrode contour as starting points, respectively emitting a preset number of rays to the periphery, and obtaining the number of intersection points of the rays and the focus box to determine whether the electrode is attached to the focus; when the electrode is attached to the focus, corresponding ablation parameters are determined according to the contour and depth of the focus, and the ablation parameters comprise the pulse voltage, the pulse width, the pulse string number and the discharge frequency. The problem that in the prior art, ablation parameters are not accurately determined is solved. | 1.A method for determining ablation parameters of high-frequency pulse electric field ablation, characterized in that the method comprises: acquiring collected medical image data containing depth information, and determining a corresponding lesion box containing three-dimensional information of the lesion based on the medical image data; analyzing and segmenting the lesion box to obtain the contour and depth of the lesion and the contour of the electrode used to ablate the lesion; taking a plurality of equally spaced points on the boundary of the electrode contour as starting points, emitting a preset number of rays to the surroundings respectively, and obtaining the number of intersections between the rays and the lesion box to determine whether the electrode fits the lesion; when the electrode fits the lesion, determining the corresponding ablation parameters based on the contour and depth of the lesion, wherein the ablation parameters include pulse voltage, pulse width, number of pulse trains, and number of discharges. 2.The method for determining ablation parameters of high-frequency pulsed electric field ablation according to claim 1 is characterized in that the step of obtaining the collected medical image data containing depth information and determining the corresponding lesion box containing three-dimensional information of the lesion according to the medical image data includes: inputting the medical image data into a pre-trained lesion localization model to obtain a corresponding lesion box containing three-dimensional information of the lesion; wherein the training process of the lesion localization model is: collecting a preset number of historical medical image data, each of which corresponds to a preset frame to frame the lesion range and mark the lesion type to obtain a training data set; dividing the training data set into a training set, a validation set and a test set according to a preset ratio, and then inputting the training data set into a preset neural network for training until the model parameters tend to be stable to obtain the lesion localization model. 3.The method for determining ablation parameters of high-frequency pulse electric field ablation according to claim 1 is characterized in that the step of analyzing and segmenting the lesion box to obtain the contour and depth of the lesion and the contour of the electrode used to ablate the lesion includes: inputting the lesion box into a pre-trained lesion segmentation model to obtain the corresponding contour and depth of the lesion and the contour of the electrode used to ablate the lesion; wherein the training process of the lesion segmentation model is: collecting a preset number of historical lesion boxes and the corresponding contours and depths of the lesions and the contours of the electrodes used to ablate the lesions to obtain a training data set; dividing the training data set into a training set, a validation set and a test set according to a preset ratio, and then inputting the training data set into a preset neural network for training until the model parameters tend to be stable to obtain the lesion segmentation model. 4.The method for determining ablation parameters of high-frequency pulse electric field ablation according to claim 1 is characterized in that the step of obtaining the number of intersections between the ray and the lesion box to determine whether the electrode and the lesion are in contact with each other includes: when the number of intersections between the ray and the lesion box is an even number, it is determined that the electrode and the lesion are not in contact with each other; when the number of intersections between the ray and the lesion box is an odd number, it is determined that the electrode and the lesion are in contact with each other. 5.The method for determining ablation parameters of high-frequency pulse electric field ablation according to claim 1 is characterized in that the step of determining the corresponding ablation parameters according to the contour and depth of the lesion comprises: inputting the contour and depth of the lesion into a preset first ablation parameter determination model to obtain the corresponding ablation parameters; wherein the training process of the first ablation parameter determination model is: collecting historical ablation parameters and the corresponding ablation depth and ablation length as a training data set, and inputting the training data into a first ablation parameter determination model; The data set is input into the naive Bayesian network for training to obtain the first ablation parameter determination model; the naive Bayesian network is based on the Bayesian probability formula, and the optimal relationship between the ablation depth, ablation length and ablation parameters is adjusted through multiple training; wherein, the mathematical expression of the Bayesian formula is: (4)·Ρ(Α) ~Ρ(^)~ wherein, P (A IB) represents the probability of event 4 occurring under the condition that event 8 occurs, P (BlΑ) represents the probability of event 8 occurring under the condition that event 4 occurs, Ρ (Α) is the prior probability of event 4 occurring independently, Ρ (Β) is the marginal probability of event 8 occurring, event Α= {pulse voltage, pulse width, number of pulse trains, number of discharges}, event 8= {ablation length, ablation depth}. 6.The method for determining ablation parameters of high-frequency pulse electric field ablation according to claim 1 is characterized in that, after the step of determining the corresponding ablation parameters according to the contour and depth of the lesion when the electrode is in contact with the lesion, it also includes: after completing an ablation discharge using the ablation parameters, obtaining the impedance data, contour and depth of the lesion within a preset time period after the end of this ablation discharge; inputting the impedance data and the contour and depth of the lesion after this ablation discharge into a pre-trained second ablation parameter determination model to obtain the corresponding secondary ablation parameters. 7.The ablation parameter determination method of high-frequency pulse electric field ablation according to claim 6 is characterized in that the training process of the second ablation parameter determination model is: collecting historical impedance data, secondary ablation parameters and corresponding ablation depth and ablation length as a training data set, and inputting the training data set into the Markov network for training to obtain the second ablation parameter determination model; the Markov network is based on the Bayesian probability formula, and adjusts the optimal relationship between the ablation depth, ablation length and secondary ablation parameters through multiple training; wherein, the mathematical expression of the Bayesian formula is: wherein /(4/8, represents the probability of event 4, / occurring under the condition that event 8 occurs /6/4,. represents the probability of event I occurring under the condition that event 4 occurs / (4 represents the probability of event I occurring under the condition that event 4 occurs / (B11) represents the probability of event / occurring under the condition that event 8 occurs, event Α= {pulse voltage, pulse width, number of pulse trains, number of discharges}, event 8= {ablation length, ablation depth}, / is impedance data. 8.A device for determining ablation parameters of high-frequency pulse electric field ablation, characterized in that the device comprises: an acquisition module, used to acquire collected medical image data containing depth information, and determine the corresponding lesion box containing three-dimensional information of the lesion according to the medical image data; a segmentation module, used to analyze and segment the lesion box to obtain the contour and depth of the lesion and the contour of the electrode used to ablate the lesion; a collision module, used to take a plurality of equally spaced points on the boundary of the electrode contour as starting points, respectively emit a preset number of rays to the surroundings, and obtain the number of intersections between the rays and the lesion box to determine whether the electrode and the lesion are in contact; a determination module, used to determine the corresponding ablation parameters according to the contour and depth of the lesion when the electrode and the lesion are in contact; Or it is used to determine corresponding ablation parameters according to the contour and depth of the lesion when the electrode is in contact with the lesion, and to obtain the impedance data as well as the contour and depth of the lesion after the corresponding ablation parameters complete one ablation discharge, and to determine corresponding secondary ablation parameters based on the impedance data and the contour and depth of the lesion; wherein the ablation parameters and the secondary ablation parameters both include pulse voltage, pulse width, number of pulse trains and number of discharges. 9.A readable storage medium having a computer program stored thereon, characterized in that when the program is executed by a processor, the steps of the method as claimed in any one of claims 1 to 7 are implemented. 10. An electronic device, characterized in that it includes a memory, a processor, and a computer program stored in the memory and executable on the processor, wherein when the processor executes the program, the steps of the method described in any one of claims 1 to 7 are implemented. | 2024 | JIANGXI SHUNJIE MEDICAL TECH | NA | NA | NA | China | The patent focuses on a high-frequency PFA parameter-determination method that pulls 3D depth-bearing medical images, builds a focus box around the lesion, and analyzes lesion outline/depth and electrode outline. It addresses the difficulty of choosing PFA parameters that match the actual 3D lesion geometry rather than using one-size-fits-all settings. Geometry from imaging then drives the choice of ablation parameters. | הפטנט מתמקד בשיטת קביעת פרמטרי PFA בתדר גבוה המושכת תמונות רפואיות תלת-ממדיות בעלות עומק, בונה תיבת מיקוד סביב הנגע ומנתחת מתאר/עומק נגע ומתאר אלקטרודה. הוא מתמודד עם הקושי בבחירת פרמטרי PFA המתאימים לגיאומטריית הנגע התלת-ממדית האמיתית במקום להשתמש בהגדרות אחידות. הגיאומטריה מההדמיה ואז מניעה את בחירת פרמטרי האבלציה. | 1 | [7 empty] | 1 |  |  | 1 | [10 empty] | 1 | [7 empty] | 1 |  | JIANGXI SHUNJIE MEDICAL TECH CO LTD | JIANGXI SHUNJIE MEDICAL TECH CO LTD | JIANGXI SHUNJIE MEDICAL TECH CO LTD | JIANGXI SHUNJIE MEDICAL TECH CO LTD | JIANGXI SHUNJIE MEDICAL TECH CO LTD | 2024-12-18 | 2024 | 2025-03-07 | 2025 | 2024-12-18 | 2024 | Application | Utility Invention | inactive | CN202411869784 | 319596422 | CN119564320B | CN
393 | CN113994380A | METHOD FOR DETERMINING AN ABLATION REGION BASED ON DEEP LEARNING | The invention relates to a method for evaluating in post-treatment an ablation of a portion of an anatomy of interest (130) of an individual (110), the anatomy of interest comprising at least one lesion (165). The evaluation method comprises in particular a step of automatically determining a contour of the ablation region by means of an automatic learning method, such as a neural network, analyzing the post-treatment image of the anatomy of interest of the individual, said automatic learning method being preloaded (290) during a so-called training phase using a database comprising a plurality of post-operative medical images of an anatomy of identical interest of a set of patients, each medical image of the database being associated with an ablation region of the anatomy of interest of said patient. The invention also relates to an electronic device (150) comprising a processor and a computer memory storing instructions of such an evaluation method. | 1A post-treatment assessment method (200) for ablation of a portion of a target anatomy (130) of an individual (110), the target anatomy comprising at least one lesion (165, 510, 600), a portion of the target anatomy The ablation is defined by an ablation zone (160), characterized in that the evaluation method comprises the steps of: acquiring (210) a post-operative medical image of the target anatomy of the individual, including all or part of the ablation zone; The machine learning method automatically determines the contour of the ablation area, analyzes the individual's post-treatment images of the target anatomy, pre-trained in the training phase using a database of multiple post-operative medical images of the same target anatomy for a set of patients (290) In the machine learning method, each medical image in the database is associated with an ablation region of a target anatomy of the patient. 2. The post-treatment evaluation method of claim 1, wherein the training phase includes a preceding step of training using medical images of the same target anatomy containing unablated lesions. 3. The post-treatment assessment method according to any one of the preceding claims, further comprising the step (220) of registering a post-operative image of a target anatomy of the individual with a medical image (170, 500), the medical Images obtained before surgical treatment are called preoperative medical images. 4. The post-treatment assessment method of any preceding claim, further comprising assessing between the ablation area and the lesion, between the ablation area and the target anatomy, or between the lesion and the lesion Steps to assess recurrence risk by relative characteristics between target anatomical structures. 5. A post-treatment assessment method according to any preceding claim, further comprising, when a risk of recurrence is demonstrated, according to between the ablation area and the lesion, the ablation area and the target anatomy The step of determining the location of recurrence based on relative characteristics between the lesion or between the lesion and the target anatomy. 6. The post-treatment evaluation method according to any one of claims 4 and 5, wherein the relative feature is the ablation margin between the ablation area and the lesion. 7. The post-treatment evaluation method according to any one of claims 4 to 6, wherein the relative features are the centroid of the lesion and the centroid of the ablation area. 8. The post-treatment evaluation method of any one of claims 4 to 7, wherein the relative characteristics are the uniformity and sharpness of the edges of the ablation zone relative to surrounding healthy tissue. 9. The post-treatment evaluation method according to any one of claims 4 to 8, wherein the relative feature is the ratio of the volume of the lesion to the volume of the ablation area. 10. The post-treatment assessment method of any one of claims 4 to 9, wherein the relative feature is the location of the lesion relative to the center of the target anatomy. 11. The post-treatment assessment method of any one of claims 3 to 10, further comprising the step of segmenting the lesion in a preoperative medical image of the subject's target anatomy. 12. The post-treatment assessment method of any one of claims 3 to 11, further comprising the step of detecting the lesion in a preoperative medical image of the individual target anatomy. 13. The post-treatment evaluation method according to any one of claims 3 to 12, wherein all or part of the training post-operative medical images in the database are cropped around an ablation region comprising at least one lesion, the cropping of the image being cropped This is performed using a common frame of predetermined dimensions, within which the collection of ablation region centers in the cropped post-operative medical image forms clusters of distinct points. 14. The post-treatment assessment method according to any one of claims 3 to 13, wherein, for the set of medical images in the database, the parts of an individual's body included in the images are divided into a plurality of single-sized basic unit, The number of base cells is divided into two nearly equal parts between the part of the human body bounded by the ablation area and the rest of the individual's body included in the image. 15. The post-treatment evaluation method according to any one of claims 3 to 14, wherein the post-operative medical image database comprises at least one pre-operative medical image containing at least one unablated lesion. 16. The post-treatment assessment method of any one of claims 4 to 15, further comprising the step of determining a supplemental ablation mask when recurrence risk is demonstrated. 17. A post-treatment assessment method according to any preceding claim, further comprising the step of proposing a path to be followed by a medical instrument to a target point of the supplemental ablation mask. 18. A post-treatment assessment method according to any preceding claim, wherein the medical image is a three-dimensional image. 19. A post-treatment assessment method according to any one of claims 3 to 18, wherein each post-operative image is acquired using the same image acquisition technique. 20. An electronic device (150) comprising a processor and a computer memory storing instructions for a post-treatment assessment method according to any preceding claim. 21. twenty one. An electronic device according to the preceding claim, wherein the device may be a control device, a navigation system, a robotic device or an augmented reality device. | 2020 | QUANTUM SURGICAL | Medical and Healthcare | Small-size company | https://quantumsurgical.com | France | The patent focuses on a deep-learning method that automatically determines ablation-region contours from post-treatment images using a trained neural network. It addresses the gap in post-treatment assessment of percutaneous ablation, where recurrence prediction is hard without precise contour analysis. By comparing pre- and post-treatment images, it produces a quantitative post-ablation assessment. | הפטנט מתמקד בשיטת למידה עמוקה הקובעת אוטומטית מתארי אזור אבלציה מתמונות לאחר טיפול באמצעות רשת עצבית מאומנת. הוא מתמודד עם הפער בהערכה לאחר טיפול של אבלציה פרקוטנית, שבה ניבוי הישנות קשה ללא ניתוח מתאר מדויק. על ידי השוואת תמונות לפני ואחרי הטיפול, הוא מייצר הערכה כמותית לאחר אבלציה. | [8 empty] | 1 | [11 empty] | 1 |  | 1 | [5 empty] | 1 |  | 1 |  | QUANTUM SURGICAL | QUANTUM SURGICAL | QUANTUM SURGICAL | QUANTUM SURGICAL | QUANTUM SURGICAL | 2021-05-20 | 2021 | 2022-01-28 | 2022 | 2020-05-20 | 2020 | Application | Utility Invention | inactive | CN20218003924 | 317518931 | CA3176333A1¦CN113994380B¦EP4154263A1¦FR3110763A1¦FR3110763B1¦IL298313A¦JP2023526909A¦JP7834030B2¦KR20230013041A¦US20230177681A1¦WO2021234305A1 | CN
394 | US20240001487A1 | MEDICAL TREATMENT APPARATUS AND TREATMENT PROBE THEREOF | Disclosed is a medical treatment apparatus. The medical treatment apparatus includes a magnetic resonance imaging (MRI) device, configured for imaging of a specific region including a target object and generating a magnetic resonance image; a laser interstitial thermal therapy (LITT) device, including an LITT probe, the LITT probe being positioned close to the target object based on the magnetic resonance image, and being configured to treat the target object by emitting laser; and a temperature measurement element, the temperature measurement element and the LITT probe being integrated as an integrated probe, and the temperature measurement element being configured to obtain a temperature of the target object. | 1. A laser interstitial thermal therapy (LITT) lateral ablation probe, which is used in an LITT device, wherein the LITT lateral ablation probe comprises: a probe main body; and a coating layer covered on an end surface of the probe main body, wherein the end surface of the probe main body and an axis of the probe main body form a first angle, the probe main body comprises a probe core and a hard cladding structure located at an outer peripheral of the probe core, the probe core is made of pure silicon oxide material, the hard cladding structure is made of technology enhanced clad silica (TECS) material, and the coating layer comprises noble metal. 2. The LITT lateral ablation probe of claim 1, wherein the coating layer has a dual-layer structure, a layer of the dual-layer structure which is close to the end surface of the probe main body comprises pure silver, and the other layer of the dual-layer structure which is away from the end surface of the probe main body comprises silicon monoxide. 3. The LITT lateral ablation probe of claim 1, wherein the LITT lateral ablation probe is obtained by brightening and cleaning the end surface of the probe main body, and then further covering the coating layer on the end surface through magnetron sputtering. 4. A laser interstitial thermal therapy (LITT) lateral ablation probe, which is used in an LITT device, wherein the LITT lateral ablation probe comprises: a probe main body; and a lens connected to an end surface of the probe main body, wherein the end surface of the probe main body is perpendicular to an axis of the probe main body, the lens has a first vertex angle, the probe main body comprises a probe core and a hard cladding structure located at an outer peripheral of the probe core, and the lens is a sapphire lens. 5. The LITT lateral ablation probe of claim 4, wherein the lens is a chamfered cylindrical lens, or a chamfered hemi-spherical or hemi-ellipsoidal lens. 6. The LITT lateral ablation probe of claim 4, wherein the LITT lateral ablation probe is obtained by brightening and cleaning the end surface of the probe main body, connecting the probe main body and the lens through fusion welding, and processing a welded surface between the probe main body and the lens using fire polishing. 7. A laser interstitial thermal therapy (LITT) circumferential ablation probe, which is used in an LITT device, wherein the LITT circumferential ablation probe comprises: a probe main body, the main body of the probe comprising a cone-shaped surface, one or more grooves being set on the cone-shaped surface, the one or more grooves being evenly distributed on the cone-shaped surface in an arabesquitic pattern, wherein the LITT circumferential ablation probe emits laser out of the one or more grooves. 8. The LITT circumferential ablation probe of claim 7, wherein a diameter of the cone-shaped surface decreases from an initial diameter to a preset diameter, the initial diameter is 630 μm, and the preset diameter is 100 μm. 9. The LITT circumferential ablation probe of claim 7, wherein the arabesquitic pattern comprises a single thread pattern, a cross-thread pattern, or a rhombic grid pattern, or a combination thereof. 10. The LITT circumferential ablation probe of claim 7, wherein the cone-shaped surface is formed by performing operations comprising: fixing two ends of a first workpiece with cone-shaped surface to be carved on a sliding rail and a fixing device, respectively; controlling, by a laser device controller, a laser device, to emit laser, wherein a power of the emitted laser is 30 W, a wavelength of the emitted laser is 10600 nm, the emitted laser is attenuated to 6.6-11.2 W by a laser power attenuator, the attenuated laser is split into two laser beams by a diffraction splitting lens unit, and the two laser beams have equal power and are focused onto a surface of the first workpiece by a focusing lens unit; and driving, by a motion driver, the first workpiece to move, the motion driver being controlled by a motion driver controller. 11. The LITT circumferential ablation probe of claim 7, wherein the one or more grooves that are evenly distributed on the cone-shaped surface in the arabesquitic pattern are formed by performing operations comprising: fixing two ends of a second workpiece with arabesquitic pattern to be carved on a sliding rail and a fixing device, respectively; controlling, by a laser device controller, a laser device, to emit laser, wherein a power of the emitted laser is 30 W, a wavelength of the emitted laser is 10600 nm, the emitted laser is focused on a surface of the second workpiece by a lens set, the lens set comprises a first concave lens, a first convex lens, and a second concave lens, the first concave lens and the first convex lens are used to expand a diameter of the emitted laser into a first diameter, the second concave lens is used to focus the laser beam of the first diameter to form a spot on the surface of the second workpiece, and the spot has a diameter of 45 μm; and driving, by a motion driver, the second workpiece to move, the motion driver being controlled by a motion driver controller. 12. The LITT circumferential ablation probe of claim 7, wherein a spatial intensity distribution of the laser emitted by the LITT circumferential ablation probe is obtained through operations comprising: obtaining signals output by a laser measurement sensor, and generating a polar intensity of a specific orientation of the laser emitted by the LITT circumferential ablation probe, wherein the LITT circumferential ablation probe is connected to a helium-neon laser device that generates laser with a wavelength of 632.8 nm, a slit diaphragm is set between the laser measurement sensor and the LITT circumferential ablation probe, and a slit size of the slit diaphragm is 0.3 mm; and determining the spatial intensity distribution of the laser emitted by the LITT circumferential ablation probe by driving, by a motion driver, the slit diaphragm and the laser measurement sensor to move along a circumferential direction and an axial direction of the cone-shape surface of the LITT circumferential ablation probe, the motion driver being controlled by a motion driver controller, wherein the slit diaphragm and the laser measurement sensor are fixedly connected to the motion driver. 13. A medical treatment apparatus, comprising: a magnetic resonance imaging (MRI) device, configured for imaging of a specific region comprising a target object and generating a magnetic resonance image; a laser interstitial thermal therapy (LITT) device, comprising: an LITT probe, the LITT probe being positioned close to the target object based on the magnetic resonance image, and configured to treat the target object by emitting laser; and a temperature measurement element, the temperature measurement element and the LITT probe being integrated as an integrated probe, and the temperature measurement element being configured to obtain a temperature of the target object, wherein the LITT probe comprises an LITT lateral ablation probe or an LITT circumferential ablation probe. 14. The medical treatment apparatus of claim 13, wherein the LITT lateral ablation probe comprises: a probe main body; and a coating layer covered on an end surface of the probe main body, wherein the end surface of the probe main body and an axis of the probe main body form a first angle, the probe main body comprises a probe core and a hard cladding structure located at an outer peripheral of the probe core, and the coating layer comprises noble metal. 15. The medical treatment apparatus of claim 14, wherein the coating layer has a dual-layer structure, a layer of the dual-layer structure which is close to the end surface of the probe main body comprises pure silver, and the other layer of the dual-layer structure which is away from the end surface of the probe main body comprises silicon monoxide. 16. The medical treatment apparatus of claim 14, wherein the LITT lateral ablation probe is obtained by brightening and cleaning the end surface of the probe main body, and then further covering the coating layer on the end surface through magnetron sputtering. 17. The medical treatment apparatus of claim 13, wherein the LITT lateral ablation probe comprises: a probe main body; and a lens connected to an end surface of the probe main body, wherein the end surface of the probe main body is perpendicular to an axis of the probe main body, the lens has a first vertex angle, the probe main body comprises a probe core and a hard cladding structure located at an outer peripheral of the probe core, and the lens is a sapphire lens. 18. The medical treatment apparatus of claim 17, wherein the lens is a chamfered cylindrical lens, or a chamfered hemi-spherical or hemi-ellipsoidal lens. 19. The medical treatment apparatus of claim 17, wherein the LITT lateral ablation probe is obtained by brightening and cleaning the end surface of the probe main body, connecting the probe main body and the lens through fusion welding, and processing a welded surface between the probe main body and the lens using fire polishing. 20. The medical treatment apparatus of claim 13, wherein the LITT circumferential ablation probe comprises: a probe main body, the main body of the probe comprising a cone-shaped surface, one or more grooves being set on the cone-shaped surface, the one or more grooves being evenly distributed on the cone-shaped surface in an arabesquitic pattern, wherein the LITT circumferential ablation probe emits laser out of the one or more grooves. 21. The medical treatment apparatus of claim 20, wherein a diameter of the cone-shaped surface decreases from an initial diameter to a preset diameter, the initial diameter is 630 μm, and the preset diameter is 100 μm. 22. The medical treatment apparatus of claim 20, wherein the arabesquitic pattern comprises a single thread pattern, a cross-thread pattern, or a rhombic grid pattern, or a combination thereof. 23. The medical treatment apparatus of claim 20, wherein the cone-shaped surface is formed by performing operations comprising: fixing two ends of a first workpiece with cone-shaped surface to be carved on a sliding rail and a fixing device, respectively; controlling, by a laser device controller, a laser device, to emit laser, wherein a power of the emitted laser is 30 W, a wavelength of the emitted laser is 10600 nm, the emitted laser is attenuated to 6.6-11.2 W by a laser power attenuator, the attenuated laser is split into two laser beams by a diffraction splitting lens unit, and the two laser beams have equal power and are focused onto a surface of the first workpiece by a focusing lens unit; and driving, by a motion driver, the first workpiece to move, the motion driver being controlled by a motion driver controller. 24. The medical treatment apparatus of claim 20, wherein the one or more grooves that are evenly distributed on the cone-shaped surface in the arabesquitic pattern are formed by performing operations comprising: fixing two ends of a second workpiece with arabesquitic pattern to be carved on a sliding rail and a fixing device, respectively; controlling, by a laser device controller, a laser device, to emit laser, wherein a power of the emitted laser is 30 W, a wavelength of the emitted laser is 10600 nm, the emitted laser is focused on a surface of the second workpiece by a lens set, the lens set comprises a first concave lens, a first convex lens, and a second concave lens, the first concave lens and the first convex lens are used to expand a diameter of the emitted laser into a first diameter, the second concave lens is used to focus the laser beam of the first diameter to form a spot on the surface of the second workpiece, and the spot has a diameter of 45 μm; and driving, by a motion driver, the second workpiece to move, the motion driver being controlled by a motion driver controller. 25. The medical treatment apparatus of claim 20, wherein a spatial intensity distribution of the laser emitted by the LITT circumferential ablation probe is obtained through operations comprising: obtaining signals output by a laser measurement sensor, and generating a polar intensity of a specific orientation of the laser emitted by the LITT circumferential ablation probe, wherein the LITT circumferential ablation probe is connected to a helium-neon laser device that generates laser with a wavelength of 632.8 nm, a slit diaphragm is set between the laser measurement sensor and the LITT circumferential ablation probe, and a slit size of the slit diaphragm is 0.3 mm; and determining the spatial intensity distribution of the laser emitted by the LITT circumferential ablation probe by driving, by a motion driver, the slit diaphragm and the laser measurement sensor to move along a circumferential direction and an axial direction of the cone-shape surface of the LITT circumferential ablation probe, the motion driver being controlled by a motion driver controller, wherein the slit diaphragm and the laser measurement sensor are fixedly connected to the motion driver. | 2022 | SHANGHAI HROCT LITT MEDICAL TECHNOLOGY | NA | NA | NA | China | The patent focuses on a laser interstitial thermal therapy (LITT) apparatus combining an MRI scanner, a LITT probe positioned by MR imaging, and temperature monitoring, with a probe core of pure silicon oxide and hard cladding. It addresses the need for a robust, MR-guided LITT probe that delivers laterally oriented ablation. MR imaging guides probe placement while temperature feedback controls the laser dose. | הפטנט מתמקד במכשיר טיפול תרמי אינטרסטיציאלי בלייזר (LITT) המשלב סורק MRI, גשוש LITT הממוקם על ידי הדמיית MR וניטור טמפרטורה, עם ליבת גשוש מתחמוצת סיליקון טהורה וציפוי קשיח. הוא מתמודד עם הצורך בגשוש LITT איתן מונחה MR המספק אבלציה בכיוון רוחבי. הדמיית MR מנחה את מיקום הגשוש בעוד היזון חוזר של טמפרטורה שולט במינון הלייזר. |  |  |  |  | 1 |  |  |  | 1 |  |  |  |  | 1 | [14 empty] | 1 |  |  |  | SHANGHAI HROCT LITT MEDICAL TECH CO LTD | SHANGHAI HROCT LITT MEDICAL TECH CO LTD | SHANGHAI HROCT LITT MEDICAL TECHNOLOGY CO., LTD. | SHANGHAI HROCT LITT MEDICAL TECHNOLOGY CO LTD | SHANGHAI HROCT LITT MEDICAL TECHNOLOGY CO LTD | 2022-11-03 | 2022 | 2024-01-04 | 2024 | 2022-07-02 | 2022 | Application | Utility Invention | active | US202217980444 | 293634501 | CN117338407A¦CN117530770A¦CN117770947A¦US20240000507A1¦WO2024060314A1 | US
395 | CN117883175A | Radio frequency ablation system and computer readable medium | The invention relates to a radio frequency ablation system and a computer readable medium, the system comprises a radio frequency host, an electrode needle device and an electrode slice module, the radio frequency host comprises a main control module, a liquid cooling module, a temperature detection module, a man-machine interaction module, an electrode needle device connecting end and an electrode slice module connecting end; the temperature detection module is used for collecting the ablation temperature of the working end of the needle device and sending the ablation temperature to the main control module; the liquid cooling module adjusts the flow of liquid conveyed to the electrode needle device according to the ablation temperature and the control signal; the man-machine interaction module is used for sending a control signal to the main control module, the man-machine interaction module is further used for displaying system information, and the main control module is further connected with the electrode needle device connecting end and the electrode slice module connecting end. The computer readable medium is used for realizing the control method when the system carries out ablation. The radiofrequency ablation system and the computer readable medium have the advantages of being accurate, convenient, easy to operate and high in safety. | 1. A radiofrequency ablation system, characterized in that it includes a radiofrequency host, an electrode needle device and an electrode sheet module; the radiofrequency host includes a main control module, a liquid cooling module, a temperature detection module, a human-computer interaction module, and an electrode needle device connection terminal and the connection end of the electrode sheet module; the temperature detection module is connected to the main control module and the connection end of the electrode needle device respectively, and the temperature detection module is used to collect the temperature connected to the connection end of the electrode needle device. The ablation temperature of the working end of the electrode needle device is sent to the main control module; the liquid cooling module is connected to the electrode needle device and the main control module respectively. The liquid cooling module is connected to the control module, and the liquid cooling module adjusts the liquid flow delivered to the corresponding electrode needle device according to the ablation temperature and control signal; the human-computer interaction module is connected to the main control module, and the human-computer interaction module is connected to the main control module. The computer interaction module is used to receive the control signal and send the control signal to the main control module, and the human-computer interaction module is also used to display system information, where the system information includes the ablation temperature ; The main control module is also connected to the connecting end of the electrode needle device and the connecting end of the electrode sheet module respectively. 2. The radiofrequency ablation system according to claim 1, wherein the radiofrequency host further includes an electrical signal acquisition module, and the electrical signal acquisition module is connected to the main control module, the electrode needle device connection end and the electrode sheet module respectively. The electrical signal acquisition module is used to collect electrical signals from the electrode needle device connection end and the electrode sheet module connection end, and generate corresponding electrical signal information according to the electrical signal, and the The electrical signal acquisition module sends the electrical signal to the main control module. 3. The radiofrequency ablation system according to claim 2, characterized in that the main control module determines the electrode needle device connected to the connection end of the electrode needle device and the electrode sheet module based on the electrical signal. The connection state of the electrode sheet module at the connection end, and the connection state of the electrode needle device connected to the connection end of the electrode needle device and/or the electrode sheet module connected to the connection end of the electrode sheet module When an abnormality occurs, a corresponding first warning message is generated and sent to the human-computer interaction module. 4. The radiofrequency , wherein the main control module determines whether the tissue is coagulated and/or carbonized based on the electrical signal, and generates a signal when the tissue is coagulated and/or carbonized. The corresponding second warning prompt information is sent to the human-computer interaction module. 5. The radiofrequency ablation system according to claim 2, wherein the radiofrequency host further includes a radiofrequency generation module, and the radiofrequency generation module is respectively connected to the main control module, the electrode needle device connection end and the electrode sheet module. The electrical signal acquisition module is used to collect the real-time power value output by the radio frequency generating module, and send the real-time power value to the main control module, and the main control module determines the real-time power value according to the real-time power value. The power deviation value between the preset power value and the power value determines whether the radio frequency generating module is working normally, and when the radio frequency generating module works abnormally, a corresponding third warning message is generated and sent to the human-computer interaction module. 6. The radiofrequency ablation system according to claim 5, characterized in that the radiofrequency host further includes a power module, and the power module is respectively connected with the electrode needle device, electrode sheet module, main control module, liquid cooling module, The temperature detection module, the human-computer interaction module, the electrical signal acquisition module and the radio frequency generation module are connected; after the radio frequency ablation system is powered on, the electrical signal acquisition module collects the connection ends of the electrode needle device and the electrode sheet module respectively. The power-on status of the connection end, the main control module, the liquid cooling module, the temperature detection module, the human-computer interaction module and the radio frequency generation module is generated, and the corresponding power-on detection signal is generated and transmitted to the main control module. When the electrode needle device connection When any abnormality occurs in any of the terminals, electrode module connection terminals, main control module, liquid cooling module, temperature detection module, human-computer interaction module, electrical signal acquisition module and radio frequency generation module, the corresponding third message will be generated. Four early warning prompts are sent to the human-computer interaction module. 7. The radio frequency ablation system according to claim 6, wherein after the radio frequency ablation system is powered on, the main control module receives heartbeat data regularly sent by each monitored node in the radio frequency ablation system. When the heartbeat data of the monitored node is refreshed abnormally, it is determined that the monitored node is abnormal, and a corresponding fifth warning message is generated and sent to the human-computer interaction module; wherein the monitored node includes a computer connected to The electrode needle device at the connection end of the electrode needle device, the electrode sheet module connected to the electrode sheet module connection end, the liquid cooling module, the temperature detection module, the human-computer interaction module, the At least one component of the electrical signal acquisition module and the radio frequency generation module. 8. The radiofrequency ablation system according to claim 2, wherein the liquid cooling module includes a water pump, and the electrical signal acquisition module obtains the real-time rotational speed pulse value of the water pump in real time, and sends the real-time rotational speed pulse value. To the main control module, the main control module determines whether the water pump is working normally based on the pulse difference between the real-time speed pulse value and the preset speed pulse value corresponding to the preset flow rate of the water pump, and When the water pump works abnormally, a corresponding sixth warning message is generated and sent to the human-computer interaction module. 9. The radiofrequency ablation system according to claim 1, wherein the human-computer interaction module includes a touch screen, a control switch unit and a speaker; the touch screen and the speaker are respectively connected to the main control module; The control switch unit is respectively connected to the main control module and the electrode needle device. 10. The radiofrequency ablation system according to claim 1, characterized in that: the radiofrequency host further includes an electrode needle multi-channel detection control module; the electrode needle device connection end includes several electrode needle interfaces; the electrode needle device One or more electrode needle mechanisms can be selectively connected to one of the electrode needle interfaces; the temperature detection module and the main control module are connected to each of the electrodes through the electrode needle multi-channel detection control module The electrode needle multi-channel detection control module is used to send the connection status of the electrode needle mechanism detected at each of the electrode needle interfaces to the main control module, and connect it to the corresponding electrode needle The ablation temperature signal of the working end of the electrode needle mechanism at the interface is transmitted to the temperature detection module, so that the temperature detection module obtains the corresponding ablation temperature according to each of the ablation temperature signals; the main control module can according to The connection status of the electrode needle mechanism detected from each of the electrode needle interfaces, the ablation temperature of the working end of the electrode needle mechanism connected to the corresponding electrode needle interface, and the control signal are adjusted and sent to the corresponding The liquid flow rate of the electrode needle mechanism is controlled, and the working status of each electrode needle mechanism connected to the corresponding electrode needle interface is controlled. 11. The radiofrequency ablation system according to claim 10, wherein when two or more electrode needle mechanisms are connected to corresponding electrode needle interfaces, each of the pairs is connected to the corresponding electrode needle interface. Controlling the working status of the electrode needle mechanism at the electrode needle interface includes: controlling the electrode needle mechanism connected to the electrode needle interface to work in turn according to the system's preset power, and in the radiofrequency ablation system When any component is abnormal, the electrode needle mechanism is controlled to stop working. 12. A computer-readable medium, characterized in that a computer program is stored on the computer-readable medium, and when the computer program is executed by a processor, the radiofrequency ablation system according to any one of claims 1 to 11 is implemented. Control method, the control method includes the following steps: Real-time monitoring of the ablation temperature of the working end of the electrode needle device connected to the connection end of the electrode needle device, and sending the ablation temperature to the main control module and the A human-computer interaction module, the human-computer interaction module is used to display the ablation temperature; the liquid cooling module adjusts the liquid flow rate delivered to the corresponding electrode needle device according to the ablation temperature and the control signal. 13. The computer-readable medium according to claim 12, wherein the control method further includes the following steps: collecting electrical signals at the connecting end of the electrode needle device and the connecting end of the electrode sheet module in real time, and Corresponding electrical signal information is generated according to the electrical signal, and the electrical signal information includes loop status information of the connecting end of the electrode needle device and the connecting end of the electrode sheet module; determining the electrode needle device based on the loop status information and the connection status of the electrode sheet module and the radio frequency host, wherein when the loop status information shows that the electrode needle device connection end and/or the electrode sheet module connection end does not form a loop, then It is determined that the connecting end of the electrode needle device and/or the connecting end of the electrode sheet module is not connected to the corresponding electrode needle device and/or the electrode sheet module; when the circuit status information shows that the connecting end of the electrode needle device and/or when the connecting end of the electrode sheet module forms a loop, it is determined that the connecting end of the electrode needle device and/or the connecting end of the electrode sheet module is connected to the corresponding electrode needle device and/or electrode sheet module . 14. The computer-readable medium according to claim 13, wherein the control method further includes the following steps: when the radiofrequency ablation system performs an ablation operation, the circuit status information displays the connection end of the electrode needle device and/or the connecting end of the electrode sheet module does not form a loop, then determine the electrode needle device connected to the connecting end of the electrode needle device and/or the electrode sheet module connected to the connecting end of the electrode sheet module If the connection status of the group is abnormal, the corresponding first warning message will be generated to issue an alarm. 15. The computer-readable medium according to claim 13, wherein the electrical signal information further includes an impedance value formed by the connection end of the electrode needle device and the connection end of the electrode sheet module, and the control method further It includes the following steps: Determining the ablation status of the tissue according to the impedance value. When detecting that the impedance value exceeds a system preset threshold, it is determined that the tissue is coagulated and/or carbonized, and a corresponding second warning prompt is generated. information and make an alarm. 16. The computer-readable medium according to claim 13, wherein the radio frequency host further includes a radio frequency generation module, and the control method further includes the following steps: collecting the real-time power value output by the radio frequency generation module, and based on the The power deviation value between the real-time power value and the preset power value is used to determine whether the radio frequency generating module is working normally, and when the radio frequency generating module is working abnormally, a corresponding third warning message is generated to issue an alarm. 17. The computer-readable medium according to claim 16, characterized in that the control method further includes the following steps: after the radiofrequency ablation system is powered on, separately collecting the connecting end of the electrode needle device and the electrode sheet module The power-on status of the connection end, main control module, liquid cooling module, temperature detection module, human-computer interaction module and radio frequency generation module, and generates a pair of According to the power-on detection signal, when the electrode needle device connection end, the electrode sheet module connection end, the main control module, the liquid cooling module, the temperature detection module, the human-computer interaction module, the electrical signal acquisition module and the radio frequency generation module When an abnormality occurs in any component of If any component among the module, the human-computer interaction module and the radio frequency generation module cannot feed back the corresponding power-on detection signal, it is determined that the component is abnormal. 18. The computer-readable medium according to claim 16, wherein the control method further includes the following steps: regularly receiving heartbeat data regularly sent by each monitored node in the radio frequency ablation system, when any of the monitored nodes When there is a refresh abnormality in the heartbeat data of the monitoring node, it is determined that the monitored node is abnormal, and a corresponding fifth warning message is generated to issue an alarm; wherein the monitored node includes all the nodes connected to the connection end of the electrode needle device. The electrode needle device, the electrode sheet module connected to the electrode sheet module connection end, the liquid cooling module, the temperature detection module, the human-computer interaction module, the electrical signal acquisition module and the At least one component in the radio frequency generation module. 19. The computer-readable medium according to claim 13, wherein the liquid cooling module includes a water pump, and the control method further includes the following steps: obtaining the real-time speed pulse value of the water pump in real time, and based on the real-time speed The pulse difference between the pulse value and the preset speed pulse value corresponding to the preset flow rate of the water pump is used to determine whether the water pump is working normally, and when the water pump is working abnormally, the corresponding sixth warning message is generated. Call the police. 20. The computer-readable medium according to claim 13, wherein the electrode needle device includes two or more electrode needle mechanisms, and the control method further includes the following steps: when two or more electrode needle mechanisms When the electrode needle mechanism is connected to the connection end of the electrode needle device, the ablation temperature and the control signal collected from the working end of each electrode needle mechanism are adjusted and transmitted to the corresponding electrode needle mechanism. liquid flow, and control the working status of each electrode needle mechanism so that each electrode needle mechanism works at a preset operating temperature and continues to work for a preset working time. 21. twenty one. The computer-readable medium according to claim 13, wherein the control method further includes the following steps: when the radiofrequency ablation system is in manual working mode, the radiofrequency host receives the information according to the human-computer interaction module. The real-time control signal received controls the switch of the electrode needle device and the adjustment of the working power; when the radiofrequency ablation system is in the automatic working mode, the radiofrequency host controls the switch according to the preset control signal received by the human-computer interaction module. , control the preset working time of the electrode needle device, monitor the ablation temperature of the electrode needle device in real time, and adjust the liquid flow rate delivered by the liquid cooling module to the corresponding electrode needle device according to the detected ablation temperature. , so that the electrode needle device works within the preset ablation temperature range. | 2022 | SHANGHAI MEIWEIDA MEDICAL TECH | NA | NA | NA | China | The patent focuses on an RF ablation system with a host integrating main control, liquid cooling, temperature detection, HMI, and connectors for the electrode needle and electrode sheet. It addresses the integration burden in clinical RFA, where cooling, monitoring, and HMI are often spread across separate units. The host coordinates ablation, cooling, and live temperature reads in one device. | הפטנט מתמקד במערכת אבלציית RF עם מארח המשלב בקרה ראשית, קירור נוזלי, גילוי טמפרטורה, HMI ומחברים למחט האלקטרודה ולסדין האלקטרודה. הוא מתמודד עם נטל האינטגרציה ב-RFA קליני, שבו קירור, ניטור ו-HMI מפוזרים לעיתים קרובות על פני יחידות נפרדות. המארח מתאם אבלציה, קירור וקריאות טמפרטורה חיות במכשיר אחד. | 1 |  |  |  | 1 | [5 empty] | 1 | [10 empty] | 1 | [8 empty] | 1 |  | SHANGHAI MEIWEIDA MEDICAL TECH CO LTD | SHANGHAI MEIWEIDA MEDICAL TECH CO LTD | SHANGHAI MEIWEIDA MEDICAL TECH CO LTD | SHANGHAI MEIWEIDA MEDICAL TECH CO LTD | SHANGHAI MEIWEIDA MEDICAL TECH CO LTD | 2022-10-08 | 2022 | 2024-04-16 | 2024 | 2022-10-08 | 2022 | Application | Utility Invention | active | CN202211222360 | 304313721 | [NA] | CN
396 | EP4378393A1 | METHOD AND SYSTEM FOR DETERMINING A TEMPERATURE AND/OR TEMPERATURE DISTRIBUTION IN AN ABLATION REGION | The invention relates to a method and system for determining a temperature and/or temperature distribution (TA) in an ablation area (5), wherein the temperature and/or temperature distribution (TA) based on first and second image data sets (6) and temperatures provided for a first sensor position (S) and a second sensor position (S'). | 1. Method for determining a temperature and/or a temperature distribution (TA) in an ablation area (5), comprising the steps: - providing (100) a first image data set (6) of an object region (8), wherein the first image data set (6) comprises medical image data (7) recorded with a different spectral distribution of an X-ray radiation, wherein the object region (8) comprises the ablation region (5), - determining (200) a first sensor position (S) in the object area (8) based on the first image data set (6), - determining (300) a first density, a first density distribution, a first attenuation value, a first attenuation value distribution and/or a first tissue distribution in a surrounding area of ​​the first sensor position (S) based on the first image data set (6), - providing (400) a first temperature for the first sensor position (S), - providing (500) a second image data set (6) of the object area (8), wherein the second image data set (6) comprises medical image data (7) recorded with a different spectral distribution of an X-ray radiation, - determining (600) a second sensor position (S') in the object area (8) based on the second image data set (6), - determining (700) a second density, a second density distribution, a second attenuation value, a second attenuation value distribution and/or a second tissue distribution in a surrounding area of ​​the second sensor position (S') based on the second image data set (6), - providing (700) a second temperature for the second sensor position (S'), - Determining (1000) a temperature and/or a temperature distribution (TA)in the ablation region (5) based on the first temperature, the first sensor position (S), the second temperature, the second sensor position (S'), the first density, the first density distribution, the first attenuation value, the first attenuation value distribution, the first tissue distribution, the second density, the second density distribution, the second attenuation value, the second attenuation value distribution and/or the second tissue distribution, - Providing (1100) the determined temperature and/or temperature distribution (TA). 2. Method according to claim 1,characterized in that in the step of determining (700) a second density distribution, a second attenuation value distribution and/or a second tissue distribution, the second density distribution, the second attenuation value distribution and/or the second tissue distribution in the object region (8) is determined. 3. Method according to claim 1 or 2,characterized in that in the step of determining (300) a first density, a first density distribution, a first attenuation value, a first attenuation value distribution and/or a first tissue distribution, the first density, the first density distribution, the first attenuation value, the first attenuation value distribution and/or the first tissue distribution in the object region (8) is determined. 4. Method according to one of the preceding claims,marked by: - determining a density change, an attenuation value change and/or a tissue distribution change based on the first density, the first density distribution, the first attenuation value, the first attenuation value distribution, the first tissue distribution, the second density, the second density distribution, the second attenuation value, the second attenuation value distribution and/or the second tissue distribution, wherein the temperature and/or temperature distribution (TA) in the ablation region (5) is determined based on the first temperature, the second temperature, the density change, the attenuation value change and/or the tissue distribution change. 5. Method according to one of the preceding claims,marked by: - determining a calibration function based on the first temperature, the first sensor position (S), the second temperature, the second sensor position (S'), the first density, the first density distribution, the first attenuation value, the first attenuation value distribution, the first tissue distribution, the second density, the second density distribution, the second attenuation value, the second attenuation value distribution and/or the second tissue distribution,where the calibration function assigns a temperature to a density, an attenuation value and/or a tissue. 6. Method according to one of the preceding claims,characterized in that the provision (100, 500) of a first and/or a second image data set (7) of an object region (8) comprises an artifact removal, wherein the artifact removal removes and/or reduces image artifacts and/or metal artifacts in the medical image data (7). 7. Method according to one of the preceding claims,characterized in that the first temperature and/or the second temperature is provided by a sensor device, wherein the sensor device comprises a first sensor element (11) and a second sensor element (11), wherein the first sensor element (11) and the second sensor element (11) are spaced apart at a sensor distance, wherein the first sensor position (S) is a first sensor element position (SE1) for the first sensor element (11) and a second sensor element position (SE2) for the second sensor element (11), wherein the first temperature and/or the second temperature comprises a first temperature value for the first sensor element (11) and a second temperature value for the second sensor element (11). 8. Method according to one of the preceding claims,characterized in that the first temperature and/or the second temperature is provided by a sensor device, wherein the sensor device is comprised by an ablation device (2), wherein the ablation device comprises an ablation section (10) for thermal ablation. 9. Method according to one of the preceding claims,marked by: - providing a third image data set of the object area, wherein the third image data set (6) comprises medical image data (7) recorded with a different spectral distribution of an X-ray radiation, - determining a third temperature and/or a third temperature distribution in the ablation region (5) based on the first temperature, the first sensor position (S), the second temperature, the second sensor position (S'), the first density, the first density distribution, the first attenuation value, the first attenuation value distribution, the first tissue distribution, the second density, the second density distribution, the second attenuation value,the second attenuation value distribution and/or the second tissue distribution, - Determining a planned temperature profile in the ablation area (5) based on the first, second and third temperature and/or temperature distribution (TA), - Providing the planned temperature profile. 10. Method according to claim 9,characterized in that Ablation operating data and/or tissue parameters are provided, whereby the planned temperature profile is determined based on the ablation operating data and/or tissue parameters. 11. Method according to one of claims 9 or 10,characterized in that Based on the planned temperature profile, a planned residual ablation time is determined, whereby the planned residual ablation time is provided. 12. Method according to one of the preceding claims,characterized in that based on the first temperature, the second temperature, the third temperature, the first temperature distribution, the second temperature distribution and/or the third temperature distribution, heat loss areas in the ablation area (5) are determined, wherein the determined heat loss areas are provided. 13. System (1) for determining a temperature and/or a temperature distribution (TA) in an ablation area (5), comprising an image data provision module (20), an image data evaluation module (21), a sensor data provision module (22), a position determination module (23), a temperature determination module (24) and a provision module (25), wherein: - The image data provision module (20) is designed to provide a first image data set (6) of an object region (8), wherein the first image data set (6) comprises medical image data (7) recorded with a different spectral distribution of an X-ray radiation, wherein the object region (8) comprises the ablation region (5), - The image data evaluation module (21) is designed to determine a first density, a first density distribution, a first attenuation value, a first attenuation value distribution and/or a first tissue distribution in at least one section of the object region based on the first image data set (6), - The position determination module (23) is designed to determine a first sensor position (S) in the object area (8) based on the first image data set (6), - The sensor data provision module (22) is designed to provide a first temperature for the first sensor position (S), - The image data provision module (20) is designed to provide a second image data set (6) of the object area (8), wherein the second image data set (6) comprises medical image data (7) recorded with a different spectral distribution of an X-ray radiation, - The position determination module (23) is designed to determine a second sensor position (S') in the object area (8) based on the second image data set (6), - The sensor data provision module (22) is designed to provide a second temperature for the second sensor position (S'), - The image data evaluation module (21) is designed to determine a second density, a second density distribution, a second attenuation value, a second attenuation value distribution and/or a second tissue distribution in a surrounding area of ​​the second sensor position (S') based on the second image data set (6), - The temperature determination module (24) is designed to determine a temperature and/or a temperature distribution (TA) in the ablation area (5) based on the first temperature, the first sensor position (S), the second temperature, the second sensor position (S'), the first density, the first density distribution, the first attenuation value, the first attenuation value distribution, the first tissue distribution, the second density, the second density distribution, the second attenuation value, the second attenuation value distribution and/or the second tissue distribution, - The provision module (25) is designed to determine the specific temperature and/or temperature distribution (TA). 14. Ablation arrangement (1) comprising an ablation device (2) and a system according to claim 13, wherein the ablation device (2) has a sensor device for providing the first, second and/or third temperature and/or temperature distribution. 15. Computer program product comprising program code means for causing a system according to claim 13, an ablation arrangement (1) according to claim 14 and/or a computer to carry out the method for determining a temperature and/or temperature distribution (TA) in an ablation area (5). | 2022 | Siemens | Medical and Healthcare | Large-Size Company | https://www.siemens.com | Germany | The patent focuses on a method that determines temperature or temperature distribution in an ablation region using two image datasets and temperature readings from two sensor positions. It addresses the limitation of single-source thermometry, which can be incomplete or misleading. Combining imaging with two-point sensor data produces a richer temperature map. | הפטנט מתמקד בשיטה הקובעת טמפרטורה או התפלגות טמפרטורה באזור אבלציה באמצעות שני מערכי תמונות וקריאות טמפרטורה משני מיקומי חיישן. הוא מתמודד עם המגבלה של תרמומטריה ממקור יחיד, שיכולה להיות לא שלמה או מטעה. שילוב הדמיה עם נתוני חיישן בשתי נקודות מייצר מפת טמפרטורה עשירה יותר. |  |  |  |  | 1 |  |  |  | 1 | [14 empty] | 1 | [6 empty] | 1 |  | SIEMENS HEALTHINEERS AG | Siemens Healthineers | SIEMENS HEALTHINEERS AG | Siemens Healthineers | Siemens | 2022-11-30 | 2022 | 2024-06-05 | 2024 | 2022-11-30 | 2022 | Application | Utility Invention | active | EP20220210596 | 310674889 | [NA] | EP
397 | CN120241030A | Laser ablation method based on magnetic resonance image temperature calculation and magnetic resonance guidance | The embodiment of the invention provides a magnetic resonance image-based temperature calculation and magnetic resonance guided laser ablation method, and the method comprises the steps: periodically scanning a target region containing a focus through a magnetic resonance fast imaging sequence, and obtaining an initial magnetic resonance image sequence of a target period, the target period is any period of periodically scanning the target area; the initial magnetic resonance image sequence of the target period is corrected according to a reference period, a corrected magnetic resonance image sequence of the target period is obtained, and the reference period is a period in which a reference area containing a focus is scanned by magnetic resonance in advance to obtain a space standard image; the corrected magnetic resonance image sequence of the target period and the corrected magnetic resonance image sequence of the historical period are used to generate the temperature map of the target period, the historical period is any period before the target period, distortion generated in the rapid scanning process is eliminated, and the accuracy and convenience of temperature monitoring are improved. | 1. A method for calculating temperature based on magnetic resonance images, characterized in that the method comprises: periodically scanning a target area containing a lesion through a magnetic resonance rapid imaging sequence to obtain an initial magnetic resonance image sequence of the target cycle, wherein the target cycle is any cycle of periodic scanning of the target area; correcting the initial magnetic resonance image sequence of the target cycle according to a reference cycle to obtain a corrected magnetic resonance image sequence of the target cycle, wherein the reference cycle is a cycle of obtaining a spatial standard image in advance by using magnetic resonance to scan a reference area containing the lesion; using the corrected magnetic resonance image sequence of the target cycle and the corrected magnetic resonance image sequence of a historical cycle to generate a temperature map of the target cycle, wherein the historical cycle is any cycle before the target cycle. 2.The method according to claim 1 is characterized in that the initial magnetic resonance image sequence of the target period is corrected according to the reference period to obtain the corrected magnetic resonance image sequence of the target period, including: acquiring a spatial standard image obtained by pre-scanning; correcting the initial magnetic resonance image sequence of the target period according to the spatial standard image to obtain the corrected magnetic resonance image sequence of the target period. 3.The method according to claim 1 is characterized in that the correcting the initial magnetic resonance image sequence of the target cycle according to the reference cycle to obtain the corrected magnetic resonance image sequence of the target cycle includes: obtaining the distortion correction information of the previous cycle according to the reference cycle and the previous cycle of the target cycle; correcting the initial magnetic resonance image sequence of the target cycle using the distortion correction information to obtain the corrected magnetic resonance image sequence of the target cycle; determining and storing the distortion correction information of the target cycle based on the initial magnetic resonance image sequence of the target cycle and the spatial standard image, wherein the distortion correction information of the target cycle is used to correct the initial magnetic resonance image sequence of the next cycle of the target cycle. 4.The method according to claim 3 is characterized in that the distortion correction information of the target period is determined and stored based on the initial magnetic resonance image sequence of the target period and the spatial standard image, including: determining the first tissue position information of the target tissue in the initial magnetic resonance image sequence of the target period; determining the second tissue position information of the target tissue in the spatial standard image; calculating the spatial transformation matrix based on the first tissue position information and the second tissue position information, using the spatial transformation matrix as the distortion correction information, and storing it. 5.The method according to claim 1 is characterized in that the correcting the initial magnetic resonance image sequence of the target period according to the reference period to obtain the corrected magnetic resonance image sequence of the target period includes: selecting any period as a correction reference period; obtaining constant distortion correction information according to the reference period and the correction reference period; correcting the initial magnetic resonance image sequence of the target period using the constant distortion correction information to obtain the corrected magnetic resonance image sequence of the target period. 6.The method according to claim 5 is characterized in that the constant distortion correction information is obtained according to the reference period and the correction base period, comprising: when the target period is earlier than the correction base period, not correcting the initial magnetic resonance image sequence of the target period, continuing to scan the next period until the target period is the correction base period, and scanning to obtain the initial magnetic resonance image sequence of the correction base period; acquiring the spatial standard image of the reference period, and obtaining the constant distortion correction information according to the spatial standard image and the correction base period. An initial magnetic resonance image sequence is used to obtain constant distortion correction information of the correction reference period. 7.The method according to claim 1 is characterized in that the use of the corrected magnetic resonance image sequence of the target cycle and the corrected magnetic resonance image sequence of the historical cycle to generate the temperature map of the target cycle includes: establishing a temperature difference map using the phase difference values of the corrected magnetic resonance image sequence of the target cycle and the corrected magnetic resonance image sequence of the historical cycle; generating the temperature map of the target cycle based on the temperature difference map and a baseline temperature map, wherein the baseline temperature map is the temperature map of the historical cycle. 8.The method according to any one of claims 1 to 7 is characterized in that the spatial standard image is obtained by scanning the reference area based on a spin echo sequence; the initial magnetic resonance image sequence of the target period is obtained by scanning the target area based on gradient echo imaging. 9.The method according to any one of claims 1 to 7 is characterized in that the periodic scanning of the target area containing the lesion by means of a rapid magnetic resonance imaging sequence to obtain an initial magnetic resonance image sequence of the target period includes: determining the target layer and/or the target sampling column within the layer to be sampled in the target period according to a set downsampling rule; scanning the target layer and/or the target sampling column within the layer by means of the rapid magnetic resonance imaging sequence to obtain a downsampling sequence of the target period; and completing and reconstructing the downsampling sequence by means of a reconstruction algorithm to obtain an initial magnetic resonance image sequence of the target period. 10. A method for calculating temperature based on magnetic resonance images, characterized in that the method includes: by setting a scanning mode, adopting a downsampling strategy to scan a target area containing a lesion, and obtaining a downsampling sequence of a target cycle, wherein the target cycle is any cycle of periodic scanning of the target area, the downsampling strategy is determined based on the set scanning mode, and the downsampling sequence of the target cycle contains image information of at least one optical fiber; reconstructing a complete magnetic resonance image sequence according to the downsampling sequence of the target cycle to obtain a reconstructed magnetic resonance image sequence of the target cycle; acquiring a reconstructed magnetic resonance image sequence of a historical cycle, and generating a temperature map of the target cycle based on the reconstructed magnetic resonance image sequence of the target cycle and the reconstructed magnetic resonance image sequence of the historical cycle, wherein the historical cycle is any cycle before the target cycle, and the temperature map is used to provide information support for the laser ablation process of the at least one optical fiber. 11.The method according to claim 10 is characterized in that the setting of the scanning mode is layer scanning, and the downsampling strategy is inter-layer downsampling; the setting of the scanning mode and the adoption of the downsampling strategy to scan the target area containing the lesion to obtain the downsampling sequence of the target period include: determining the target layer to be sampled in the target period according to the set inter-layer downsampling rule; scanning the target layer by the set scanning mode to obtain the downsampling sequence of the target period. 12.The method according to claim 11 is characterized in that the set inter-layer downsampling rule is to sample odd layers and even layers respectively at intervals; determining the target layer to be sampled in the target cycle according to the set inter-layer downsampling rule comprises: determining the sampling layer scanned in the previous cycle; when the sampling layer is an odd layer, determining the target layer to be sampled in the target cycle to be an even layer; when the sampling layer is an even layer, determining the target layer to be sampled in the target cycle to be an odd layer. 13. The method according to claim 11 is characterized in that the set inter-layer downsampling rule is the layer to be sampled When the number of layers to be sampled exceeds the layer number threshold, the first strategy is adopted, and when the number of layers to be sampled does not exceed the layer number threshold, the second strategy is adopted; the target layer to be sampled in the target cycle is determined according to the set inter-layer downsampling rule, including: determining the number of layers to be sampled in the target area; when the number of layers to be sampled exceeds the layer number threshold, determining the target layer to be sampled in the target cycle according to the first strategy; When the number of layers to be sampled does not exceed the layer number threshold, the target layers to be sampled in the target cycle are determined according to the second strategy. 14.The method according to claim 10 is characterized in that the set scanning mode is column scanning, and the downsampling strategy is inter-layer downsampling and/or intra-layer downsampling; the downsampling strategy is used to scan the target area containing the lesion by setting the scanning mode to obtain a downsampling sequence of the target period, including: determining the target layer and/or the target sampling column within the layer to be sampled in the target period according to the set inter-layer downsampling rules and/or intra-layer downsampling rules; scanning the target layer and/or the target sampling column within the layer by the set scanning mode to obtain the downsampling sequence of the target period. 15.A magnetic resonance-guided laser ablation method, characterized in that the method comprises: obtaining a temperature map of a target period, wherein the temperature map of the target period is obtained by the method for calculating temperature based on a magnetic resonance image as described in any one of claims 1-9 or 1014; generating an ablation map according to the temperature map of the target period, wherein the ablation map is used to provide information support for the laser ablation process of at least one optical fiber. 16.A magnetic resonance-guided laser ablation system, characterized in that the system includes a host, and the host is provided with an application for implementing the method for calculating temperature based on magnetic resonance images as described in any one of claims 1-9 or 10-14, or the magnetic resonance-guided laser ablation method as described in claim 15. 17. A computer-readable storage medium, characterized in that the computer-readable storage medium is arranged in a host of a magnetic resonance-guided laser ablation system, and stores computer-executable instructions, which, when executed by a processor, implement the method for calculating temperature based on magnetic resonance images as described in any one of claims 1-9 or 10-14, or the magnetic resonance-guided laser ablation method as described in claim 15. | 2023 | SINOVATION BEIJING MEDICAL EQUIPMENT | Medical and Healthcare | NA | http://www.sinovationmed.com/?lang=en | China | The patent focuses on an MR-guided laser ablation method that uses a fast MR imaging sequence to scan the lesion repeatedly and computes temperature from successive MR images. It addresses the gap that MR-guided laser interstitial thermotherapy (MRgLITT) lacks an efficient way to derive temperature from MR images for real-time control. MR-derived thermometry then guides laser ablation in real time. | הפטנט מתמקד בשיטת אבלציית לייזר מונחית MR המשתמשת ברצף הדמיית MR מהיר לסריקה חוזרת של הנגע ומחשבת טמפרטורה מתמונות MR עוקבות. הוא מתמודד עם הפער שתרמותרפיה אינטרסטיציאלית בלייזר מונחית MR (MRgLITT) חסרה דרך יעילה לגזור טמפרטורה מתמונות MR לבקרה בזמן אמת. תרמומטריה מבוססת MR ואז מנחה אבלציית לייזר בזמן אמת. |  |  |  |  | 1 |  |  |  | 1 |  |  |  |  | 1 | [6 empty] | 1 | [9 empty] | 1 |  | SINOVATION BEIJING MEDICAL EQUIPMENT CO LTD | SINOVATION BEIJING MEDICAL EQUIPMENT CO LTD | SINOVATION BEIJING MEDICAL EQUIPMENT CO LTD | SINOVATION BEIJING MEDICAL EQUIPMENT CO LTD | SINOVATION BEIJING MEDICAL EQUIPMENT CO LTD | 2023-12-27 | 2023 | 2025-07-04 | 2025 | 2023-12-27 | 2023 | Application | Utility Invention | active | CN202311828361 | 322187829 | [NA] | CN
398 | US20260083490A1 | METHOD AND DEVICE FOR ABLATING ANATOMICAL TARGETS FROM WITHIN A BLOOD VESSEL WITH IN-PROCEDURE FEEDBACK | A cryo-ablation system treats conditions associated with sympathetic nervous system hyperactivity using a catheter with a configurable heat transfer element that transitions from a flat delivery profile to deployed configurations including helical shapes. The system includes deployable temperature probes that directly measure tissue temperature at target locations inside or outside blood vessels. A computational control system processes temperature data from multiple sensors and deployable probes to calculate real-time lethal isotherm progression using heat transfer modeling algorithms. The system provides vessel occlusion during treatment and real-time feedback for treatment optimization. Methods include introducing the catheter into a renal artery, deploying the heat transfer element, occluding the vessel, deploying temperature probes to target locations within or outside the vessel, delivering cryogenic temperatures while monitoring probe temperatures, and terminating treatment based on computational analysis of temperature data to achieve target ablation depths. | 1. A cryo-ablation system for treating resistant hypertension, comprising: a cryo-ablation balloon configured to be placed within a blood vessel; a deployable temperature probe outside the cryo-ablation catheter, to perform direct temperature measurement of tissue, at a known distance from the balloon. 2. The cryo-ablation catheter of claim 1, including an occlusive cryo-ablation balloon configured to be placed within a blood vessel; the occlusive cryo-ablation balloon being selectively transitionable from a substantially flat configuration to a cylindrical, conical, or helical shape for creating a continuous, circumferential, or helical elongated cryogenic treatment region of at least 5 mm in length; a temperature monitoring system configured to measure temperatures at the temperature probe; and a pressure sensing system to detect arterial occlusion distal to the cryo-ablation balloon. 3. The cryo-ablation system of claim 1, wherein the catheter is an over-the-wire, rapid exchange, deflectable, or shapable configuration. 4. The cryo-ablation system of claim 1, further comprising a shapeable structure including a metallic core wire encapsulated in an atraumatic polymer body, allowing manual shaping of the catheter at a desired angle or curvature. 5. The cryo-ablation system of claim 1, further including a compliant occlusion balloon located on the catheter shaft proximal to the cryo-ablation balloon, wherein the occlusion balloon is configured to be manually inflated by the operator using a saline or contrast solution. 6. The cryo-ablation system of claim 1, wherein the cryo-ablation balloon is configured to transition to a helical shape. 7. The cryo-ablation system of claim 6, further comprising means to control the pitch, outer diameter, and length of the cryo-ablation balloon when deployed in the helical shape. 8. The cryo-ablation system of claim 1, further comprising an ultrasound imaging system to visualize progression of tissue freezing, measure a diameter of frozen tissue volume, and calculate a position of a lethal isotherm to provide real-time feedback during a cryo-ablation procedure. 9. The cryo-ablation system of claim 2, wherein the temperature monitoring system includes one or more temperature probes, at least one of which is situated outside the cryo-ablation balloon at a known distance to compute the distance of a lethal isotherm from the balloon surface. 10. The cryo-ablation system of claim 9, wherein the temperature monitoring system is configured to display the distance of the lethal isotherm as a guide during the cryo-ablation procedure. 11. The cryo-ablation system of claim 2, wherein the temperature monitoring system includes one or more deployable temperature probes that are advanceable through a vessel wall 1-10 mm. 12. The cryo-ablation system of claim 1, wherein the cryo-ablation balloon has a deployed diameter sufficient to completely occlude the blood vessel. 13. A method for treating resistant hypertension, comprising: introducing a cryo-ablation catheter with an occlusive cryo-ablation balloon into a renal artery; transitioning the cryo-ablation balloon from a flat configuration to a cylindrical, conical, or helical configuration within the renal artery to create a continuous cryogenic treatment region; occluding the renal artery using the cryo-ablation balloon to reduce blood flow and increase the efficacy of the cryo-ablation; achieving cryogenic temperatures to ablate renal sympathetic nerves adjacent to the renal artery; measuring the temperature at a known distances outside the cryo-ablation balloon to determine tissue freezing; and monitoring progression of tissue freezing at the known distances and adjusting cryo-ablation parameters based on real-time feedback from the temperature measurements. 14. The method of claim 13, further comprising using ultrasound to monitor progression of tissue freezing and using a pressure sensing system to confirm arterial occlusion distal to the cryo-ablation balloon. 15. The method of claim 14, further comprising manually shaping the cryo-ablation catheter to navigate and engage the distal branches of the renal artery. 16. The method of claim 14, further comprising the step of inflating a compliant occlusion balloon proximal to the cryo-ablation balloon to achieve vessel occlusion if the cryo-ablation balloon does not fully occlude the renal artery. 17. The method of claim 14, further comprising the step of computing a distance of a lethal isotherm from the cryo-ablation balloon and displaying this information to an operator for guiding duration and extent of cryo-ablation. 18. The method of claim 13, wherein cryo-ablation is guided by an ultrasound imaging externally or within the renal artery to visualize the progression of tissue freezing and provide real-time feedback for the procedure. 19. The method of claim 13, wherein the cryo-ablation is performed to treat medical conditions selected from the group consisting of resistant hypertension, heart failure, central sympathetic hyperactivity, metabolic syndrome, and obesity by cryo-ablating targets proximate to the renal artery. | 2024 | STEER MEDICAL | Medical and Healthcare | Start Up | http://www.steermedical.com | Canada | The patent focuses on a cryoablation catheter for sympathetic-nervous-system disorders (e.g., resistant hypertension), with a heat-transfer element that goes from a flat delivery profile to deployed shapes (including helical), plus deployable temperature probes. It addresses the difficulty of safely cryoablating perivascular sympathetic nerves with adequate in-procedure feedback. The probes measure temperature inside or outside the vessel for in-procedure feedback. | הפטנט מתמקד בקטטר קריו-אבלציה להפרעות מערכת העצבים הסימפתטית (למשל יתר לחץ דם עמיד), עם אלמנט העברת חום העובר מפרופיל מסירה שטוח לצורות פרושות (כולל ספירלית), בתוספת גששי טמפרטורה הניתנים לפריסה. הוא מתמודד עם הקושי בקריו-אבלציה בטוחה של עצבים סימפתטיים פריוסקולריים עם היזון חוזר תוך-פרוצדורלי מספק. הגששים מודדים טמפרטורה בתוך או מחוץ לכלי הדם להיזון חוזר תוך-פרוצדורלי. |  |  | 1 |  | 1 | [9 empty] | 1 | [10 empty] | 1 | [6 empty] | STEER MEDICAL INC | STEER MEDICAL INC | STEER MEDICAL INC | STEER MEDICAL INC | STEER MEDICAL INC | 2025-09-23 | 2025 | 2026-03-26 | 2026 | 2024-09-24 | 2024 | Application | Utility Invention | active | US202519337355 | 326960204 | [NA] | US
399 | US20260069350A1 | Laparoscopic Applicator, and System for Laparoscopic Ablation | The present invention describes a device for use in laparoscopic surgery, and a system for laparoscopic ablation. A method for perivascular and/or perineural neuromodulation is also disclosed. | 1. Device for use in laparoscopic surgery, having a shaft, which has a distal end and a proximal end, a holding unit, which is arranged at the distal end of the shaft and is adapted to hold and/or to release a tissue, at least two electrodes integrated in the holding unit, a force transmission unit, which is arranged on the shaft, is connected to the holding unit and is adapted to be moved relative to the shaft, an electric cable, which is arranged in the shaft, is connected to the at least two electrodes and is adapted to transmit an electrical signal that is to be received via the proximal end of the shaft to the at least two electrodes. 2. Device according to claim 1, having an articulated arrangement, which is arranged between the holding unit and the distal end of the shaft and is adapted to connect the distal end of the shaft, in particular movably, to the holding unit, wherein a movement of the force transmission unit leads to a movement of the holding unit and/or to the tissue being held and/or released by means of the holding unit. 3. Device according to claim 1, wherein the holding unit is configured as a finger gripper, in particular a three-fingered gripper, and the force transmission unit is configured as a pull- and push-rod, wherein a movement of the pull- and push-rod leads to the opening or closing of the finger gripper in order to grip or release the tissue. 4. Device according to claim 3, wherein fingers of the finger gripper each have an inner face oriented, in particular when the finger gripper is in a closed state, towards the tissue, in which inner face there is integrated at least one of the at least two electrodes, in particular one of a plurality of electrodes. 5. Device according to claim 1, wherein the holding unit is configured as a suction unit, wherein a movement of the force transmission unit leads to an angling of the suction unit relative to the shaft, and/or the suction unit is adapted to suction a tissue or to detach itself from the tissue. 6. Device according to claim 5, wherein the suction unit has: a suction cup, at least one suction hole, which is arranged in the suction cup, a suction channel, which is connected to the at least one suction hole, and/or a suction connection provided at the end of the suction channel, via which an external vacuum pump can be connected in order to generate a negative pressure in the suction cup and suction the tissue against the suction cup. 7. Device according to claim 6, wherein the suction cup is configured to be oval, in particular stadium-shaped or circular, along a vertical axis, and the at least two electrodes are arranged in the suction cup along the vertical axis and/or transverse to the vertical axis. 8. Device according to claim 6, wherein the suction cup is configured to be oval, in particular stadium-shaped or circular, along a vertical axis, and at least three electrodes are arranged in the suction cup along the vertical axis and/or transverse to the vertical axis, in particular spaced apart from the at least two electrodes. 9. Device according to claim 1, wherein the holding unit is configured as a clamshell gripper, in particular a double half-clamshell gripper, wherein a movement of the force transmission unit leads to an angling of the clamshell gripper relative to the shaft, and/or to the opening and/or to the closing of the clamshell gripper. 10. Device according to claim 9, wherein the clamshell gripper extends along a transverse axis and is adapted to open or to close along a vertical axis, wherein the at least two electrodes are integrated in an inside of at least one clamshell, in particular along a longitudinal axis. 11. Device according to claim 9, wherein the clamshell gripper extends along a transverse axis and is adapted to open or to close along a vertical axis, wherein at least three electrodes are integrated in an inside of at least one clamshell, in particular along and/or transverse to a longitudinal axis, spaced apart from the at least two electrodes. 12. Device according to claim 9, wherein the force transmission unit is configured as a first pull- and push-rod, which is arranged in the shaft and is connected to the articulated arrangement, wherein a movement of the first pull- and push-rod leads to the opening or closing of the clamshell gripper, and/or the force transmission unit is configured as a second pull- and push-rod, which is arranged in the shaft and is connected, in particular solely, to the clamshell gripper, wherein a movement of the second pull- and push-rod leads to an angling of the clamshell gripper relative to the shaft. 13. Device according to claim 1, wherein the force transmission unit and the holding unit are configured as a contiguous shape memory material, in particular shape memory alloy, and the holding unit forms a trap structure, in particular a trap basket, wherein a movement of the force transmission unit leads to the trap structure, in particular the trap basket, moving back into the shaft and folding, or leads to the trap structure, in particular the trap basket, moving out of the shaft and unfolding. 14. Device according to claim 13, wherein the trap structure has two individual and three contiguous splines, wherein the at least two electrodes and/or at least three electrodes are integrated in the splines, in particular in the contiguous splines. 15. System for laparoscopic ablation, having: a device according to claim 1; a signal generator arrangement which is connected or can be connected to the device; and a control and evaluation unit which is connected or can be connected to the device and/or to the signal generator arrangement. 16. System according to claim 15, wherein the signal generator arrangement has a first signal generator for generating a radiofrequency, RF, signal and a second signal generator for generating a signal for pulsed field ablation. 17. System according to claim 15, wherein the control and evaluation unit is configured to activate the first signal generator and/or the second signal generator to deliver a signal. 18. System according to claim 17, wherein the control and evaluation unit is configured to activate the first signal generator and/or the second signal generator to deliver a signal in dependence on at least one electrode temperature. 19. System according to claim 18, wherein the control and evaluation unit is configured to activate the first signal generator to deliver a signal if the electrode temperature falls below a temperature limit value and to activate the second signal generator to deliver a signal if the electrode temperature assumes or exceeds the temperature limit value. 20. System according to claim 15, wherein the control and evaluation unit is adapted to apply an electric current to the at least two electrodes and/or to measure a voltage via the at least three electrodes, wherein a nerve activity of the tissue and/or the local tissue impedance can be determined or is determined from the electric current and the electric voltage. 21. Method for perivascular and/or perineural neuromodulation, comprising: inserting an applicator into a patient; positioning a distal end of the applicator at a vessel or tissue or nerve of the patient such that at least two electrodes of the applicator are in contact perivascularly and/or perineurally with a perimeter of the vessel or tissue or nerve or lie against a perimeter of the vessel or tissue or nerve; carrying out a denervation by delivering energy via the at least two electrodes, wherein the delivery of energy is effected on the basis of a protocol which has at least one pulse of a pulsed field ablation (PFA); withdrawing the applicator. 22. Method according to claim 21, wherein an irreversible electroporation is carried out. 23. Method according to claim 21, wherein a reversible electroporation is carried out. 24. Method according to claim 21, wherein the applicator is inserted and/or withdrawn by a laparoscopic procedure. 25. Method according to claim 22, wherein the laparoscopic applicator is inserted into a patient's torso by means of a trocar. 26. Method according to claim 21, wherein the applicator is inserted and/or withdrawn by means of an open surgical procedure. 27. Method according to claim 21, wherein, after the applicator has been inserted, geometric unfolding of a distal end of the applicator is initiated and a form of the distal end that conforms to a vessel or tissue or nerve is thus achieved. 28. Method according to claim 21, wherein the electrodes of the applicator are brought into contact with a target vessel or target tissue or target nerve. 29. Method according to claim 21, further comprising: checking a positioning of the applicator by means of a local impedance measurement. 30. Method according to claim 21, wherein the protocol comprises only PFA pulses. 31. Method according to claim 21, wherein the protocol comprises a combination of PFA pulses and radiofrequency (RF) pulses. 32. Method according to claim 21, further comprising: choosing whether the delivery of energy for the denervation is to be carried out on the basis of a protocol by means of PFA pulses or on the basis of a protocol by means of a combination of PFA pulses and RF pulses. 33. Method according to claim 21, wherein a delivery of energy is effected once or multiple times at the same or different locations of the vessel or tissue or nerve. 34. Method according to claim 21, wherein perivascular and/or perineural neuromodulation occurs by carrying out the denervation. 35. Method according to claim 21, further comprising: before the denervation is carried out, quantifying a conductivity of the vessel or tissue or nerve by means of at least two stimulation and measuring electrodes. 36. Method according to claim 21, further comprising: after the denervation has been carried out, characterising the denervation by means of at least two stimulation and measuring electrodes. 37. Method according to claim 21, further comprising: measuring a temperature at a chosen ablation site. | 2024 | STOCKERT | Medical and Healthcare | Small-size company | http://www.stockert.de | Germany | The patent focuses on a laparoscopic applicator and system for laparoscopic ablation, with a holding unit at the distal shaft that grips tissue and integrates electrodes for energy delivery. It addresses the limitations of conventional laparoscopic devices that lack integrated tissue holding plus electrical neuromodulation. The device targets nerves around vessels via a laparoscopic approach with combined grip and energy. | הפטנט מתמקד באפליקטור ומערכת לפרוסקופיים לאבלציה לפרוסקופית, עם יחידת אחיזה בפיר הדיסטלי האוחזת ברקמה ומשלבת אלקטרודות למסירת אנרגיה. הוא מתמודד עם המגבלות של מכשירים לפרוסקופיים קונבנציונליים החסרים אחיזת רקמה משולבת בתוספת נוירומודולציה חשמלית. המכשיר ממקד עצבים סביב כלי דם דרך גישה לפרוסקופית עם אחיזה ואנרגיה משולבות. | 1 |  |  |  | 1 | [6 empty] | 1 | [17 empty] | 1 |  |  | STOCKERT GMBH | STOCKERT GMBH | STOCKERT GMBH | STOCKERT GMBH | STOCKERT GMBH | 2024-11-12 | 2024 | 2026-03-12 | 2026 | 2024-09-11 | 2024 | Application | Utility Invention | active | US202418944629 | 326727139 | CN121647795A¦DE102024126105A1¦EP4710881A1¦JP2026052677A | US
400 | CN115802962A | Ultrasonic surgical aspirator for probing and ablating tissue | Systems and methods for probing a patient's problem using an ultrasound surgical aspirator that includes a tip for treating tissue of a patient and a console coupled to a handpiece. The console provides an AC drive signal to the ultrasound handpiece that vibrates the tip against the patient tissue. The AC drive signal includes a first component having a resonant frequency of the ultrasonic handpiece and a second component having a probing frequency less than the resonant frequency. The console measures a voltage and a current of the AC drive signal, calculates a resistance associated with the ultrasound handpiece based on the measured voltage and current, and then provides at least one of an audible, visual, or tactile indication based on the calculated resistance. | 1. An ultrasonic tool system for probing patient tissue, the system comprising: an ultrasonic handpiece comprising a tip for treating a distal region of patient tissue and at least one driver, the tip being coupled to the driver and an AC drive signal is applied to the driver to vibrate the tip, the ultrasonic handpiece defines a first passage for providing suction at a distal region of the tip and a first passage for supplying fluid to the distal region of the tip. a second pathway; and a console coupled to the ultrasonic handpiece and configured to generate said AC drive signal applied to said at least one driver of the ultrasonic handpiece to vibrate the tip of the ultrasonic handpiece, said The console includes: a first sensor for measuring a voltage of the AC drive signal; a second sensor for measuring a current of the AC drive signal; and a processor coupled to the first and second sensors and configured to: providing said AC drive signal to said at least one driver of an ultrasonic handpiece, said AC drive signal comprising a first component having a resonant frequency of the ultrasonic handpiece and a second component having a probe frequency less than said resonant frequency; Measure voltage and current of the AC drive signal using the first and second sensors; calculate a resistance associated with the ultrasonic handpiece based on the measured voltage and measured current; and provide an audible, visual, or tactile sense based on the calculated resistance at least one of the instructions. 2. The system of claim 1 , wherein the processor is configured to provide at least one of an auditory, visual, or tactile indication based on the calculated resistance by being configured to: identifying a property of the patient tissue; and providing at least one of an audible, visual, or tactile indication of the identified property. 3. The system of claim 1 or 2, wherein the AC drive signal provided to the ultrasound handpiece is defined by a base signal having a resonant frequency amplitude modulated according to the probe frequency. 4. The system of any one of claims 1-3, wherein the resonant frequency is about 25 kHz and the probing frequency is about 4 Hz. 5. The system of any one of claims 1-4, wherein the AC drive signal provided to the ultrasonic handpiece is configured to induce vibration of the tip insufficient to ablate patient tissue. 6. The system of claim 5 , wherein the AC drive signal provided to the ultrasonic handpiece is controlled by being configured to induce vibrations with a peak-to-peak displacement of less than or equal to 100 microns at the distal region of the tip. Configured to induce vibration of the tip insufficient to ablate patient tissue. 7. The system of any one of claims 1-6, wherein the AC drive signal is defined as a first AC drive signal, and the processor is configured to: determine that the ultrasonic tool system is set whether to operate in an exploratory mode or an ablation mode; in response to determining that the ultrasonic tool system is configured to operate in an exploratory mode, providing a first AC drive signal to an ultrasonic handpiece; and in response to determining that the ultrasonic tool system is configured to Operating in the ablation mode, a second AC drive signal is provided to the ultrasonic handpiece, the second AC drive signal being configured to induce vibrations of the tip sufficient to ablate tissue of the patient. 8. The system of claim 7 , wherein the second AC drive signal provided to the ultrasonic handpiece is configured to induce a peak-to-peak displacement greater than 100 microns and less than or equal to 300 microns at the distal region of the tip. Micron vibrations configured to induce vibrations of the tip sufficient to ablate tissue of the patient. 9. The system of claim 7 or 8, further comprising a switch communicatively coupled to the processor, the switch having a first setting and a second setting, wherein the processor is configured to: determine that the ultrasound tool system is set to operate in a probe mode in response to the switch being set to the first setting; and in response to the switch being set To the second setting, it is determined that the ultrasound tool system is set to operate in an ablation mode. 10. The system of any one of claims 7-9, wherein the processor is configured to, in response to determining that the ultrasound tool system is configured to operate in an ablation mode: through a first pathway defined by the ultrasound handpiece Suction is provided at the distal region of the tip; and fluid is supplied to the distal region of the tip through a second channel defined by the ultrasonic handpiece. 11. The system according to any one of claims 1-9, wherein the processor, by being configured to perform Based on the measured voltage and measured current, the equivalent current through the mechanical part of the ultrasonic handpiece is calculated; and based on the calculated equivalent current through the mechanical part of the ultrasonic handpiece, the calculation and ultrasonic The resistance associated with the handpiece. 12. The system of claim 11 , wherein the processor, by being configured to perform the following, is configured to calculate an ultrasonic handpiece based on the calculated equivalent current through a mechanical component of the ultrasonic handpiece resistance associated with the piece: Calculate the first amplitude of the measured voltage at the probe frequency, the second amplitude of the calculated equivalent current through the mechanical part of the ultrasonic handpiece at the probe frequency, and the second amplitude at the probe frequency the phase difference between the measured voltage and the calculated equivalent current through the mechanical components of the ultrasonic handpiece; and based on the calculated first amplitude, the calculated second amplitude, and the calculated phase difference, calculate the ultrasonic The real part of the impedance of the handpiece. 13. The system according to any one of claims 1-12, wherein the processor is configured to provide an audible, visual or tactile indication based on the calculated resistance by being configured to perform At least one of: calculating a difference between the calculated resistance and the no-load resistance of the ultrasonic handpiece; and providing at least one of an audible, visual, or tactile indication based on the calculated difference. 14. The system of claim 13 , wherein the processor is configured to provide at least one of an auditory, visual, or tactile indication based on the calculated difference by being configured to: The identified difference identifies a property of the patient tissue; and provides at least one of an audible, visual, or tactile indication of the identified property. 15. The system of any one of claims 1-14, wherein the console further comprises a memory coupled to the processor that stores tissue property data indicative of potential tissue properties, and for potential tissue One or more values specific to the potential tissue property for each of the potential tissue properties, and providing at least one of an auditory, visual or tactile indication based on the calculated resistance comprises: the one or more values of the property and the calculated resistance, identifying one of the potential tissue properties indicated by the tissue property data as a property of the patient tissue; and providing an audible sense of the identified property of the patient tissue , at least one of visual or tactile indications. 16. The system of claim 15 , wherein the processor is configured based on the one or more values specific to the underlying tissue property and the calculated resistance by being configured to perform One of the potential tissue properties indicated by the tissue property data is identified: calculating a difference between the calculated resistance and the no-load resistance of the ultrasonic handpiece; and identifying a potential tissue property indicated by the tissue property data based on the one or more values specific to the potential tissue property and the calculated difference The one underlying tissue property in . 17. The system of claim 13 , 14 , or 16 , wherein the processor is configured to determine the absence of the ultrasound handpiece by being configured to perform the following in response to the ultrasound handpiece being connected to the console. Load resistance: with the ultrasonic handpiece in an unloaded state, an AC drive signal is provided to the ultrasonic handpiece; with the ultrasonic handpiece in an unloaded state, measured using the first and second sensors a second voltage and a second current of the AC drive signal; and calculating an unloaded resistance of the ultrasonic handpiece based on the measured second voltage of the AC drive signal and the measured second current of the AC drive signal. 18. The system of claim 13 , 14 , or 16 , wherein the processor, by being configured to read from a memory integrated with the ultrasound handpiece in response to the ultrasound handpiece being connected to the console, indicates the The data of the no-load resistance is configured to determine the no-load resistance of the ultrasonic handpiece. 19. A method of probing patient tissue using an ultrasonic tool system comprising an ultrasonic handpiece having a tip for treating tissue of a patient and at least one driver, the tip being coupled to the driver and An AC drive signal is applied to the driver to vibrate the tip, the method comprising: supplying fluid to a distal region of the tip through at least a portion of the ultrasonic handpiece; at the distal region of the tip through at least a portion of the ultrasonic handpiece providing suction; providing an AC drive signal to the ultrasonic handpiece, the AC drive signal comprising a first component having a resonant frequency of the ultrasonic handpiece and a second component having a probing frequency less than the resonant frequency; measuring the AC drive at the tip a voltage and a current of the signal; calculating a resistance associated with the ultrasonic handpiece based on the measured voltage and the measured current; and providing at least one of an audible, visual, or tactile indication based on the calculated resistance. 20. The method of claim 19 , wherein providing at least one of an audible, visual, or tactile indication based on the calculated resistance comprises: identifying a property of the patient's tissue based on the calculated resistance; and providing an indication of the identified property. At least one of auditory, visual or tactile indications. 21. twenty one. The method of claim 19 or 20, wherein providing the AC drive signal to the ultrasound handpiece comprises: generating a base signal having a resonant frequency; and modulating the amplitude of the base signal in accordance with the probe frequency. 22. twenty two. The method of any one of claims 19-21, wherein the resonant frequency is about 25 kHz and the probing frequency is about 4 Hz. 23. twenty three. The method of any one of claims 19-22, wherein providing an AC drive signal to the ultrasonic handpiece comprises providing an AC drive signal to the ultrasonic handpiece that induces vibrations of the tip that are insufficient to ablate patient tissue. 24. twenty four. The method of claim 23 , wherein providing an AC drive signal to the ultrasonic handpiece that induces vibrations of the tip that are insufficient to ablate patient tissue comprises: providing the ultrasonic handpiece with an AC drive signal that induces peak-to-peak vibration at a distal region of the tip. AC drive signal for vibration with displacement less than or equal to 100 microns. 25. The method according to any one of claims 19-24, further comprising: determining that the ultrasound tool system is set to operate in a probing mode; In response to determining that the ultrasound tool system is configured to operate in a probing mode, providing an AC drive signal to the ultrasound handpiece; determining that the ultrasound tool system is configured to operate in an ablation mode; and in response to determining that the ultrasound tool system is configured to operate in an ablation mode, Another AC drive signal is provided to the ultrasonic handpiece that induces vibrations of the tip sufficient to ablate the patient's tissue. 26. The method of claim 25 , wherein providing the another AC drive signal to the ultrasonic handpiece to induce vibrations of the tip sufficient to ablate the patient's tissue comprises providing the ultrasonic handpiece to induce a peak at a distal region of the tip. - A vibrating AC drive signal with a peak displacement greater than 100 microns and less than or equal to 300 microns. 27. The method of claim 25 or 26, wherein the ultrasonic tool system further comprises a switch having a first setting and a second setting, and the method further comprises: monitoring a state of the switch to determine whether the switch is set to Whether the first setting or the second setting; determining that the switch is set to the first setting; in response to determining that the switch is set to the first setting, determining that the ultrasound tool system is set to operate in a probe mode; determining that the switch is set to the second setting; responding Upon determining that the switch is set to the second setting, it is determined that the ultrasound tool system is set to operate in the ablation mode. 28. The method of any one of claims 25-27, further comprising, in response to determining that the ultrasonic tool system is configured to operate in an ablation mode, supplying fluid to a distal region of the tip through at least a portion of the ultrasonic handpiece and Suction is provided at a distal region of the tip through at least a portion of the ultrasonic handpiece. 29. The method of any one of claims 19-28, wherein calculating a resistance associated with the ultrasonic handpiece based on the measured voltage and the measured current comprises: calculating based on the measured voltage and the measured current An equivalent current through a mechanical component of the ultrasonic handpiece; based on the calculated equivalent current through a mechanical component of the ultrasonic handpiece, a resistance associated with the ultrasonic handpiece is calculated. 30. The method of claim 29 , wherein calculating the resistance associated with the ultrasonic handpiece based on the calculated equivalent current through a mechanical component of the ultrasonic handpiece comprises: One amplitude, the calculated equivalent current through the mechanical part of the ultrasonic handpiece, the second amplitude at the probe frequency, and the difference between the measured voltage and the calculated equivalent current through the mechanical part of the ultrasonic handpiece at the probe frequency a phase difference at a frequency; and calculating a real part of an impedance of the ultrasonic handpiece based on the calculated first amplitude, the calculated second amplitude, and the calculated phase difference. 31. The method of any one of claims 19-30, wherein providing at least one of an audible, visual, or tactile indication based on the calculated resistance comprises: calculating the calculated resistance and the no-load resistance of the ultrasonic handpiece providing at least one of an auditory, visual or tactile indication based on the calculated difference. 32. The method of claim 31 , wherein providing at least one of an auditory, visual, or tactile indication based on the calculated difference comprises: identifying a property of the patient tissue based on the calculated difference; providing an auditory indication of the identified property. , at least one of visual or tactile indications. 33. The method according to any one of claims 19-32, wherein the ultrasound tool system further comprises a memory storing tissue property data indicating potential tissue properties, and for one of the potential tissue properties One or more values specific to each potential tissue property, and providing at least one of an audible, visual, or tactile indication based on the calculated resistance comprises: based on all values specific to the potential tissue property The one or more values and the calculated resistance identify one of the potential tissue properties indicated by the tissue property data as a property of the patient tissue; and provide an audible, visual, or at least one of the tactile indications. 34. The method of claim 33 , wherein identifying one of the potential tissue properties indicated by the tissue property data based on the one or more values specific to the potential tissue property and the calculated resistance comprises: calculating a difference between the calculated resistance and an unloaded resistance of the ultrasound handpiece; and identifying a potential tissue property indicated by the tissue property data based on the one or more values specific to the potential tissue property and the calculated difference The one underlying tissue property in . 35. The method of claim 31 , 32 or 34, further comprising determining the no-load resistance of the ultrasonic handpiece by: positioning the ultrasonic handpiece in an unloaded state; where the ultrasonic handpiece is positioned in an unloaded state Bottom: providing an AC drive signal to the ultrasonic handpiece; measuring a second voltage and a second current of the AC drive signal provided to the ultrasonic handpiece; and measuring a second voltage and a measured second current based on the AC drive signal Calculate the no-load resistance of the ultrasound handpiece. 36. The method of claim 31 , 32 or 34, further comprising determining the unloaded resistance of the ultrasound handpiece by reading data indicative of the unloaded resistance from a memory integrated with the ultrasound handpiece. 37. An ultrasonic tool system for probing patient tissue, the system comprising: an ultrasonic handpiece comprising a tip for treating a distal region of patient tissue and at least one driver, the tip being coupled to the driver and An AC drive signal is applied to the driver to vibrate the tip, the ultrasonic handpiece defines a first passage for providing suction at a distal region of the tip and a second passage for supplying fluid to the distal region of the tip. two channels; a console coupled to the ultrasonic handpiece and configured to generate the AC drive signal applied to the at least one driver of the ultrasonic handpiece to vibrate the tip of the ultrasonic handpiece, the console comprising : a first sensor for measuring the voltage of the AC drive signal; a second sensor for measuring the current of the AC drive signal; a processor coupled to the first and second sensors and configured to: providing to the at least one driver of the ultrasonic handpiece, the AC drive signal induces vibrations insufficient to ablate patient tissue at a distal region of the tip; measuring the voltage and current of the AC drive signal using the first and second sensors; and At least one of an audible, visual or tactile indication is provided based on the measured voltage and current. 38. The system of claim 37 , wherein the processor is configured to provide at least one of an audible, visual or tactile indication based on the measured voltage and current by being configured to: Identifying a property of the patient's tissue based on the measured voltage and current; providing at least one of an audible, visual, or tactile indication of the identified property. 39. The system of claim 38, wherein the processor is configured by being configured to perform to identify the characteristics of the patient's tissue based on the measured voltage and current: to calculate the equivalent current through the mechanical components of the ultrasonic handpiece based on the measured voltage and measured current; and to calculate the equivalent current through the mechanical components of the ultrasonic handpiece based on the calculated The equivalent current to identify the characteristics of the patient tissue. 40. The system of claim 39 , wherein the AC drive signal provided to the ultrasonic handpiece includes a first component having a resonant frequency of the ultrasonic handpiece and a second component having a probe frequency less than the resonant frequency, and the processor , configured to identify a characteristic of patient tissue based on the calculated equivalent current through a mechanical component of the ultrasonic handpiece by: calculating a first amplitude of the measured voltage at the probe frequency, calculating The second amplitude at the probe frequency of the calculated equivalent current through the mechanical part of the ultrasonic handpiece, and the phase at the probe frequency between the measured voltage and the calculated equivalent current through the mechanical part of the ultrasonic handpiece difference; identifying a characteristic of the patient tissue based on the calculated first amplitude, the calculated second amplitude, and the calculated phase difference. 41. The system according to any one of claims 37-40, wherein the processor is configured to provide an audible, visual or tactile indication based on the measured voltage and current by being configured to perform at least one of: identifying a characteristic of the patient tissue based on the measured voltage, the measured current, and the unloaded resistance of the ultrasonic handpiece; and providing at least one of an audible, visual, or tactile indication of the identified characteristic. 42. The system of any one of claims 37-41, wherein the console further comprises a memory coupled to the processor that stores tissue property data indicative of potential tissue properties, and for potential tissue One or more values specific to each of the potential tissue properties in the properties, and the processor is configured to provide an auditory sense based on the measured voltage and current by being configured to perform the following at least one of, a visual or a tactile indication: identifying one of the potential tissue properties indicated by the tissue property data based on the one or more values specific to the potential tissue property and the measured voltage and current , as a property of the patient tissue; and providing at least one of an audible, visual, or tactile indication of the identified property of the patient tissue. 43. The system of claim 42 , wherein the processor is configured to identify an ultrasonic handpiece caused by Said one of the potential tissue properties indicated by the tissue property data is said one of the potential tissue properties. 44. The system of claim 41 or 43, wherein the processor is configured to determine the no-load resistance of the ultrasonic handpiece by being configured to perform the following in response to the ultrasonic handpiece being connected to the console : When the ultrasonic handpiece is in a no-load state, an AC drive signal is provided to the ultrasonic handpiece; When the ultrasonic handpiece is in an unloaded state, the first and second sensors are used to measure the AC drive signal provided to the ultrasonic handpiece a second voltage and a second current of the AC drive signal; and calculating an unloaded resistance of the ultrasonic handpiece based on the measured second voltage and the measured second current of the AC drive signal. 45. The system of claim 41 or 43, wherein the processor, by being configured to read from a memory integrated with the ultrasound handpiece in response to the ultrasound handpiece being connected to the console, indicates no load resistance. data configured to determine the no-load resistance of the ultrasonic handpiece. 46. A method of probing tissue in a patient using an ultrasonic tool system comprising an ultrasonic handpiece, The ultrasonic handpiece has a tip for treating tissue of a patient and at least one driver, the tip is coupled to the driver and an AC drive signal is applied to the driver to vibrate the tip, the method comprising: handheld by ultrasound supplying fluid to the distal region of the tip by at least a portion of the piece; providing suction at the distal region of the tip through at least a portion of the ultrasonic handpiece; providing an AC drive signal to the ultrasonic handpiece, the AC drive signal being at the distal region of the tip Inducing insufficient vibration at the side region to ablate patient tissue; measuring voltage and current of the AC drive signal; and providing at least one of an audible, visual, or tactile indication based on the measured voltage and current. 47. The method of claim 46 , wherein providing at least one of an audible, visual, or tactile indication based on the measured voltage and current comprises: identifying a characteristic of the patient's tissue based on the measured voltage and current; providing the identified at least one of an auditory, visual or tactile indication of a characteristic of the 48. The method of claim 47, wherein identifying characteristics of the patient's tissue based on the measured voltage and current comprises: calculating an equivalent current through a mechanical component of the ultrasonic handpiece based on the measured voltage and the measured current; The calculated equivalent currents through the mechanical components of the ultrasound handpiece identify characteristics of the patient tissue. 49. The method of claim 48, wherein the AC drive signal provided to the ultrasonic handpiece includes a first component having a resonant frequency of the ultrasonic handpiece and a second component having a probe frequency less than the resonant frequency, and based on calculating Characterizing the patient tissue by calculating the equivalent current through the mechanical components of the ultrasonic handpiece includes: calculating the first amplitude of the measured voltage at the probe frequency, the calculated equivalent current through the mechanical component of the ultrasonic handpiece at the second amplitude at the probe frequency, and the phase difference at the probe frequency between the measured voltage and the calculated equivalent current through the mechanical part of the ultrasonic handpiece; based on the calculated first amplitude, the calculated first Two amplitude and phase differences are calculated to identify patient tissue properties. 50. The method of any one of claims 46-49, further comprising identifying a characteristic of the patient's tissue based on the measured voltage, the measured current, and the unloaded resistance of the ultrasound handpiece. 51. The method according to any one of claims 46-50, wherein the ultrasound tool system further comprises a memory storing tissue property data indicative of potential tissue properties, and for each of the potential tissue properties One or more values specific to a potential tissue property, and providing at least one of an audible, visual, or tactile indication based on the measured voltage and current includes: The one or more values and the measured voltage and current identify one of the potential tissue properties indicated by the tissue property data as a property of the patient tissue; providing an audible, At least one of visual or tactile indications. 52. The method of claim 51 , further comprising identifying a tissue indicated by tissue property data based on the one or more values specific to the underlying tissue property, the measured voltage and current, and the no-load resistance of the ultrasonic handpiece. The one of the potential tissue properties of . 53. The method of claim 50 or 52, further comprising determining the no-load resistance of the ultrasonic handpiece by: positioning the ultrasonic handpiece in an unloaded state; and when the ultrasonic handpiece is positioned in an unloaded state : providing an AC drive signal to the ultrasonic handpiece; measuring a second voltage and a second current of the AC drive signal provided to the ultrasonic handpiece; calculating an ultrasonic handpiece based on the measured second voltage and the measured second current of the AC drive signal The no-load resistance of the device. 54. The method of claim 50 or 52, further comprising determining the no-load resistance of the ultrasound handpiece by reading data indicative of the no-load resistance from a memory integrated with the ultrasound handpiece. 55. An ultrasonic tool system comprising: an ultrasonic handpiece comprising a tip having a distal region for treating tissue of a patient and at least one driver, the tip coupled to the driver and an AC drive signal applied to the driver vibrating the tip, the ultrasonic handpiece defining a first channel for providing suction at a distal region of the tip and a second channel for supplying fluid to the distal region of the tip; a console coupled to an ultrasonic handpiece and configured to generate the AC drive signal applied to the at least one driver of the ultrasonic handpiece to vibrate the tip of the ultrasonic handpiece; and a switch communicatively coupled to the console, the switch having a first A setting and a second setting, wherein, in response to the switch being set to the first setting, the console is configured to operate the ultrasound handpiece in an exploratory mode, and in response to the switch being set to the second setting, the console is configured to Configured to operate the ultrasound handpiece in ablation mode. 56. A method for operating an ultrasonic tool system comprising an ultrasonic handpiece having a tip for treating tissue of a patient and at least one driver, the tip coupled to the driver and An AC drive signal is applied to the driver to vibrate the tip, the ultrasonic tool system includes a switch having a first setting and a second setting, the method comprising: providing the AC drive signal to the at least one of the ultrasonic handpieces a driver to vibrate the tip of the ultrasonic handpiece; supply fluid to the distal region of the tip through at least a portion of the ultrasonic handpiece; provide suction at the distal region of the tip through at least a portion of the ultrasonic handpiece ; monitoring the state of the switch to determine whether the switch is set to the first setting or the second setting; determining that the switch is set to the first setting; in response to determining that the switch is set to the operating the ultrasound handpiece in an exploratory mode; determining that the switch is set to the second setting; and in response to determining that the switch is set to the second setting, operating the ultrasound handpiece in an ablation mode Ultrasound Handpiece. | 2020 | Stryker | Medical and Healthcare | Large-Size Company | https://www.stryker.com | USA | The patent focuses on an ultrasonic surgical aspirator whose console drives a vibrating tip with both a resonant-frequency and a sub-resonant probe component, while measuring voltage. It addresses the limitation that surgical aspirators typically only treat tissue without sensing it, missing diagnostic information. The probe component lets the same tool sense tissue in addition to ablating/aspirating it. | הפטנט מתמקד בשואב כירורגי אולטרסוני שהקונסולה שלו מניעה קצה רוטט עם רכיב גשוש בתדר תהודה ובתת-תהודה, תוך מדידת מתח. הוא מתמודד עם המגבלה ששואבים כירורגיים בדרך כלל רק מטפלים ברקמה ללא חישתה, ומפספסים מידע אבחוני. רכיב הגשוש מאפשר לאותו כלי לחוש רקמה בנוסף לאבלציה/שאיבה. | 1 |  | 1 | [12 empty] | 1 | [5 empty] | 1 | [8 empty] | 1 |  | STRYKER CORP | Stryker | STRYKER CORP | Stryker | Stryker | 2021-06-04 | 2021 | 2023-03-14 | 2023 | 2020-06-04 | 2020 | Application | Utility Invention | inactive | CN20218042977 | 317125923 | AU2021283983A1¦CA3181229A1¦CN115802962B¦EP4161411A1¦EP4161411B1¦EP4523638A2¦EP4523638A3¦JP2023528485A¦JP2026048663A¦JP7777546B2¦US20230200835A1¦WO2021248062A1 | CN
401 | CN115969500A | Multifunctional PFA pulse ablation catheter assembly | The invention discloses a multifunctional PFA pulse ablation catheter assembly which comprises an ablation catheter. The ablation catheter comprises a catheter body; the ablation catheter head end integrally extends outwards from the catheter body to form a preset shape and is used for ablating surrounding cells; the mapping catheter penetrates through the inside of the ablation catheter, forms a preset shape, extends to the far end of the head end of the ablation catheter, and is used for recording changes of surrounding electric signals; a plurality of electrodes are arranged outside the ablation catheter and the mapping catheter, and the electrodes of the ablation catheter form a plurality of electric fields. According to the device, the mapping catheter is used for replacing a guide steel wire to guide and support the ablation catheter, meanwhile, a three-dimensional mapping system and other mapping catheters are replaced to evaluate a pulmonary vein isolation result after ablation, cost is reduced, and meanwhile, changes of cardiac cavity potential can be recorded during ablation. | 1. A multifunctional PFA pulse ablation catheter assembly, characterized in that: it comprises an ablation catheter (1); the ablation catheter (1) comprises a tube body (101); the ablation catheter head end (102), from the tube body (101) extending outward integrally to form a predetermined shape for ablation of surrounding cells; the mapping catheter (2) passes through the inside of the ablation catheter (1) and forms a predetermined shape, and extends to the head end of the ablation catheter ( 102) for recording changes in surrounding electrical signals; the external settings of the ablation catheter (1) and the mapping catheter (2) are all provided with a plurality of electrodes (3), and the electrodes (3) of the ablation catheter (1) Multiple electric fields are formed. 2. A kind of multifunctional PFA pulse ablation catheter assembly according to claim 1, is characterized in that: described ablation catheter tip (102) is annular, petal-shaped, divergent shape or ball net shape. 3. A kind of multifunctional PFA pulse ablation catheter assembly according to claim 1, is characterized in that: the head end of described mapping catheter (2) is ring-shaped. 4. A kind of multifunctional PFA pulse ablation catheter assembly according to claim 1, it is characterized in that: the number of electrodes outside the head end (102) of the ablation catheter is 8-30, and the electrodes on the head end of the mapping catheter (2) The quantity is 8-20pcs. 5. A kind of multifunctional PFA pulse ablation catheter assembly according to claim 1, is characterized in that: the length of described electrode (B) is 1-5mm, and described electrode (B) is wrapped in described ablation catheter (1) and described mapping The outside of the catheter (2). 6. A kind of multifunctional PFA pulse ablation catheter assembly according to claim 1, is characterized in that: the head end of described mapping catheter (2) is soft material. 7. A multifunctional PFA pulse ablation catheter assembly according to claim 1, characterized in that: the diameter of the mapping catheter (2) is smaller than the diameter of the lumen of the tube body (101). | 2022 | SUZHOU XINXIAN MEDICAL TECH | Medical and Healthcare | NA | http://xxeswl.com | China | The patent focuses on a multifunctional PFA ablation catheter assembly: an ablation catheter with a preset distal shape, plus a mapping catheter inside it that takes a defined shape and extends to the tip. It addresses the workflow burden of needing separate mapping and ablation catheters in cardiac procedures. The same device records signals (mapping) and delivers PFA energy. | הפטנט מתמקד במכלול קטטר אבלציה PFA רב-תכליתי: קטטר אבלציה עם צורה דיסטלית מוגדרת מראש, בתוספת קטטר מיפוי בתוכו הלוקח צורה מוגדרת ומשתרע לקצה. הוא מתמודד עם נטל זרימת העבודה של צורך בקטטרי מיפוי ואבלציה נפרדים בפרוצדורות לבביות. אותו מכשיר רושם אותות (מיפוי) ומספק אנרגיית PFA. | 1 | [10 empty] | 1 | [12 empty] | 1 | [7 empty] | SUZHOU XINXIAN MEDICAL TECH CO LTD | SUZHOU XINXIAN MEDICAL TECH CO LTD | SUZHOU XINXIAN MEDICAL TECH CO LTD | SUZHOU XINXIAN MEDICAL TECH CO LTD | SUZHOU XINXIAN MEDICAL TECH CO LTD | 2022-06-09 | 2022 | 2023-04-18 | 2023 | 2022-06-09 | 2022 | Application | Utility Invention | active | CN202210653906 | 288270988 | [NA] | CN
402 | CA3213212A1 | LASER SYSTEM AND METHOD FOR DETECTING AND PROCESSING INFORMATION | The present disclosure provides a laser system suitable for modification of a calcified blood vessel, comprising: a laser source; a feedback controller, configured to regulate a dosimetry of the laser source to produce spatially and/or temporally modulated laser light; a catheter comprising a first optical delivery element, the first optical delivery element configured to guide the modulated laser light to an in-vivo object in the blood vessel; and a detecting element, configured to detect one or more physical, chemical, mechanical and/or dimensional characteristics of an area of the in-vivo object in real-time, wherein the feedback controller is configured to process the real-time detected information pertaining to the one or more physical, chemical, mechanical and/or dimensional characteristics of the area in real-time, wherein the feedback controller is further configured to regulate in real-time the dosimetry of the laser source based on the real-time-detected information for a controlled formation of a porous structure and/or a zone of denaturized tissue in the in-vivo object. | 1. i. A laser system suitable for modification of a calcified blood vessel (201), comprising a laser source (101) a feedback controller (106), configured to regulate a dosimetry of the laser source (101) to produce spatially and/or temporally modulated laser light a catheter (103) comprising a first optical delivery element (102), the first optical delivery element (102) configured to guide the modulated laser light to an in-vivo object (202) in the blood vessel (201); and a detecting element (105), configured to detect one or more physical, chemical, mechanical and/or dimensional characteristics of an area of the in-vivo object in realtime, wherein the feedback controller (106) is configured to process the real-time detected information pertaining to the one or more physical, chemical, mechanical and/or dimensional characteristics of the area in real-time, wherein the feedback controller (106) is further configured to regulate in real-time the dosimetry of the laser source (101) based on the real-time-detected information for a controlled formation of a porous structure and/or a zone of denaturized tissue in the in-vivo object (202). 2. The laser system of claim 1, wherein the one or more physical, chemical, mechanical and/or dimensional characteristics comprise at least one of: a position of the in-vivo object (202) a composition of the in-vivo object (202) a dimension of the in-vivo object (202) a temperature of the in-vivo object (202) and/or of an environment of the in-vivo object (202) a stress distribution of the area of the in-vivo object (202) a stress distribution near a vessel wall (202a) of the blood vessel (201) a light scattering induced by the area of the in-vivo object (202) a conductivity of the in-vivo object (202) a Young’s modulus of the area of the in-vivo object (202) a characteristic pertaining to a porous structure and/or a zone of denaturized tissue on the in-vivo object (202) a lumen area of the blood vessel (201) Date Reçue/Date Received 2023-09-20 a compliance of the blood vessel (201) a strength of the in-vivo object (202); and a plasticity threshold of the in-vivo object (202). 3. The laser system of claims 1 or 2, wherein the feedback controller (106) is configured to control, based on the real-time detected information, a distance between the optical delivery element (102) and the in-vivo object (202) in the course of an irradiation by the laser source. 4. The laser system of any one of the preceding claims, wherein the first optical delivery element (102) comprises a bundle of optical fibers and/or is configured to multiplex a plurality of laser outputs of the laser source (101) into one fiber at an input of the first optical delivery element (102). 5. The laser system of any one of the preceding claims, further comprising a balloon (104), configured to be inflated and deflated in the bloodvessel (201); wherein the feedback controller (106) is further configured to control a gas pressure in the balloon (104), and/or to implement a desired positioning of the balloon (104) in realtime based on the real-time detected information. 6. The laser system of any one of the preceding claims, wherein the feedback controller (106) comprises and/or is coupled to a remote highperformance computer, a hybrid quantum-classical computational facility, and/or a quantum computer (io6d); and/or the feedback controller (106) comprises and/or is connected to a storage device, the storage device storing an offline settings table, wherein the settings table is calculated by a remote high-performance computer, a remote hybrid quantum-classical computational facility, and/or a remote quantum computer (io6d). 7. A method for detecting and processing information, comprising a) detecting one or more physical, chemical, mechanical and/or dimensional characteristics of an area of an in-vivo object (202) in a calcified blood vessel (201); and b) processing the detected information pertaining to the physical, chemical, mechanical and/or dimensional characteristics to acquire a property of a porous structure formation and/or a formation of a zone of denaturized tissue in the in-vivo object (202), Date Reçue/Date Received 2023-09-20 wherein the information pertaining to the physical, chemical, mechanical and/or dimensional characteristics of the in-vivo object (202) area is detected and processed in real-time during the porous structure formation and/or the formation of the zone of denaturized tissue. 8. The method of claim 7, wherein the processing of the detected information comprises generating a value for a dosimetry of a laser source (101) in real-time based on the detected information pertaining to the physical, chemical, mechanical and/or dimensional characteristics of the in-vivo object (202) area, wherein the porous structure formation and/or the formation of the zone of denaturized tissue is induced by temporally and/or spatially modulated laser light generated by the laser source (101). 9. The method of claim 8, wherein the generating of the value for the dosimetry of a laser source (101) comprises: generating a time interval between two laser pulses to be longer than a time it takes for a fluid to fill the porous structure after a porous structure formation, wherein the time interval is generated based on the pore size distribution of the porous structure. 10. The method of any one of the claims 7 to 9, wherein the detecting of the physical, chemical, mechanical and/or dimensional characteristics of the area of the in-vivo object comprises detecting a temperature in the area, wherein the value for the dosimetry of the laser source (101) is generated if the temperature is within a predetermined range. 11. The method of any one of claims 7 to 10, wherein the physical, chemical, mechanical and/or dimensional characteristics comprises a characteristic of a scattered light, and the processing of the detected information further comprises: calculating a pore size distribution of the porous structure based on the scattered light. 12. The method of any one of claims 7 to 11, wherein the processing of the detected information further comprises calculating a stress distribution and/or a temperature distribution in the area of the in-vivo object (202); and/or mapping a stress distribution to a temperature distribution and/or evaluating the correlation between a stress distribution and a temperature distribution. Date Reçue/Date Received 2023-09-20 13- The method of any one of daims 7 to 12, further comprising detecting a physical, chemical, mechanical and/or dimensional characteristic of the area of the in-vivo object (202) in the calcified blood vessel (201) before, during, and/or after the porous structure formation and/or the formation of the zone of 5 denaturized tissue in the in-vivo object (202) processing the detected information pertaining to the physical, chemical, mechanical and/or dimensional characteristic of the area of the in-vivo object (202) in the calcified blood vessel (201) before, during and/or after the porous structure formation and/or the formation of the zone of denaturized tissue in the in-vivo object (202) to 10 identify a location where a stress is higher than a predetermined value. 14. The method of any one of the daims 7 to 13, further comprising detecting a physical, chemical, mechanical and/or dimensional characteristic of the area of an in-vivo object (202) on the blood vessel wall (202a) after the blood vessel 15 wall (202a) has undergone a mechanical action. 15 The method of any one of the claims 7 to 14, wherein the processing of the detected information is performed in real time using a remote high-performance computer, a hybrid quantum-classical computational facility, and/or a quantum computer (io6d). Date Reçue/Date Received 2023-09-20 | 2022 | TERRA QUANTUM | AI-driven healthcare company | Start Up | https://terraquantum.swiss | Switzerland | The patent focuses on a laser system for modifying calcified blood vessels using a laser source, a feedback controller producing spatially/temporally modulated light, a catheter optical delivery element, and a detector. It addresses the high mortality of coronary heart disease and the difficulty of treating calcified plaque safely with conventional PCI. The detector's signal feeds back into laser dosing for safer, more effective plaque modification. | הפטנט מתמקד במערכת לייזר לשינוי כלי דם מסויידים באמצעות מקור לייזר, בקר היזון חוזר המייצר אור מודולרי מרחבית/זמנית, אלמנט מסירה אופטית של קטטר וגלאי. הוא מתמודד עם תמותה גבוהה ממחלת לב כלילית ועם הקושי בטיפול בלוחית מסויידת באופן בטוח עם PCI קונבנציונלי. אות הגלאי מוזן בחזרה למינון הלייזר לשינוי לוחית בטוח ויעיל יותר. |  | 1 | [11 empty] | 1 | [16 empty] | 1 |  | TERRA QUANTUM AG | TERRA QUANTUM AG | TERRA QUANTUM AG | TERRA QUANTUM AG | TERRA QUANTUM AG | 2023-09-20 | 2023 | 2024-03-23 | 2024 | 2022-09-23 | 2022 | Application | Utility Invention | active | CA20233213212 | 318824047 | AU2023229530A1¦AU2023229530B2¦AU2023229602A1¦AU2023229602B2¦AU2023229607A1¦AU2023229607B2¦CA3213220A1¦CA3213431A1¦CN117716325A¦CN117752412A¦CN117752491A¦CN117752954A¦CN117769694A¦EP4342407A1¦EP4342408A1¦EP4342435A1¦EP4377769A1¦EP4377771A1¦JP2024046744A¦JP2024046746A¦JP2024046747A¦JP7755322B2¦JP7807088B2¦KR20240041830A¦KR20240041831A¦KR20240042353A¦TW202319888A¦TW202324065A¦US12096112B2¦US12132983B2¦US12501150B2¦US20230031871A1¦US20230032467A1¦US20240099771A1¦US20240099888A1¦US20240130789A1¦US20250024138A1¦US20250030940A1¦WO2023009806A1¦WO2023009853A1 | CA
403 | CN119700291A | CT-guided tumor ablation dose regulation and control method and system | The invention relates to the technical field of tumor treatment, and discloses a CT-guided tumor ablation dose regulation and control method and system, and the method comprises the steps: obtaining the three-dimensional contour data and gray value distribution of a tumor through CT scanning, calculating the volume and average density of the tumor, and obtaining the initial ablation dose. Initial parameters of ablation equipment are set, an ablation program is started, data are updated at specific time intervals in the ablation process, and the real-time ablation dosage is calculated. And comparing the real-time ablation dose with the initial ablation dose, and adjusting the power of ablation equipment according to a specific formula and a comparison result. The system comprises an imaging and parameter acquisition module, an initial dose calculation module, an ablation equipment control module, a real-time monitoring module, a dose comparison and regulation module and the like. The ablation dosage can be accurately regulated and controlled according to the real-time condition of a tumor, the ablation treatment effect is improved, the damage risk to normal tissue is reduced, an alarm can be given to prompt an operator when the dosage is abnormal, and the high clinical application value is achieved. | 1. L A method for controlling the dose of tumor ablation under CT guidance, characterized in that the method comprises: starting a CT scanning device to perform three-dimensional imaging of the tumor site, obtaining three-dimensional contour data of the tumor, and calculating the tumor volume based on the contour data; calculating the average density of the tumor through the gray value distribution of the CT image; and calculating the initial ablation dose based on the tumor volume and the average density. Add calculation; setting the initial parameters of the ablation device according to the initial ablation dose, starting the ablation program, and during the ablation process, starting the CT scanning device again at a specific interval of time nine to update the three-dimensional contour data and gray value distribution of the tumor, repeating the above volume, density and dose calculation steps to obtain the real-time ablation dose Oyeoi; comparing the real-time ablation dose Oyeoi with the initial ablation dose Dinit, and when D3I > Dinit XQ +. ), reducing the power of the ablation device, where & is the preset upper limit floating ratio of the dose; when D3I < Dinit XQ +.init When X(1-(3)), the power of the ablation device is increased, and X is the preset floating ratio of the lower limit of the dose. 2.The tumor ablation dose control method according to claim 1 is characterized in that the tumor volume V is calculated according to the contour data through the formula V = £1 Si × hi, where Si is the cross-sectional area of the i-th layer, hi is the distance between two adjacent cross-sectional layers, and η is the number of layers. 3The method for controlling the dose of tumor ablation according to claim 1 is characterized in that the average density of the tumor is calculated by combining the gray value distribution of the CT image with the formula: Pj is the density value corresponding to the jth pixel point, and m is the total number of pixels selected; the method for selecting the pixel points is: taking the geometric center of the tumor as the center and r as the radius to draw a circle, and selecting the pixel points in the circle to participate in the calculation, and r is determined according to the maximum diameter R of the tumor through the formula r = and. 4. The tumor ablation dose control method according to claim 1 is characterized in that the formula for calculating the initial ablation dose Dmit is: Dinit - k XVX ρ; where Α; is the empirical coefficient, the value of Α; is related to the tissue type of the tumor, and different types of tumors correspond to different constant values. 5. The tumor ablation dose control method according to claim 1 is characterized in that the step of setting the initial parameters of the ablation device according to the initial ablation dose includes: calculating the initial power R value t according to the energy conversion efficiency of the ablation device through the formula Pinit = 4, where AT is the preset initial ablation time; setting the initial coolant flow rate according to the heat dissipation characteristics of the ablation needle, and the flow rate and power satisfy the relationship of CX = CX, where c is the heat dissipation coefficient. 6.The tumor ablation dose control method according to claim 1 is characterized in that in the step of using force to turn on the CT scanning device to update data at each specific time interval, the value of force is related to the growth rate of the tumor, and the faster the growth rate, the smaller the value of force. 7The method for controlling the dose of tumor ablation according to claim 1 is characterized in that after the step of reducing or increasing the power of the ablation device, it further comprises: continuing to monitor the real-time ablation dose reoi, and if the monitoring results show that reoi still exceeds or falls below the corresponding dose range after N consecutive times, an alarm is issued to prompt the operator, and N is the preset alarm threshold number. 8.A CT-guided tumor ablation dose control system, characterized in that it includes the following modules: an imaging and parameter acquisition module, which turns on a CT scanning device to perform three-dimensional imaging of the tumor site, obtains three-dimensional contour data of the tumor, calculates the tumor volume based on the contour data, and calculates the average density of the tumor through the grayscale value distribution of the CT image; an initial dose calculation module, which calculates the initial ablation dose Dmit using a preset formula based on the tumor volume and average density; the formula for calculating the initial ablation dose Dmit is: ^init - k XVX ρ·, where V is the tumor volume, 0 is the average density of the tumor, Α; is an empirical coefficient, and the value of Α; is related to the tissue type of the tumor, and different types of tumors correspond to different constant values; an ablation device control module, which sets the initial parameters of the ablation device according to the initial ablation dose Dinit and starts the ablation program; a real-time monitoring module, which turns on the CT scanning device again at a specific interval to update the three-dimensional contour data and grayscale value distribution of the tumor, and repeats the above volume, density and dose calculation steps to obtain the real-time ablation dose OreaZ; a dose comparison and control module, which converts the real-time ablation dose Dreai Compared with the initial ablation dose Dmit, when Dreal > Dinit X (1 + .), reduce the power of the ablation device, where & is the preset upper dose floating ratio; when Dreai < Dinit When X(1-(3)), the power of the ablation device is increased, and X(1-(3)) is the floating ratio of the preset lower limit of the dose. 9. The tumor ablation dose control system according to claim 8 is characterized in that the ablation device control module includes a power setting submodule, which calculates the initial power Pinit according to the energy conversion efficiency of the ablation device through the formula Pinit = ΔT X (7, where AT is the preset initial ablation time; according to the heat dissipation characteristics of the ablation needle, sets the initial coolant flow rate Gt, the flow rate and power satisfy the relationship of Gt = c X t, where c is the heat dissipation coefficient. 10.The tumor ablation dose control system according to claim 8 is characterized in that the imaging and parameter acquisition module: calculates the tumor volume V according to the contour data through the formula V = Si × hi, where Si is the cross-sectional area of the i-th layer, % is the spacing between two adjacent cross-sectional layers, and η is the number of layers; calculates the average tumor density O through the gray value distribution of the CT image in combination with the formula o = £Taozhu, where spoon is the density value corresponding to the j-th fm pixel point, and m is the total number of selected pixel points; the method for selecting pixel points is: draw a circle with the geometric center of the tumor as the center and r as the radius, select the pixel points within the circle to participate in the calculation, and r is determined according to the maximum diameter R of the tumor through the formula r = and. | 2025 | Sun Yat-sen University | University | NA | https://www.sysu.edu.cn | China | The patent focuses on a CT-guided tumor-ablation dose-control method that calculates tumor volume and density from CT contour and gray-value data to set the initial ablation dose, then updates parameters at intervals. It addresses the limitations of conventional CT-guided ablation, where dose is set once and not adjusted based on procedure data. The system updates parameters dynamically during ablation based on new CT data. | הפטנט מתמקד בשיטת בקרת מינון אבלציית גידולים מונחית CT המחשבת נפח וצפיפות גידול מנתוני מתאר וערכי אפור CT לקביעת מינון האבלציה ההתחלתי, ואז מעדכנת פרמטרים במרווחים. הוא מתמודד עם המגבלות של אבלציה מונחית CT קונבנציונלית, שבה המינון נקבע פעם אחת ואינו מותאם על בסיס נתוני פרוצדורה. המערכת מעדכנת פרמטרים באופן דינמי במהלך האבלציה על בסיס נתוני CT חדשים. | [8 empty] | 1 | [12 empty] | 1 | [6 empty] | 1 |  |  |  | THE SIXTH AFFILIATED HOSPITAL SUN YAT SEN UNIV | Sun Yat-sen University | THE SIXTH AFFILIATED HOSPITAL SUN YAT SEN UNIV | Sun Yat-sen University | Sun Yat-sen University | 2025-02-21 | 2025 | 2025-03-28 | 2025 | 2025-02-21 | 2025 | Application | Utility Invention | active | CN202510196901 | 320365316 | [NA] | CN
404 | WO2025263780A1 | HIGH-FREQUENCY THERMAL COAGULATION CONTROL DEVICE AND SYSTEM USING THREE-DIMENSIONAL BRAINWAVE ELECTRODE BASED ON REAL-TIME TEMPERATURE MONITORING | The present invention may be characterized by: communicating with a plurality of first sensors which are provided on one side of the inside of a rod-type electrode member in a high-frequency thermal coagulation device and detect brain waves of a high-frequency irradiation site in the brain, the rod-type electrode member being inserted into the high-frequency irradiation site and including a plurality of electrodes; receiving brain wave information acquired through the plurality of first sensors; communicating with a plurality of second sensors which are provided on the other side of the inside of the rod-type electrode member in the high-frequency thermal coagulation device and detect the temperature of the high-frequency irradiation site; receiving temperature information acquired through the plurality of second sensors; and when the plurality of electrodes reach the high-frequency irradiation site, controlling an energy supply unit so that a current is applied to the plurality of electrodes at a preset current intensity in connection with the brain wave information and the temperature information. | 1. A communication unit for performing communication with a high-frequency thermal coagulation device; and a processor for performing operations related to controlling the high-frequency thermal coagulation.Including, the communication unit, which is inserted into a high-frequency irradiation site of the brain among the high-frequency thermal coagulation devices and is provided on one side of the inside of a rod-type electrode member including a plurality of electrodes, performs communication with a plurality of first sensors that detect brain waves of the high-frequency irradiation site, and which is provided on the other side of the inside of the rod-type electrode member among the high-frequency thermal coagulation devices and detects the temperature of the high-frequency irradiation site, and performs communication with an energy supply unit that selectively applies current at a preset current intensity to the plurality of electrodes among the high-frequency thermal coagulation devices, and the processor receives brain wave information acquired through the plurality of first sensors, receives temperature information acquired through the plurality of second sensors, and controls the energy supply unit so that when the plurality of electrodes reach the high-frequency irradiation site, the current is applied to the plurality of electrodes at a preset current intensity in connection with the brain wave information and the temperature information.High frequency thermal coagulation control device. 2. In the first paragraph, the high-frequency thermal coagulation control device, characterized in that the processor receives the brain wave information through the plurality of first sensors provided corresponding to the plurality of electrodes on one inner side of the electrode member. 3. In the second paragraph, the high-frequency thermal coagulation control device, characterized in that the processor receives the brain wave information through the plurality of first sensors respectively provided corresponding to the plurality of electrodes at preset equal intervals. 4. In the second paragraph, the high-frequency thermal coagulation control device, characterized in that the processor receives the brain wave information through the plurality of first sensors respectively provided corresponding to the plurality of electrodes at different preset intervals. 5. In the second paragraph, the high-frequency thermal coagulation control device, characterized in that the processor receives the temperature information through the plurality of second sensors provided corresponding to the plurality of electrodes among the inner other sides of the electrode member. 6. In the fifth paragraph, the high-frequency thermal coagulation control device, characterized in that the processor receives the temperature information through the plurality of second sensors respectively provided corresponding to the plurality of electrodes at preset equal intervals. 7. In the fifth paragraph, the high-frequency thermal coagulation control device, characterized in that the processor receives the temperature information through the plurality of second sensors respectively provided corresponding to the plurality of electrodes at different preset intervals. 8. In the first paragraph, the processor controls the energy supply unit so that, when the plurality of electrodes reach the high-frequency irradiation site, a first current of the monopolar type and a second current of the bipolar type are selectively applied to the plurality of electrodes at a preset current intensity linked to the brain wave information and the temperature information, a high-frequency thermal coagulation control device. 9. In the first paragraph, the processor controls the energy supply unit so that, when the plurality of electrodes reach the high-frequency irradiation site, a first current of the monopolar type and a second current of the bipolar type are alternately applied to the plurality of electrodes at a preset current intensity linked to the brain wave information and the temperature information, a high-frequency thermal coagulation control device. 10. In the first paragraph, the processor controls the energy supply unit so that, when the plurality of electrodes reach the high-frequency irradiation site, a first current of the monopolar type and a second current of the bipolar type are sequentially applied to the plurality of electrodes at a preset current intensity linked to the brain wave information and the temperature information, a high-frequency thermal coagulation control device. 11. A method for controlling high-frequency thermal coagulation of a high-frequency thermal coagulation control device, comprising: a step of performing communication with a plurality of first sensors that detect brain waves of the high-frequency irradiation area, which are inserted into a high-frequency irradiation area of the brain among high-frequency thermal coagulation devices and are provided on one side of the inside of a rod-type electrode member including a plurality of electrodes; A step of receiving brain wave information acquired through the plurality of first sensors; A step of performing communication with a plurality of second sensors that are provided on the inner side of the rod-type electrode member of the high-frequency thermal coagulation device and detect the temperature of the high-frequency irradiation area; A method comprising: a step of receiving temperature information acquired through the plurality of second sensors; and a step of controlling the energy supply unit so that, when the plurality of electrodes reach the high-frequency irradiation site, the current is applied to the plurality of electrodes at a preset current intensity linked to the brain wave information and the temperature information. 12. In the 11th paragraph, the control step is characterized in that, when the plurality of electrodes reach the high-frequency irradiation site, the energy supply unit is controlled so that a first current of the monopolar type and a second current of the bipolar type are selectively applied to the plurality of electrodes at a preset current intensity linked to the brain wave information and the temperature information. 13. In the 11th paragraph, the control step is characterized in that, when the plurality of electrodes reach the high-frequency irradiation site, the energy supply unit is controlled so that a first current of the monopolar type and a second current of the bipolar type are alternately applied to the plurality of electrodes at a preset current intensity linked to the brain wave information and the temperature information. 14. In the 11th paragraph, the control step is characterized in that, when the plurality of electrodes reach the high-frequency irradiation site, the energy supply unit is controlled so that a first current of the monopolar type and a second current of the bipolar type are sequentially applied to the plurality of electrodes at a preset current intensity linked to the brain wave information and the temperature information. 15. A high-frequency thermal coagulation control system comprising a high-frequency thermal coagulation device and a high-frequency thermal coagulation control device for controlling the high-frequency thermal coagulation device, wherein the high-frequency thermal coagulation device comprises: a rod-type electrode member inserted into a high-frequency irradiation site of the brain and including a plurality of electrodes; A plurality of first sensors provided on one inner side of the electrode member of the above-mentioned rod type and detecting brain waves of the high-frequency irradiation area; A system comprising: a plurality of second sensors provided on the inner side of the rod-type electrode member and detecting the temperature of the high-frequency irradiation area; and an energy supply unit selectively applying current to the plurality of electrodes; wherein the high-frequency thermal coagulation control device receives brain wave information acquired through the plurality of first sensors, receives temperature information acquired through the plurality of second sensors, and controls the energy supply unit so that, when the plurality of electrodes reach the high-frequency irradiation area, the current is applied to the plurality of electrodes at a preset current intensity linked to the brain wave information and the temperature information. | 2024 | Yonsei University | University | NA | https://www.yonsei.ac.kr | South Korea | The patent focuses on a high-frequency thermal coagulation system using a rod electrode with built-in sensors that detect brain waves at the irradiation site, with real-time temperature monitoring. It addresses the gap that conventional thermal coagulation in the brain doesn't track brainwave changes during treatment. Real-time brainwave info during thermocoagulation guides parameters and monitors safety. | הפטנט מתמקד במערכת קרישה תרמית בתדר גבוה המשתמשת באלקטרודת מוט עם חיישנים מובנים המזהים גלי מוח באתר ההקרנה, עם ניטור טמפרטורה בזמן אמת. הוא מתמודד עם הפער שקרישה תרמית קונבנציונלית במוח אינה עוקבת אחר שינויי גלי מוח במהלך הטיפול. מידע על גלי מוח בזמן אמת במהלך תרמוקואגולציה מנחה פרמטרים ומנטר בטיחות. | 1 |  |  |  | 1 | [24 empty] | 1 |  |  | UIF UNIV INDUSTRY FOUNDATION YONSEI UNIV | Yonsei University | UIF UNIV INDUSTRY FOUNDATION YONSEI UNIV | Yonsei University | Yonsei University | 2025-04-14 | 2025 | 2025-12-26 | 2025 | 2024-06-18 | 2024 | Application | Utility Invention | active | WO2025KR05022 | 325401668 | KR20250178162A | WO
405 | US20250311929A1 | SYSTEM AND METHOD OF DUAL-MODE EAT/US-GUIDED ELECTROPORATION | The present embodiments relate generally to medical care and more particularly to a new imaging modality referred to herein as electroacoustic tomography (EAT) that not only depicts electrical field energy distribution in real time, but also enables clear discrimination between IRE and RE zones in situ during the treatment. According to some aspects, EAT exploits the phenomenon that the amplitude of acoustic emission generated by an electric field is proportional to the electrical energy deposition in tissue. After detecting these acoustic waves with ultrasound transducers, an image of the electric field distribution can be reconstructed in real-time. This will allow real time monitoring the IRE and provides feedback control of the treatment. Some embodiments are directed to a novel EAT combined with ultrasound (EAT/US) guided IRE intervention that permits a) real time ultrasound image-guided needle placement, and b) intraoperational discrimination of IRE and RE zones during the treatment of cancer. | 1. A system for electroacoustic tomography (EAT) including: a device adapted to receive ultrasound signals reflected from transmitted ultrasound waves; a device adapted to receive electro-acoustic signals generated from an electrical pulse generation device; an EAT imaging unit adapted to generate an electroacoustic image based on the received electro-acoustic signals; an ultrasound imaging unit adapted to generate an ultrasound image based on the received ultrasound signals; and an image combining device adapted to receive and combine the generated electroacoustic image and ultrasound image. 2. The system of claim 1, further comprising: a device adapted to generate an electrical field induced acoustic signal; and a device adapted to transmit ultrasound waves. 3. The system of claim 1, further comprising a trigger sequencing logic device for switching a connection of an ultrasound transducer between ultrasound transmission/reception and electro-acoustic signal acquisition. 4. The system of claim 3, wherein the trigger sequencing logic device is configured to perform switching such that the ultrasound signal and the electroacoustic signal are received in different time windows. 5. The system of claim 1, wherein the ultrasound imaging unit includes: a transducer adapted to receive ultrasound signals in an ultrasound reception window controlled by a flip-flop sequencing logic switch; an ultrasonic beamformer adapted to produce a series of radio frequency signals from the ultrasonic signals; and an ultrasound image construction unit adapted to produce the ultrasound image from the radio frequency signals. 6. The system of claim 1, wherein the EAT imaging unit includes: an acquisition stage adapted to collect an electroacoustic signal in response to an electrical pulse signal trigger; a multi-channel pre-amplification stage adapted to amplify the electroacoustic signal; an averaging/filtering stage adapted to process the amplified electroacoustic signal; and an EAT image construction unit adapted to perform electro-acoustic imaging reconstruction. 7. The system of claim 1, wherein the image combiner uses spatially resampling to produce averaged electroacoustic imaging reconstructions so that the electroacoustic image can be superimposed on top of the ultrasound image. 8. The system of claim 7, wherein the image combiner produces combined image frames for display. 9. A method for electroacoustic tomography (EAT) including: depicting electrical field energy distribution in real time; enabling clear discrimination between irreversible electroporation (IRE) zones and reversible electroporation (RE) zones in situ during treatment, including exploiting the phenomenon that the amplitude of acoustic waves generated by an electric field is proportional to the electrical energy deposition in tissue; and after detecting the acoustic waves with ultrasound transducers, constructing an image of the electric field distribution in real-time. 10. The method of claim 9, further comprising enabling real time ultrasound image-guided needle placement. 11. A method for electroacoustic tomography (EAT) including: receiving ultrasound signals reflected from transmitted ultrasound waves; receiving electro-acoustic signals generated from an electrical pulse generation device; generating, by an EAT imaging unit, an electroacoustic image based on the received electro-acoustic signals; generating, by an ultrasound imaging unit, an ultrasound image based on the received ultrasound signals; and receiving, by an image combining device, the generated electroacoustic image and ultrasound image and combining them to generate a combined image. 12. The method of claim 11, further comprising: generating an electrical field induced acoustic signal; and transmitting ultrasound waves. 13. The method of claim 11, further comprising switching a connection of an ultrasound transducer between ultrasound transmission/reception and electro-acoustic signal acquisition. 14. The method of claim 13, wherein switching includes performing the switching such that the ultrasound signal and the electroacoustic signal are received in different time windows. 15. The method of claim 11, wherein generating the ultrasound image includes: receiving ultrasound signals in an ultrasound reception window controlled by a flip-flop sequencing logic switch; producing a series of radio frequency signals from the ultrasonic signals; and producing the ultrasound image from the radio frequency signals. 16. The method of claim 11, wherein generating the electroacoustic image includes: collecting an electroacoustic signal in response to an electrical pulse signal trigger; amplifying the electroacoustic signal; processing the amplified electroacoustic signal in an averaging/filtering stage; and performing electro-acoustic imaging reconstruction based on the processed electroacoustic signal. 17. The method of claim 11, wherein combining includes using spatially resampling to produce averaged electroacoustic imaging reconstructions so that the electroacoustic image can be superimposed on top of the ultrasound image. 18. The method of claim 17, further including producing combined image frames for display. | 2022 | University of California | University | NA | https://www.universityofcalifornia.edu | USA | The patent focuses on a dual-mode electroacoustic-tomography (EAT)/ultrasound electroporation system: EAT shows electric-field energy distribution in real time and discriminates IRE versus reversible electroporation zones. It addresses the difficulty of distinguishing irreversible from reversible electroporation in tissue during electroporation treatments. The acoustic-emission amplitude scales with deposited electrical energy, enabling the discrimination. | הפטנט מתמקד במערכת אלקטרופורציה דו-מצבית טומוגרפיה אלקטרואקוסטית (EAT)/אולטרסאונד: EAT מראה התפלגות אנרגיית שדה חשמלי בזמן אמת ומבחין בין אזורי IRE לאלקטרופורציה הפיכה. הוא מתמודד עם הקושי בהבחנה בין אלקטרופורציה בלתי הפיכה להפיכה ברקמה במהלך טיפולי אלקטרופורציה. משרעת הפליטה האקוסטית מתמתחת עם אנרגיה חשמלית מופקדת, ומאפשרת את ההבחנה. |  |  | 1 | 1 | [7 empty] | 1 | [11 empty] | 1 |  |  |  |  | 1 |  |  |  | UNIV CALIFORNIA | University of California | UNIV CALIFORNIA | University of California | University of California | 2023-05-30 | 2023 | 2025-10-09 | 2025 | 2022-05-31 | 2022 | Application | Utility Invention | active | US202318869803 | 323701983 | WO2023235332A1 | US
406 | US12072283B2 | System, method, computer-accessible and apparatus for providing near-infrared spectroscopy for anatomical mapping of the epicardium | An exemplary catheter can be provided, which can include, for example a source fiber(s) configured to (i) receive a near infrared spectroscopic (NIRS) radiation, and (ii) provide the NIRS radiation to a portion(s) of a sample(s), a detection fiber(s) configured to receive a return radiation from the sample(s) that can be based on the NIRS radiation that was provided to the portion(s) of the sample(s), and an ablation electrode(s) configured to ablate the sample(s) based on the return radiation. The source fiber(s), the detection fiber(s), and the ablation electrode(s) can be integrated into the single sheath. The ablation electrode(s) can be a radiofrequency ablation electrode. | 1. A catheter comprising: at least one source fiber configured to (i) receive a near infrared spectroscopic (NIRS) radiation, and (ii) provide the NIRS radiation to at least one portion of at least one sample; at least one detection fiber configured to receive a return radiation from the at least one sample that is based on the NIRS radiation that was provided to the at least one portion of the at least one sample; and at least one ablation electrode configured to ablate the at least one sample based on the return radiation. 2. The catheter of claim 1, further comprising a single sheath, wherein the at least one source fiber, the at least one detection fiber, and the at least one ablation electrode are integrated into the single sheath within the catheter. 3. The catheter of claim 1, wherein the at least one ablation electrode is a radiofrequency ablation electrode. 4. The catheter of claim 1, wherein a configuration of the at least one source fiber and the at least one detection fiber are spaced apart at a specific distance from one another. 5. The catheter of claim 4, wherein the specific distance is between about 0 mm to about 4 mm. 6. The catheter of claim 1, further comprising at least one broadband light source configured to generate the NIRS radiation being received by the at least one source fiber. 7. The catheter of claim 1, further comprising at least one continuous-wave light source configured to generate the NIRS radiation being received by the at least one source fiber. 8. The catheter of claim 1, further comprising a charge coupled device configured to (i) receive the return radiation from the at least one detection fiber, and (ii) generate imaging information based on the return radiation, wherein the imaging information is associated with the at least one portion of the at least one sample. 9. The catheter of claim 1, wherein: the at least one source fiber includes at least six source fibers, the at least one detection fiber includes at least six detection fibers, and the at least one ablation electrode includes at least twelve ablation electrodes, and wherein a combination of the at least six source fibers and the at least six detection fibers facilitates a parallel or simultaneous detection of information regarding the at least one portion of the at least one sample. 10. The catheter of claim 9, wherein: one of the at least six source fibers and one of the at least six detection fibers form a pair of fibers, one of the at least twelve ablation electrodes is disposed on a first side of the pair of fibers, another one of the at least twelve ablation electrodes is disposed on a second side of the pair of fibers, and the first side is disposed opposite to the second side with respect to the pair of fibers. 11. A method for ablating at least one portion of at least one sample, comprising: generating at least one near infrared spectroscopic (NIRS) radiation; providing the at least one NIRS radiation to the at least one portion using at least one source fiber provided in a catheter; receiving at least one return radiation from the at least one portion that is based on the at least one NIRS radiation that was provided to the at least one portion of the at least one sample using at least one detection fiber provided in the catheter; and ablating the at least one portion using at least one ablation electrode provided in the catheter based on the at least one return radiation. 12. The method of claim 11, wherein the at least one ablation electrode includes at least one radiofrequency ablation (RFA) electrode integrated into the catheter. 13. The method of claim 11, further comprising determining a three-dimensional position of the catheter with respect to the at least one portion. 14. The method of claim 11, further comprising determining a contact between the catheter and the at least one portion. 15. The method of claim 11, wherein the at least one sample includes a heart, and further comprising generating at least one epicardial map of the heart based on the at least one return radiation. 16. The method of claim 15, wherein the at least one epicardial map includes at least one of (i) at least one coronary vessel in the heart, (ii) fat in the heart, (iii) scar tissue in the heart, or (iv) fibrosis in the heart. 17. The method of claim 16, further comprising determining a location of the at least one coronary vessel in the heart by determining a hemoglobin concentration in the heart. 18. The method of claim 17, further comprising determining the hemoglobin concentration using a spectral unmixing procedure. 19. The method of claim 18, wherein the spectral unmixing procedure is an inverse Monte Carlo procedure. 20. The method of claim 11, further comprising generating at least one optical substrate map of at least one portion of the at least one sample based on the at least one return radiation. 21. The method of claim 20, further comprising automatically assessing at least one substrate using the at least one optical substrate map, wherein the at least one substrate includes a pulmonary vein, a coronary artery, fat, fibrosis or a scar. 22. The method of claim 20, further comprising: tracking a position of the catheter at or within the tissue; and interpolating the at least one optical substrate map based on the at least one return radiation as a function of the tracked position. 23. The method of claim 20, wherein the generating of the at least one optical substrate map includes tracking at least one input. 24. The method of claim 11, wherein the at least one source fiber includes a plurality of source fibers, wherein the at least one detection fiber includes a plurality of detection fibers, and wherein the receiving of at least one return radiation includes a receipt of plurality of radiations, each from a respective one of the plurality of detection fibers, facilitating a parallel or simultaneous detection of information regarding the at least one portion of the at least one sample. 25. The method of claim 24, wherein the information is simultaneously detected from different locations on or in the at least one portion of the at least one sample. 26. The method of claim 11, wherein the source and detection fibers extend in the catheter for a majority of a length thereof. 27. A non-transitory computer-accessible medium having stored thereon computer-executable instructions for causing an ablation of at least one portion of at least one sample, wherein, when a computing arrangement executes the instructions, the computing arrangement is configured to perform procedures comprising: receiving information related to a backscattered radiation from the at least one portion that is based on at least one near infrared spectroscopic radiation provided to the at least one sample by at least one source fiber provided in a catheter, wherein the received information is based on at least one return radiation at least one detection fiber provided in the catheter; generating at least one epicardial map of the at least one portion based on the information; determining a first location of at least one coronary vessel using the information; and causing the ablation of the at least one portion by at least one ablation electrode provided in the catheter at a second location that excludes the at least one coronary vessel. | 2020 | Columbia University | University | Large-Size Company | https://www.columbia.edu | USA | The patent focuses on a NIRS catheter for epicardial mapping with source fibers delivering near-infrared light, detection fibers reading return light, and an ablation electrode that uses that NIRS info. It addresses the limitation that conventional cardiac ablation lacks integrated optical tissue characterization at the epicardium. The catheter ablates the same area it characterizes optically. | הפטנט מתמקד בקטטר NIRS למיפוי אפיקרדיאלי עם סיבי מקור המספקים אור אינפרא-אדום קרוב, סיבי גילוי הקוראים אור חוזר ואלקטרודת אבלציה המשתמשת במידע NIRS זה. הוא מתמודד עם המגבלה שאבלציה לבבית קונבנציונלית חסרה אפיון רקמה אופטי משולב באפיקרדיום. הקטטר מבצע אבלציה באותו אזור שהוא מאפיין אופטית. |  | 1 | [8 empty] | 1 | [10 empty] | 1 |  |  | 1 | [7 empty] | UNIV COLUMBIA | Columbia University | THE TRUSTEES OF COLUMBIA UNIVERSITY IN THE CITY OF NEW YORK | Columbia University | Columbia University | 2020-06-19 | 2020 | 2024-08-27 | 2024 | 2020-06-19 | 2020 | Grant | Utility Invention | active | US202016906429 | 277640420 | US20210396661A1 | US
407 | US12171524B2 | Devices, systems, and methods for imaging in certain endoscopic environments | Systems and methods for multi-modal imaging using an endoscope having an instrument channel, where the imaging is achieved without using the channel, are disclosed. The systems can include a multi-modal imaging paddle housing couple to a distal end of the endoscope. The housing can receive at least two imaging probes. The imaging probes can be an angle-resolved low-coherence interferometry probe (a/LCI) and an optical coherence tomography (OCT) probe. The housing can be scaled and positioned to be visible via the endoscope camera. The system and method can include locating the housing in a region of interest using the endoscope camera, acquiring OCT measurements to identify targets, and then acquiring a/LCI measurements at the identified targets. | 1. A system for multi-modal imaging using an endoscope having an instrument channel, wherein the multi-modal imaging is achieved without using the instrument channel, the system comprising: an endoscope selected from the group consisting of an anoscope, a colonoscope, a colposcope, an esophagoscope, a gastroscope, a laryngoscope, and a sigmoidoscope, the endoscope having a video camera positioned at a distal end for providing video images; a multi-modal imaging paddle housing coupled to the distal end of the endoscope, the multi-model imaging paddle housing occupying a retracted position adjacent to the distal end of the endoscope, the multi-modal imaging paddle housing optionally adapted to be movable between the retracted position and an extended position that is extended axially and/or radially from the distal end of the endoscope, the multi-modal imaging paddle housing adapted to receive a portion of at least two imaging probes, the multi-modal imaging paddle housing having a single imaging window, through which each of the at least two imaging probes acquire images, or two imaging windows, each through which a respective one of the at least two imaging probes acquires images; an angle-resolved low-coherence interferometry (a/LCI) probe having at least a distal portion received within the multi-modal imaging paddle housing; an optical coherence tomography (OCT) probe having at least a distal portion received within the multi-modal imaging paddle housing; and an imaging controller operatively coupled to the a/LCI probe and the OCT probe, wherein the multi-modal imaging paddle housing has a shape and a location in the retracted position, wherein the shape and the location combine to make the multi-modal imaging paddle housing visible in the video images provided by the video camera, such that a user can visually guide the multi-modal imaging paddle housing to a region of interest, wherein the a/LCI probe and the OCT probe are electronically coupled to the imaging controller and optically coupled to respective light sources and optical detectors via wires and optical fibers that are coupled to an exterior surface of the endoscope, wherein the a/LCI probe operates within a first wavelength range and the OCT probe operates in a second wavelength range, wherein the first and second wavelength ranges are different and non-overlapping, wherein the single imaging window transmits at least 95% of light at each wavelength in the first wavelength range and the second wavelength range, wherein one of the two imaging windows transmits at least 95% of light at each wavelength in the first wavelength range, wherein the other of the two imaging windows transmits at least 95% of light at each wavelength in the second wavelength range, wherein the instrument channel is not occupied by the a/LCI probe or the OCT probe. 2. The system of claim 1, wherein the multi-modal imaging paddle housing further comprises one or more radiofrequency ablation electrodes, wherein the one or more radiofrequency ablation electrodes are coupled to external power and controls via the wires that are coupled to the exterior surface of the endoscope. 3. The system of claim 1, wherein the multi-modal imaging paddle housing including a coupling cuff and is coupled to the distal end of the endoscope via the coupling cuff. 4. The system of claim 1, wherein the first wavelength range includes 840 nm within it and the second wavelength range includes 1310 nm within it. 5. The system of claim 1, wherein the a/LCI probe includes a reflection-only three-optic rotator prism and a two-axis scanning mirror. 6. The system of claim 5, wherein the controller is adapted to direct the a/LCI probe and the respective light sources and optical detectors to acquire A-scans by varying an angle of the reflection-only three-optic rotator prism. 7. The system of claim 1, wherein the a/LCI probe is coupled to the respective light sources and optical detectors via a pathlength-matched linear fiber array comprising a plurality of single-mode optical fibers, via coherent fiber bundles, or via a single multi-mode fiber. 8. The system of claim 1, wherein the OCT probe is a rotating probe. | 2021 | University of North Carolina | University | NA | https://www.northcarolina.edu | USA | The patent focuses on an endoscopic multi-modal imaging accessory that attaches to the distal end without using the instrument channel, holding two imaging probes (e.g., a/LCI and OCT). It addresses the gap that >90% of cancers originate in epithelium and detecting them needs sensitive imaging that doesn't sacrifice the working channel. The accessory enables dual-mode optical imaging without sacrificing the working channel. | הפטנט מתמקד באביזר הדמיה אנדוסקופי רב-מצבי המתחבר לקצה הדיסטלי ללא שימוש בערוץ המכשיר, ומחזיק שני גששי הדמיה (למשל a/LCI ו-OCT). הוא מתמודד עם הפער ש-90%+ מהסרטנים מקורם באפיתל וזיהוים זקוק להדמיה רגישה שאינה מוותרת על ערוץ העבודה. האביזר מאפשר הדמיה אופטית דו-מצבית מבלי לוותר על ערוץ העבודה. |  | 1 | [6 empty] | 1 |  | 1 | [16 empty] | 1 |  |  |  |  | UNIV DUKE¦UNIV NORTH CAROLINA CHAPEL HILL | Duke University¦University of North Carolina | THE UNIVERSITY OF NORTH CAROLINA AT CHAPEL HILL | University of North Carolina | University of North Carolina | 2022-01-31 | 2022 | 2024-12-24 | 2024 | 2021-01-29 | 2021 | Grant | Utility Invention | active | US202217589333 | 283842837 | US20220240782A1 | US
408 | EP4106625A1 | SYSTEM AND METHOD FOR GUIDING DIRECTION TO AND TREATING TARGETS FOR ABNORMAL BIOLOGICAL RHYTHMS | Ablation catheters for treating electrical rhythm disorders include an array of sensor electrodes for detecting electrical signals to determine the location of target areas for treatment. If the catheter is not optimally positioned at the target area, the controller uses the detected signals to guide movement of the catheter toward the target area. Once proper positioning is confirmed, the controller activates the ablation components within the catheter to deliver energy to modify tissue at the target region. | 1. Claims of equivalent WO 2021168380 A1 CLAIMS: 1. A system for treating a heart rhythm disorder, the device comprising: a catheter configured to be placed in contact with a tissue surface, the catheter comprising: a flexible body having a contact surface; an array of sensor electrodes arranged within the flexible body, each sensor electrode having a conductive surface substantially flush with the contact surface, each sensor electrode configured to detect electrical signals from the tissue surface; and one or more treatment elements configured to deliver energy to the tissue surface; a plurality of conductors, each conductor having a distal end connected to one of a sensor electrode and the one or more treatment element; a controller in communication with proximal ends of the plurality of conductors, the controller comprising a processor configured to: receive the detected electrical signals; determine a location of a target region of a heart rhythm disorder; determine whether the catheter is optimally positioned at the target region, and, if not optimally positioned, to compute directionality to the target region and generate movement instructions to move the catheter toward the target region; and after determining that the catheter is optimally positioned, generate treatment signals to activate the one or more treatment elements to modify tissue in the target region. 2. The system of claim 1, wherein the flexible layer is generally planar and has a shape selected from a group consisting of a rectangle, an ellipse, and an annulus. 3. The system of claim 1 or claim 2, further comprising an elongated hollow shaft having a distal end, a proximal end, and a length, wherein the catheter is disposed at the distal end, the controller is disposed at the proximal end, and the plurality of conductors is retained within and extends the length the shaft, wherein the distal end of the shaft is manipulable from the proximal end. 4. The system of claim 3, further comprising a shaft motor configured to steer the distal end of the shaft in response to movement instructions generated by the controller. 5. The system of claim 3 or claim 4, further comprising a sheath slidably disposed on the shaft, the sheath having an interior volume configured to retain the catheter in a folded condition until the catheter is deployed by sliding the sheath away from the distal end of the shaft. 6. The system of any one of claims claim 2- 5, further comprising irrigant pores formed in the flexible body, the irrigant pores in fluid communication with an irrigant source associated with the controller, wherein the irrigant source is configured to deliver irrigant through the irrigant pores to tissue at the target region in conjunction with activation of the array of treatment elements. 7. The system of any one of claims 1-6, wherein the one or more treatment elements comprise an array of ablation electrodes, and wherein a subset of the plurality of conductors connected to the one or more treatment elements are electrical conductors configured to deliver electromagnetic energy to each ablation electrode. 8. The system of claim 7, wherein the array of sensor electrodes and the array of ablation electrodes are uniformly distributed around the contact surface. 9. The system of claim 7, wherein ablation electrodes of the array of ablation electrodes are evenly interspersed among the array of sensor electrodes. 10. The system of claim 7 wherein the processor is further configured to: determine a size of the target region based on the detected electrical signals; identify one or more ablation electrodes of the array of ablation electrodes based on at least the size and the location of the target region; and activate the identified ablation electrodes. 11. The system of any one of claims 1-10, wherein the processor determines the location of the target region by: generating a directionality map of heart rhythms based on the detected electrical signals, the directionality map describing pathways of heart rhythms; generating a guidance direction in which to move the flexible body towards the target region, and integrating the directionality map to determine the location of the target region . 12. The system of claim 11, wherein the directionality map is generated by applying a trained machine learning model to the electrical signals, wherein the machine learning model is trained on training examples comprising electrical signals of a human heart and known target regions of the heart rhythm disorder. 13. The system of any one of claims 7-10, wherein each ablation electrode is configured to emit a distinct waveform. 14. The system of claim 13, wherein the controller is configured to separately address one or more subsets of ablation electrodes of the array, and wherein the treatment signals comprise a first signal to a first subset of ablation electrodes to emit a first waveform and a second signal to a second subset of ablation electrodes to emit a second waveform. 15. The system of any one of claims 1-14, wherein the array of sensor electrodes comprises at least four electrodes. 16. The system of any one of claims 1-15, wherein the sensor electrodes are configured to deliver ablation energy so that the one or more treatment elements comprise the array of sensor electrodes. 17. The system of any one of claims 1-5, wherein the one or more treatment elements comprise one or more coolant chambers formed within the flexible body and configured for retaining a coolant, and wherein the plurality of conductors comprises a subset of conductors configured to direct a coolant fluid from a coolant source to the one or more coolant chambers to deliver freezing energy to tissue at the target region. 18. The system of claim 17, wherein the flexible body has a thermally conductive material incorporated therein to enhance conduction of freezing energy to tissue in contact with the contact surface. 19. The system of any one of claims 1-5, wherein the one or more treatment elements comprises an array of cryoablation loci formed within the flexible body, and wherein the plurality of conductors comprises a subset of conductors configured to direct a coolant fluid from a coolant source to the cryoablation loci in response to treatment signals from the controller to deliver freezing energy to tissue at the target region. 20. The system of any one of claims 1-6, wherein the one or more treatment elements comprise an array of targeting fiducials distributed within the flexible body, the targeting fiducials configured for guiding delivery of ablation energy from one or more external ablation energy sources. 21. The system of any one of claims 1-20, further comprising a contact sensor configured to determine whether the contact surface and the tissue surface are in adequate contact and, if not, to provide a signal to the controller to guide movement of the flexible body to provide improved contact. 22. A method for treating a heart rhythm disorder, the method comprising: detecting electrical signals of a heart using the array of sensor electrodes of the system of any one of claims 1-21; generating a directionality map describing pathways of heart rhythms based on the detected electrical signals; integrating the directionality map to determine one of: (i) a location of a target region of the heart rhythm disorder in the directionality map, and(ii) a guidance direction to the target region of the heart rhythm disorder outside the directionality map; determining whether the flexible body is optimally positioned at the target region according to the directionality map; and responsive to determining optimal positioning, activating the one or more treatment elements of the system to deliver energy to modify tissue at the determined location of the target region. 23. The method of claim 22, wherein generating the directionality map comprises applying a trained machine learning model to the detected electrical signals, wherein the machine learning model is trained on training examples comprising electrical signals of a human heart and known locations of one or more target regions of a heart rhythm disorder. 24. The method of claim 22, further comprising: responsive to determining the guidance direction to the target region outside the directionality map, steering the device to a subsequent position in the guidance direction to the target region. 25. The method of claim 22, further comprising: detecting subsequent electrical signals of the heart with the plurality of sensing electrodes after steering the device to the subsequent position; generating a subsequent directionality map describing pathways of heart rhythms based on the subsequent electrical signals; integrating the subsequent directionality map to determine one of: (i) the location of a target region of the heart rhythm disorder in the subsequent directionality map, and(ii) a subsequent guidance direction to the target region of the heart rhythm disorder outside the subsequent directionality map; and responsive to determining the subsequent guidance direction to the source region outside the directionality map, steering the flexible body to a third position in the subsequent guidance direction to the target region. 26. The method of any of claim 22, further comprising, responsive to determining the direction of the target region outside the directionality map, providing a notification on an electronic display to move the device in the direction of the target region. 27. The method of claim 22, further comprising responsive to determining the location of the target region, identifying the one or more ablation components within a threshold proximity to the location of the target region. 28. The method of claim 27, further comprising, responsive to determining the location of the target region determining a size of the target region based on the directionality map, wherein identifying the one or more ablation components is further based on at least the size of the target region. 29. The method of claim 22, wherein each ablation component is an ablation electrode, wherein instructing the one or more ablation components comprises actuating one or more ablation electrodes to emit one or more electromagnetic waves that modify tissue at the target region of the heart rhythm disorder. 30. The method of claim 29, wherein each ablation electrode is adjustable to generate a distinct waveform, wherein activating the one or more ablation electrodes includes: instructing a first ablation electrode to generate a first waveform; and instructing a second ablation electrode to generate a second waveform. 31. The method of claim 22, wherein each ablation component is a cryoablation loci, wherein instructing the one or more ablation components comprises directing coolant to one or more cryoablation loci that modify tissue at the target region of the heart rhythm disorder. 32. The method of claim 22, further comprising, responsive to determining the location of the target region, directing coolant into a chamber on the device cooling a portion of the heart inclusive of the target region. 33. The method of claim 22, further comprising, responsive to determining the location of the target region, venting irrigant from one or more irrigation pores on the device onto the heart to cool a portion of the heart inclusive of the target region. 34. The method of claim 22, further comprising detecting subsequent electrical signals of the heart with the plurality of sensing electrodes after ablation of the target region; and determining whether the heart rhythm disorder persists based on the subsequent electrical signals. 35. The method of claim 22, further comprising, responsive to determining that the heart rhythm disorder persists: generating a subsequent directionality map describing pathways of heart rhythms based on the subsequent electrical signals; integrating the subsequent directionality map to determine one of: a location of a second target region of the heart rhythm disorder in the subsequent directionality map, and a guidance direction to the second target region of the heart rhythm disorder outside the subsequent directionality map; and responsive to determining the location of the second target region, instructing one or more ablation components of a plurality of ablation components on the device to modify tissue at the second target region at the determined location of the second target region. | 2020 | Stanford University | University | NA | https://www.stanford.edu | USA | The patent focuses on a cardiac-rhythm ablation catheter whose sensor electrode array detects electrical signals to find target regions and guides catheter movement toward them. It addresses the difficulty operators face in locating arrhythmia targets and confirming proper catheter positioning during ablation. Once positioned, the controller activates onboard ablation components to treat the rhythm disorder. | הפטנט מתמקד בקטטר אבלציה לקצב לב שמערך אלקטרודות החיישן שלו מזהה אותות חשמליים למציאת אזורי מטרה ומנחה את תנועת הקטטר אליהם. הוא מתמודד עם הקושי שמפעילים נתקלים בו באיתור מטרות הפרעות קצב ואישור מיקום קטטר נכון במהלך אבלציה. לאחר המיקום, הבקר מפעיל רכיבי אבלציה מובנים לטיפול בהפרעת הקצב. | 1 | [13 empty] | 1 | [7 empty] | 1 | [7 empty] | 1 |  | UNIV LELAND STANFORD JUNIOR | Stanford University | UNIV LELAND STANFORD JUNIOR | Stanford University | Stanford University | 2021-02-20 | 2021 | 2022-12-28 | 2022 | 2020-02-20 | 2020 | Application | Utility Invention | active | EP20210756920 | 318664107 | AU2021224243A1¦AU2021224243B2¦CA3171471A1¦CN115768346A¦EP4106625A4¦IL295739A¦IL295739B1¦JP2023517284A¦JP7836763B2¦US12193728B2¦US20210259765A1¦WO2021168380A1 | EP
409 | CN121015128A | Bimodal endoscopic probe with OCT (optical coherence tomography) imaging and plasma ablation functions | The invention provides a dual-mode endoscopic probe with OCT imaging and plasma ablation. The dual-mode endoscopic probe comprises a probe shell, a plasma ablation electrode device and an imaging device. The probe shell is provided with an imaging channel and an ablation channel which are arranged in parallel, the imaging device is arranged in the imaging channel, and the plasma ablation electrode device is arranged in the ablation channel; a coaxial needle type plasma ablation electrode is coupled with OCT imaging, plasma ablation is a minimally invasive precise ablation technology and is matched with optical coherence tomography, tomography is carried out on polyp and surrounding intestinal wall tissue before and after excision, the excision depth is judged, proper plasma excitation voltage and frequency are set, and excision time is shortened. Complications caused by incomplete polyp cutting or excessive polyp cutting are avoided, and accurate ablation cutting is achieved. | 1 A dual-modal endoscopic probe with OCT imaging and plasma ablation capabilities is characterized by comprising a probe housing, a plasma ablation electrode device, and an imaging device; the probe housing has parallelly arranged imaging channels and ablation channels, the imaging device is disposed within the imaging channel, and the plasma ablation device is disposed within the ablation channel; the imaging device includes a gradient refractive index lens, an optical fiber, a piezoelectric bicrystalline wafer, and a fixing device; the gradient refractive index lens is disposed at the distal end of the imaging channel; the optical fiber is located behind the gradient refractive index lens and is fixed to the centerline of the upper surface of the piezoelectric bicrystalline wafer, with one end of the optical fiber extending out of the surface of the piezoelectric bicrystalline wafer towards the gradient refractive index lens to form a cantilever beam structure; the piezoelectric bicrystalline wafer is fixed to the proximal end of the imaging channel by the fixing device. 2 According to claim 1, the dual-modal endoscopic probe with OCT imaging and plasma ablation is characterized in that the plasma ablation device includes a movable electrode, a heat shrink tubing, and a loop electrode; the outer layer of the movable electrode is the heat shrink tubing, and the outer layer of the heat shrink tubing wraps around the loop electrode; wherein, the end of the movable electrode extends beyond one end of the heat shrink tubing and bends towards the imaging side. 3 According to claim 1, the dual-modal endoscopic probe with OCT imaging and plasma ablation is characterized in that the gradient refractive index lens has an 8° end face and a gradient index of 0.327 mm.1。 4 The dual-mode endoscopic probe with OCT imaging and plasma ablation as described in claim 1 is characterized in that the optical fiber is a single-mode optical fiber, and the end face of the optical fiber has an 8° bevel. 5 According to claim 1, the dual-modal endoscopic probe with OCT imaging and plasma ablation is characterized in that the piezoelectric bicrystalline wafer is composed of two piezoelectric ceramics with the same polarization direction stacked together. 6 According to claim 1, the dual-modal endoscopic probe with OCT imaging and plasma ablation is characterized in that the fixing device consists of two opposing fixing plates, the inner surfaces of the two fixing plates are provided with slots, the slots are adapted to the sides of the piezoelectric bicrystalline wafer, and the outer surfaces of the two fixing plates are arc surfaces adapted to the inner surface of the imaging channel; the two sides of the piezoelectric bicrystalline wafer are sandwiched between the two fixing plates, and the piezoelectric bicrystalline wafer is adhered to the imaging channel by optical adhesive. 7 According to claim 1, the dual-modal endoscopic probe with OCT imaging and plasma ablation is characterized in that the probe housing is made of medical ABS material, and the junction of the gradient refractive index lens and the probe housing is coated with ultraviolet curing adhesive. 8 The dual-modal endoscopic probe with OCT imaging and plasma ablation according to claim 1 is characterized in that the surface of the gradient refractive index lens is coated with an anti-interference coating of inert tin oxide. | 2025 | University of Shanghai for Science and Technology | University | NA | https://www.usst.edu.cn | China | The patent focuses on a dual-mode endoscopic probe combining OCT imaging and plasma ablation in parallel channels of the housing, with a coaxial needle-type plasma electrode for ablation and an OCT probe for imaging. It addresses the high prevalence of colorectal polyps that progress to cancer when removed without precise visualization. The combined design enables minimally invasive imaging-guided plasma ablation of polyps. | הפטנט מתמקד בגשוש אנדוסקופי דו-מצבי המשלב הדמיית OCT ואבלציית פלזמה בערוצים מקבילים של הדיור, עם אלקטרודת פלזמה מסוג מחט קואקסיאלית לאבלציה וגשוש OCT להדמיה. הוא מתמודד עם השכיחות הגבוהה של פוליפים קולורקטליים המתפתחים לסרטן כשמסירים אותם ללא הדמיה מדויקת. העיצוב המשולב מאפשר אבלציית פלזמה זעיר-פולשנית מונחית הדמיה של פוליפים. |  | 1 | [16 empty] | 1 | [8 empty] | 1 |  |  |  |  | UNIV SHANGHAI SCIENCE & TECH | University of Shanghai for Science and Technology | UNIV SHANGHAI SCIENCE & TECH | University of Shanghai for Science and Technology | University of Shanghai for Science and Technology | 2025-07-18 | 2025 | 2025-11-28 | 2025 | 2025-07-18 | 2025 | Application | Utility Invention | active | CN202510990859 | 324781042 | [NA] | CN
410 | CN119326492A | Control system for ablation of duodenum through pulsed electric field | The invention relates to the field of medical physics, and discloses a control method and system for pulse electric field ablation of duodenum, and the method comprises the steps: determining a multi-target tissue, needing pulse electric field ablation, of the duodenum corresponding to a target user, collecting a target tissue image of the multi-target tissue, and constructing a multi-tissue three-dimensional model of the multi-target tissue; analyzing tissue parameters of the multi-target tissue, and analyzing electric field intensity and pulse parameters required for ablation of the multi-target tissue; multiple electrodes of the ablation equipment are configured, and multiple electrode implantation sites of the multiple electrodes are determined; constructing a digestive tract model from the corresponding oral cavity to the duodenum of a target user, analyzing an ideal transmission path of multiple electrodes, analyzing a path deviation coefficient of a simulation transmission path and the ideal transmission path, correcting the simulation transmission path, transmitting the multiple electrodes to a multi-electrode implantation site, and then performing multi-electrode implantation. And pulse electric field ablation of multi-target tissues is realized by utilizing multiple electrodes. The precise ablation effect of the pulsed electric field can be improved. | 1.A control method for pulsed electric field ablation of the duodenum, characterized in that the method comprises: determining multiple target tissues of the duodenum corresponding to the target user that need to be subjected to pulsed electric field ablation, acquiring target tissue images of the multiple target tissues, registering the target tissue images to obtain registered images, and constructing a multi-tissue three-dimensional model of the multiple target tissues based on the registered images; analyzing the tissue types of the multiple target tissues, analyzing the tissue parameters of the multiple target tissues based on the tissue types, and analyzing the electric field strength and pulse parameters required for ablating the multiple target tissues based on the tissue parameters; acquiring an ablation device for the pulsed electric field ablation, configuring multiple electrodes of the ablation device based on the electric field strength and tissue parameters, and determining multiple electrode implantation sites of the multiple electrodes based on the multi-tissue three-dimensional model; constructing a digestive tract model from the mouth to the duodenum corresponding to the target user, analyzing the ideal transmission path of the multiple electrodes based on the digestive tract model, constructing a mechanical electrode catheter of the multiple electrodes, acquiring the dynamic parameters of the mechanical electrode catheter in real time, and performing path simulation on the mechanical electrode catheter using a preset path simulation algorithm based on the dynamic parameters to obtain a simulated transmission path;The path deviation coefficients of the simulated transmission path and the ideal transmission path are analyzed, and based on the path deviation coefficients, the simulated transmission path is corrected to obtain a corrected simulated transmission path. When the corrected simulated transmission path is consistent with the ideal transmission path, based on the mechanical electrode catheter, after the multi-electrode is transmitted to the multi-electrode implantation site, the multi-electrode is used to achieve pulse electric field ablation of multiple target tissues according to the pulse parameters. 2.The control method of pulsed electric field ablation of the duodenum as described in claim 1 is characterized in that, the registering of the target tissue image to obtain a registered image includes: preprocessing the target tissue image to obtain a preprocessed tissue image; analyzing the key indicators of the preprocessed tissue image, and dividing the preprocessed tissue image into a reference image and a source image based on the key indicators; constructing an image transformation model of the source image, analyzing the transformation parameters of the image transformation model, and transforming the source image based on the transformation parameters to obtain a transformed source image; calculating the registration coefficient of the transformed source image and the reference image; when the registration coefficient is less than a preset registration coefficient threshold, optimizing the transformation parameters to obtain optimized transformation parameters, and when the optimized transformation parameters are greater than or equal to a preset registration coefficient threshold, using the optimized transformation parameters as optimal transformation parameters; based on the optimal transformation parameters, transforming the source image to obtain an optimal transformed original image; aligning the optimal source image and the reference image to obtain a registered image. 3.The control method for pulsed electric field ablation of the duodenum as described in claim 1 is characterized in that the analysis of the electric field strength and pulse parameters required for ablating the multiple target tissues based on the tissue parameters includes: extracting the multiple target tissue radius, multiple target tissue thickness, multiple target tissue dielectric constant and multiple target tissue conductivity of the multiple target tissues based on the tissue parameters; determining the electric field strength required for ablating the multiple target tissues through the multiple target tissue dielectric constant and multiple target tissue conductivity; calculating the pulse voltage required for ablating the multiple target tissues based on the electric field strength, multiple target tissue radius, multiple target tissue thickness, multiple target tissue dielectric constant and multiple target tissue conductivity; identifying the ablation requirements of the multiple target tissues, and determining the pulse width, pulse frequency and pulse sequence of the multiple target tissues based on the ablation requirements and the pulse voltage; The pulse parameters required for ablating the multiple target tissues are determined based on the pulse voltage, pulse width, pulse frequency and pulse sequence. 4.The control method for pulsed electric field ablation of the duodenum as described in claim 1 is characterized in that configuring the multiple electrodes of the ablation device based on the electric field strength and tissue parameters includes: determining the electrode material, electrode size and electrode shape of the ablation device based on the tissue parameters; determining the number of electrodes of the ablation device according to the number of multiple target tissues corresponding to the tissue parameters; configuring the electrode sensor of the ablation device, and monitoring the real-time electric field strength and real-time temperature of the multiple target tissues based on the electrode sensor; constructing an electrode safety mechanism of the ablation device through the real-time electric field strength and real-time temperature; configuring the multiple electrodes of the ablation device according to the electrode material, electrode size, electrode shape, electrode number and electrode safety mechanism. 5.The control method for pulsed electric field ablation of the duodenum as described in claim 1 is characterized in that determining the multi-electrode implantation site of the multi-electrode based on the multi-tissue three-dimensional model includes: analyzing the three-dimensional state of the multi-tissue three-dimensional model; determining the multi-electrode implantation range of the multi-electrode based on the three-dimensional state; identifying the multi-electrode initial implantation point of the multi-electrode implantation range; analyzing the straight-line distance between the multi-electrode initial implantation point and the multi-target tissues corresponding to the multi-tissue three-dimensional model according to the three-dimensional state; calculating the superimposed electric field strength of the multi-electrode based on the straight-line distance; if the superimposed electric field strength does not meet the preset electric field strength threshold, adjusting the implantation point position of the multi-electrode initial implantation point, and returning to the above-mentioned superimposed electromagnetic strength calculation step; if the superimposed electric field strength meets the preset electric field strength threshold, using the multi-electrode initial implantation point as the multi-electrode implantation site. 6.The control method for pulsed electric field ablation of the duodenum as described in claim 1 is characterized in that the construction of the digestive tract model from the mouth to the duodenum of the target user includes: using a preset medical imaging technology to collect image data from the mouth to the duodenum of the target user; performing data cleaning on the image data to obtain cleaned image data; registering the cleaned image data to obtain registered image data; based on the registered image data, extracting geometric structural features from the mouth to the duodenum of the target user; and constructing the digestive tract model from the mouth to the duodenum of the target user according to the geometric structural features. 7.The control method for pulsed electric field ablation of the duodenum as described in claim 1 is characterized in that the construction of the multi-electrode mechanical electrode catheter includes: obtaining simulation test results of the multi-electrode in the corresponding digestive tract model, and analyzing the performance requirements of the mechanical electrode catheter according to the simulation test results; determining the catheter structure for constructing the mechanical electrode catheter based on the performance requirements, wherein the catheter structure includes: a catheter geometric frame and a catheter material; integrating the microcircuit and sensor of the mechanical electrode catheter; constructing a real-time monitoring mechanism of the mechanical electrode catheter based on the microcircuit and the sensor; arranging the transmission device of the mechanical electrode catheter, and constructing a driving mechanism of the mechanical electrode catheter based on the transmission device. Advance mechanism; Construct the multi-electrode electrode mechanical catheter according to the catheter structure, real-time monitoring mechanism and advancement mechanism. 8.The control method for pulsed electric field ablation of the duodenum as described in claim 1 is characterized in that, based on the dynamic parameters, a preset path simulation algorithm is used to perform path simulation on the mechanical electrode catheter to obtain a simulated transmission path, including: extracting the current position vector, current velocity vector and current acceleration vector of the mechanical electrode catheter based on the dynamic parameters; calculating the future position vector of the mechanical electrode catheter based on the current position vector, current velocity vector and current acceleration vector; and performing path simulation on the mechanical electrode catheter based on the previous position vector and the future position vector using a preset path simulation algorithm to obtain a simulated transmission path. 9.The control method for pulsed electric field ablation of the duodenum as described in claim 1 is characterized in that the analysis of the path deviation coefficient of the simulated transmission path and the ideal transmission path includes: taking the same starting point corresponding to the simulated transmission path and the ideal transmission path as the origin of the three-dimensional coordinate system; based on the origin of the three-dimensional coordinate system, integrating the simulated transmission path and the ideal transmission path in the same three-dimensional coordinate system to obtain a three-dimensional curve of the simulated path and a three-dimensional curve of the ideal path; calculating the three-dimensional coordinate difference between the three-dimensional curve of the simulated path and the three-dimensional curve of the ideal path; calculating the sum of the absolute values of the three-dimensional coordinate differences; constructing a position deviation curve of the sum of the absolute values of the differences; analyzing the fluctuation characteristics of the position deviation curve; and calculating the path deviation coefficient of the simulated transmission path and the ideal transmission path based on the fluctuation characteristics. 10.A control system for pulsed electric field ablation of the duodenum, characterized in that it is used to execute the control method for pulsed electric field ablation of the duodenum as described in any one of claims 1 to 9, and the system includes: a tissue model construction module, used to determine the multi-target tissues that need to be subjected to pulsed electric field ablation in the duodenum of the target user, collect target tissue images of the multi-target tissues, register the target tissue images to obtain registered images, and construct a multi-tissue three-dimensional model of the multi-target tissues based on the registered images; an ablation parameter determination module, used to analyze the tissue types of the multi-target tissues, analyze the tissue parameters of the multi-target tissues based on the tissue types, and analyze the electric field strength and pulse parameters required for ablating the multi-target tissues based on the tissue parameters; an electrode implantation position confirmation module, used to obtain the ablation device for the pulsed electric field ablation, configure the multi-electrodes of the ablation device based on the electric field strength and tissue parameters, and determine the multi-electrode implantation sites of the multi-electrodes based on the multi-tissue three-dimensional model; The pulse electric field ablation module is used to construct a digestive tract model from the mouth to the duodenum of the target user, analyze the ideal transmission path of the multi-electrode based on the digestive tract model, and construct a mechanical electrode catheter for the multi-electrode.The dynamic parameters of the mechanical electrode catheter are collected in real time, and based on the dynamic parameters, a preset path simulation algorithm is used to simulate the path of the mechanical electrode catheter to obtain a simulated transmission path, and a path deviation coefficient between the simulated transmission path and the ideal transmission path is analyzed. Based on the path deviation coefficient, the simulated transmission path is corrected to obtain a corrected simulated transmission path. When the corrected simulated transmission path is consistent with the ideal transmission path, after the multi-electrode is transmitted to the multi-electrode implantation site based on the mechanical electrode catheter, the multi-electrode is used to implement the multi-electrode implantation according to the pulse parameters. Pulsed electric field ablation of target tissue. | 2024 | UNIV TIANJIN MEDICAL | University | NA | https://www.tmu.edu.cn | China | The patent focuses on a duodenal PFA control method that picks multiple target tissues, builds a 3D multi-tissue model from images, and analyzes parameters across them. It addresses the difficulty of precisely controlling pulsed-field ablation across a multi-target zone in the duodenum to reduce damage to surrounding healthy tissue. The model supports staged, parameter-tuned PFA across the duodenum. | הפטנט מתמקד בשיטת בקרת PFA תריסריונית הבוחרת מספר רקמות מטרה, בונה מודל רב-רקמתי תלת-ממדי מתמונות ומנתחת פרמטרים לרוחבן. הוא מתמודד עם הקושי בבקרה מדויקת של אבלציה בשדה פולסי על פני אזור רב-מטרתי בתריסריון להפחתת נזק לרקמה בריאה סובבת. המודל תומך ב-PFA מדורג ומכוונן בפרמטרים על פני התריסריון. | 1 | [7 empty] | 1 |  |  | 1 | [15 empty] | 1 |  |  |  |  | UNIV TIANJIN MEDICAL | UNIV TIANJIN MEDICAL | UNIV TIANJIN MEDICAL | UNIV TIANJIN MEDICAL | UNIV TIANJIN MEDICAL | 2024-12-13 | 2024 | 2025-01-21 | 2025 | 2024-12-13 | 2024 | Application | Utility Invention | active | CN202411840143 | 318702590 | [NA] | CN
411 | WO2022182876A1 | TREATMENT PLANNING FOR ATRIAL FIBRILLATION USING PATIENT-SPECIFIC MODELS BASED ON QUANTIFICATION OF FIBRILLATORY AREA | Systems and methods for developing an implementing an ablation plan for treating atrial fibrillation. A 3D digital model of cardiac tissue of a patient is segmented into a plurality of different sections and a fibrillatory area of each section is quantified. The fibrillatory area is based on a surface area of the section and a weighted surface area of each adjacent section, wherein the weighting is based on a proportion of a boundary between the section and the adjacent section that is not occupied by ablative scar tissue. An ablation plan is then defined based on the determining fibrillatory area of each section by adding planned ablation areas on the boundaries between sections until the quantified fibrillatory area of each section is below a maximum fibrillatory area threshold. The ablation plan is then displayed on a display screen to guide an ablation procedure according to the ablation plan. | 1. CLAIMSWhat is claimed is: 1. A method comprising: segmenting a 3D digital model of cardiac tissue of a patient into a plurality of different sections; quantifying, by an electronic controller, a fibrillatory area of each section of the plurality of different sections, wherein the fibrillatory area is quantified based at least in part on a surface area of the section and a weighted surface area of each adjacent section, wherein the weighted surface area of each adjacent section is calculated based on the surface area of the adjacent section and a proportion of a boundary between the section and the adjacent section that is not occupied by ablative scar tissue; defining an ablation plan adding at least one planned ablation area on the boundary between the section and the adjacent section until the quantified fibrillatory area of each section of the plurality of sections is below a maximum fibrillatory area threshold; and displaying the ablation plan overlaid onto the 3D digital model of the cardiac tissue on a display screen during performance of an ablation procedure according to the ablation plan. 2. The method of claim 1, further comprising performing the ablation procedure according to the ablation plan. 3. The method of claim 1, wherein defining the ablation plan further includes: identifying a first section of the plurality of sections with a largest quantified fibrillatory area, identifying a second section of the plurality of sections, wherein the second section is adjacent to the first section and has a larger fibrillatory area than any other section adjacent to the first section, and adding a first planned ablation area to the ablation plan, wherein the first planned ablation area is positioned along a boundary between the first section and the second section at a location where ablative scar tissue is not currently present. 4. The method of claim 3, wherein defining the ablation plan further includes: quantifying a first updated fibrillatory area for the first section, wherein the first updated fibrillatory area is quantified based at least in part on the surface area of the first section and a weighted surface area of each adjacent section, wherein the weighted surface area of the second section is calculated based on the surface area of the second section and a proportion of the boundary between the first section and the second section that is not covered with ablative scar tissue and is not covered by the first planned ablation area; and adding a second planned ablation area to the ablation plan in response to determining that the first updated fibrillatory area for the first section exceeds the maximum fibrillatory area threshold. 5. The method of claim 1, wherein defining the ablation plan further includes iteratively adding additional planned ablation areas on boundaries between the different sections until the quantified fibrillatory area of each section of the plurality of sections is below the maximum fibrillatory area. 6. The method of claim 1, wherein quantifying the fibrillatory area of each section of the plurality of sections included calculating the fibrillatory area of each section as wherein IΆ, is the fibrillatory area of a section wherein EAi is an effective area of the section wherein m is a number of sections adjacent to the section wherein g,, is a sum of lengths of a shared boundary between the section i and an adjacent section j that is not occupied by ablative scar tissue or a planned ablation area, wherein is a total length of the shared boundary between the section i and the adjacent section j, wherein EAj is an effective area of the adjacent section /, and wherein the effective area is a surface area of the section that is not occupied by ablative scar tissue or any planned ablation area. 7. The method of claim 1, further comprising: capturing image data of the cardiac tissue of the patient; generating the 3D digital model of the cardiac tissue of the patient based at least in part on the captured image data; and automatically detecting, by an electronic processor, locations of ablative scar tissue on the 3D digital model. 8. The method of claim 7, wherein capturing the image data includes capturing the image data of the cardiac tissue using late-gadolinium-enhancement (LGE) MRI, and wherein automatically detecting the locations of the ablative scar tissue on the 3D digital model includes automatically detecting locations highlighted by the LGE MRI. 9. The method of claim 7, wherein segmenting the 3D digital model of the cardiac tissue of the patient into the plurality of different sections includes positioning boundaries between the different sections to align with the detected locations of the ablative scar tissue in the 3D digital model. 10. The method of claim 1, further comprising: measuring an effective refractory period (ERP) of the patient; and determining the maximum fibrillatory area threshold based at least in part on the measured effective refractory period. 11. The method of claim 1, further comprising: applying a computer-based simulation of electrical activity to the 3D digital model of the cardiac tissue of the patient, wherein the fibrillatory area of one or more sections are adjusted during the computer-based simulation; and determining the maximum fibrillatory area threshold based at least in part on a minimum fibrillatory area required for observed rotor activity in the computer-based simulation. 12. The method of claim 1, wherein segmenting the 3D digital model of the cardiac tissue of the patient into the plurality of different sections includes segmenting the 3D digital model of a left atrium of the patient based on anatomical landmarks on the 3D model, wherein the plurality of different sections includes a section surrounding the left sided pulmonary veins, a section surrounding the right sided pulmonary veins, a section below the pulmonary veins in a lower part of a posterior wall of the left atrium, a section including an anterior wall of the left atrium, and a section including an appendage of the left atrium. 13. The method of claim 1, wherein no ablative scarring from a prior ablative procedure is present on the cardiac tissue of the patient, and wherein defining the ablation plan includes: defining an initial ablation plan including a plurality of planned ablation areas along the boundaries between the different sections according to a predefined initial ablation plan, and iteratively adding the at least one planned ablation area until the quantified fibrillatory area of each section of the plurality of sections is below the maximum fibrillatory area threshold. 14. The method of claim 13, wherein defining the initial ablation plan includes defining an initial ablation plan that includes planned ablation areas surrounding pulmonary veins of a left atrium of the patient and at least one additional planned ablation area on a border between two other sections of the plurality of different sections. 15. A computer-based system for developing and implementing an ablation plan for an ablation procedure to treat atrial fibrillation, the computer-based system comprising: a display screen; and an electronic controller configured to segment a 3D digital model of cardiac tissue of a patient into a plurality of different sections, quantify a fibrillatory area of each section of the plurality of different sections, wherein the fibrillatory area is quantified based at least in part on a surface area of the section and a weighted surface area of each adjacent section, wherein the weighted surface area of each adjacent section is calculated based on the surface area of the adjacent section and a proportion of a boundary between the section and the adjacent section that is not covered with ablative scar tissue, define an ablation plan adding at least one planned ablation area on the boundary between the section and the adjacent section until the quantified fibrillatory area of each section of the plurality of sections is below a maximum fibrillatory area threshold, and display the ablation plan overlaid onto the 3D digital model of the cardiac tissue on the display screen during performance of an ablation procedure according to the ablation plan | 2021 | University Of Utah Research Foundation | University | NA | https://tto.utah.edu | USA | The patent focuses on an atrial-fibrillation ablation planning system that segments a 3D cardiac model into sections and quantifies a fibrillatory area in each, weighted by adjacent sections. It addresses the variability of atrial fibrillation recurrence after ablation, where generic plans don't account for patient-specific fibrillatory patterns. The fibrillatory-area map produces patient-specific ablation plans. | הפטנט מתמקד במערכת תכנון אבלציה לפרפור פרוזדורים המפלחת מודל לבבי תלת-ממדי לקטעים ומכמתת אזור פיברילטורי בכל אחד, משוקלל לפי קטעים סמוכים. הוא מתמודד עם השונות של הישנות פרפור פרוזדורים לאחר אבלציה, שבה תוכניות גנריות אינן מתחשבות בתבניות פיברילטוריות ספציפיות למטופל. מפת האזור הפיברילטורי מייצרת תוכניות אבלציה ספציפיות למטופל. | [8 empty] | 1 | [12 empty] | 1 |  |  | 1 | [7 empty] | UNIV UTAH RES FOUND | University Of Utah Research Foundation | UNIV UTAH RES FOUND | University Of Utah Research Foundation | University Of Utah Research Foundation | 2022-02-24 | 2022 | 2022-09-01 | 2022 | 2021-02-25 | 2021 | Application | Utility Invention | inactive | WO2022US17705 | 284618258 | [NA] | WO
412 | DE112024002177T5 | MONITORING ABLATION PROGRESS USING LINEAR EBUS DATA | A system and method are presented for treating targeted tissue using cryoablation. A linear endobronchial ultrasound (EBUS) device is positioned adjacent the targeted tissue and is used to guide a percutaneously inserted cryoprobe into the targeted tissue. The EBUS device is partially rotated in order to create multiple image slices that are combined together. To overcome the acoustic shadow created by the ice ball, a 3D model is created of the ice ball based on the locations identified on visible portion of the ice ball's peripheral surface. These locations are mirrored across an axis defined for the cryoprobe to define an approximate location for the hidden, non-visible periphery. This model is then displayed along with the location of the targeted tissue and an indication of the killing zone of the ice ball defined by a selected isotherm. | 1. Method for treating a patient, comprising: a) placing an ultrasound device in a tracheal tree of the patient in a position to image a target tissue; b) placing a cryoprobe in the target tissue; c) creating an ice ball by cooling the cryoprobe; and d) during the creation of the ice ball: i) rotating the ultrasound device to create multiple 2D image slices of the target tissue, ii) combining the multiple 2D image slices using a computer to identify a visible surface of the ice sphere, iii) creating a 3D model of the ice ball based on the visible surface using the computer, iv) generating an isotherm for the 3D model at a selected temperature identified for effective ablation using the computer, and v) Displaying the 3D model of the ice sphere and the isotherm on a user interface using the computer. 2. Procedure according toClaim 1, where the ultrasound device is a linear endobronchial ultrasound device (EBUS). 3. Procedure according toClaim 2, wherein the linear EBUS device comprises a plurality of ultrasonic transducer elements; further wherein the plurality of ultrasonic transducer elements are of a type selected from a group consisting of PZT-based transducers, pMUT-based transducers and cMUT-based transducers. 4. Procedure according toClaim 1, which further includes displaying a representation of the target tissue relative to the 3D model of the ice sphere and the isotherm on the user interface. 5. Procedure according toClaim 1, further comprising comparing the 3D model of the ice ball on the computer with a known size and shape of the target tissue to identify portions of the target tissue outside of the isotherm; and displaying the identified portions of the target tissue on the user interface using an identifiable distinguishing visual feature. 6. Procedure according toClaim 1, in which a distal end of the ultrasound device further includes electromagnetic sensors that receive electromagnetic signals that locate the distal end in an electromagnetic field, and further comprising using the electromagnetic signals to display the 3D model of the ice ball on a representation of the target tissue on the user interface. 7. Procedure according toClaim 1, where the 3D model of the ice sphere and the isotherm displayed on the interface change over time in response to an increase in the size of the ice sphere. 8. Procedure according toClaim 1, further comprising terminating the generation of the ice sphere before an expected completion time when the user interface indicates that the isotherm includes the target tissue. 9. Procedure according toClaim 1, which further includes continuing to generate the ice ball beyond an expected end time when the user interface indicates that the isotherm does not enclose the target tissue. 10. Procedure according toClaim 1, where creating the 3D model includes: i) identifying an axis for the cryoprobe, ii) identifying 3D positions on the visible surface of the ice sphere, iii) mirroring the identified 3D positions across the axis for the cryoprobe to identify points on a hidden surface of the ice sphere, and iv) Combining the 3D positions on the visible surface with the identified points on the hidden surface. 11. Procedure according toClaim 10, where the reflection of the identified 3D positions across the axis for the cryoprobe occurs perpendicular to the axis and the points are the same distance from the axis as the identified 3D positions. 12. Procedure according toClaim 10, where only a selected number of 3D positions are reflected and the 3D model approximates positions between the selected number of 3D positions on the visible surface and the points on the hidden surface. 13. Procedure according toClaim 1, where generating the isotherm for the 3D model at the selected temperature includes: i) identifying a known distance from a circumference of the ice ball known to be at the selected temperature or colder; ii) Identify the isotherm located at this known distance from the circumference of the ice ball. 14. Procedure according toClaim 13, where the known distance is determined by testing temperatures in cryoablation tests. 15. The procedure afterClaim 13, further requiring that the selected temperature or a colder temperature be maintained for a specified period of time. 16. Procedure according toClaim 15, where the specified period of time is at least three minutes. 17. A system comprising: a) a cryoprobe; b) an ultrasound machine; c) an advertisement; d) a computer system that communicates with the cryoprobe, the ultrasound device and the display, the computer system having a processor operating under programming, the programming causing the computer system to: i) receiving 2D image slices of target tissue in a patient from the ultrasound device as it rotates within the patient, the 2D image slices including representations of an ice ball formed by the cryoprobe, ii) combining the 2D image slices to identify a visible surface of the ice ball, iii) creating a 3D model of the ice ball based on the visible surface, iv) generating an isotherm for the 3D model of the ice ball at a selected temperature identified for effective ablation, and v) Creating a user interface on the display that displays the 3D model of the ice sphere and the isotherm. 18. The system afterClaim 17, wherein the computer system further includes a model of the target tissue registered to the patient, and wherein the user interface displays a representation of the target tissue relative to the 3D model of the ice ball and isotherm. 19. The system afterClaim 18, whereby the programming also causes the computer system compares the 3D model of the ice sphere with a known size and shape of the target tissue to identify portions of the target tissue outside the isotherm; and further that the identified portions of the target tissue are displayed using an identifiable distinguishing visual feature. 20. The system afterClaim 17, where the programming causes the computer system to create the 3D model of the ice ball by causing the computer system to: i) identify an axis for the cryoprobe, ii) identify 3D positions on the visible surface of the ice sphere, iii) mirror the identified 3D positions across the axis for the cryoprobe to identify points on a hidden surface of the ice sphere, and iv) combine the 3D positions on the visible surface with the identified points on the hidden surface. 21. The system afterClaim 20, where the identified 3D positions transverse to the axis of the cryoprobe are mirrored perpendicular to the axis and the points are the same distance from the axis as the identified 3D positions. 22. The system afterClaim 20, where only a selected number of 3D positions are mirrored and the 3D model of the ice ball approximates positions between the selected number of 3D positions on the visible surface and the points on the hidden surface. 23. The system afterClaim 17, where generating the isotherm for the 3D model at the selected temperature includes: i) identifying a known distance from a circumference of the ice ball known to be at the selected temperature or colder; ii) Identify the isotherm located at this known distance from the circumference of the ice ball. 24. The system afterClaim 23, where the known distance is determined by testing temperatures in cryoablation tests. 25. The system afterClaim 23, further requiring that the selected temperature or a colder temperature be maintained for a certain period of time. 26. The system afterClaim 25, where the specified period of time is at least three minutes. 27. A computer system connected to a cryoprobe, an ultrasound device and a display, the computer system comprising: a. a processor; and b. a programming that causes the processor to: i) receiving 2D image slices of a patient's target tissue from the ultrasound device as it rotates within the patient, the 2D image slices containing representations of an ice ball formed by the cryoprobe, ii) combine the 2D image slices to identify a visible surface of the ice sphere, iii) create a 3D model of the ice sphere based on the visible surface, iv) generate an isotherm for the 3D model of the ice sphere at a selected temperature identified for effective ablation, and v) Creating a user interface on the display that shows the 3D model of the ice sphere and the isotherm. | 2023 | VERAN MEDICAL TECH | Medical and Healthcare | Small-size company | https://veranmedical.com/home/ | USA | The patent focuses on a cryoablation monitoring system that uses a linear EBUS device to guide a cryoprobe and create multiple image slices, then builds a 3D model of the iceball to overcome acoustic shadows. It addresses the fact that iceballs cast acoustic shadows during cryoablation, hiding the freeze zone from ultrasound. The 3D iceball reconstruction reveals the freeze zone behind ultrasound shadows. | הפטנט מתמקד במערכת ניטור קריו-אבלציה המשתמשת במכשיר EBUS ליניארי להנחיית גשוש קריו ויצירת מספר פרוסות תמונה, ואז בונה מודל תלת-ממדי של כדור הקרח כדי להתגבר על צללים אקוסטיים. הוא מתמודד עם העובדה שכדורי קרח מטילים צללים אקוסטיים במהלך קריו-אבלציה, ומסתירים את אזור ההקפאה מאולטרסאונד. השחזור התלת-ממדי של כדור הקרח חושף את אזור ההקפאה מאחורי צללי האולטרסאונד. |  |  | 1 |  |  |  |  | 1 | [6 empty] | 1 | [5 empty] | 1 | [5 empty] | 1 | [5 empty] | VERAN MEDICAL TECH INC | VERAN MEDICAL TECH INC | VERAN MEDICAL TECH INC | VERAN MEDICAL TECH INC | VERAN MEDICAL TECH INC | 2024-05-21 | 2024 | 2026-03-12 | 2026 | 2023-05-22 | 2023 | Application | Utility Invention | active | DE20241102177 | 316778416 | CN121398752A¦US12329568B2¦US20240389970A1¦WO2024243177A1 | DE
413 | CN120899385A | Compound eye laser ablation catheter, laser ablation device and control method thereof | The invention relates to an ablation catheter in an in-vivo laser ablation system, in particular to a compound eye laser ablation catheter, a laser ablation device and a control method of the laser ablation device. In order to overcome the defects that in an existing laser ablation technology, normal saline for cooling cannot be quantitatively controlled, local blood vessels are prone to being damaged in the ablation process, or an imaging device arranged at the head end of a catheter causes reduction of laser transmission efficiency and energy, and the ablation effect is affected due to the poor imaging monitoring effect, normal saline for cooling cannot be quantitatively controlled. The compound eye laser ablation catheter comprises a catheter body, the catheter body comprises a catheter inner cavity, a light transmission optical fiber layer and a catheter wrapping layer, the light transmission optical fiber layer comprises a plurality of light transmission optical fibers and heat conduction layers filled among the light transmission optical fibers, and a guide wire and a temperature sensor are arranged in the catheter inner cavity; an imaging module is arranged on the catheter wrapping layer and comprises at least two compound eye imaging assemblies, and the compound eye imaging assemblies are used for conducting real-time imaging on the ablation part and transmitting the image to the outside. | 1. L A compound eye laser ablation catheter, comprising a catheter body, the catheter body comprising a guidewire (2), a light-transmitting fiber layer and a catheter sheath (6) disposed outside the light-transmitting fiber layer, the light-transmitting fiber layer being used to transmit pulsed laser, and a catheter lumen (1) for injecting physiological saline being disposed along the central axis of the light-transmitting fiber layer, the guidewire (2) being disposed on the central axis of the catheter lumen (1); characterized in that: it further comprises M temperature sensors (3) and an imaging module, wherein, M22; the light-transmitting fiber layer comprising multiple light-transmitting fibers (4) and a heat-conducting layer (5) filling the spaces between the light-transmitting fibers (4), the distal end of the heat-conducting layer (5) being used to dock with a cold surface for heat conduction; The probes of the M temperature sensors (3) are located at the ablation end of the catheter body and are uniformly arranged in the catheter lumen (1) along the circumference of the guidewire (2); the imaging module is arranged on the catheter sheath (6) and close to the ablation end of the catheter body, including at least two compound eye imaging components (7) uniformly arranged along the circumference of the catheter sheath (6), the compound eye imaging components (7) are used to perform real-time imaging of the ablation area and transmit it to the outside. 2.According to claim 1, the compound eye laser ablation catheter is characterized in that: the compound eye imaging component (7) includes a compound eye lens disposed on the side wall of the catheter sheath (6) and a detector disposed corresponding to the compound eye lens; the compound eye lens is a microlens array disposed along an arc surface, and the microlens array includes multiple microlenses arranged in a hexagonal honeycomb structure; each microlens and its six neighboring microlenses are defined as a monitoring unit, and the microlenses of each monitoring unit are respectively configured with different working wavelengths. 3.The compound eye laser ablation catheter according to claim 1 or 2 is characterized in that: the number N of the optical transmission fibers (4) is calculated by the following formula: N = Stop ¥ Where, 2 is the outer radius of the optical fiber layer, j is the inner radius of the conduit, and are the radii of the optical fiber, and B is the fiber filling factor. 4.According to claim 3, the compound eye laser ablation catheter is characterized in that: the ablation end of the optical fiber layer is provided with a protective glass cover (8); The thermally conductive layer (5) is composed of thermally conductive potting compound injected into the gap between adjacent optical fibers (4); the temperature sensor (3) is a miniature thin-film temperature sensor; multiple optical fibers (4) are arranged concentrically around the central axis, with an optical fiber fill factor of 75.7%, or multiple optical fibers (4) are arranged in hexagonal layers from the inside to the outside around the central axis, with an optical fiber fill factor of 90.69%. 5A laser ablation device, characterized in that it comprises a laser source system, an intelligent temperature control system, the compound eye laser ablation catheter as described in claim 1, and an imaging data processing system; the output end of the laser source system is connected to the compound eye laser ablation catheter to provide a laser source; the intelligent temperature control system is connected to both the temperature sensor (3) and the laser source system, and is used to dynamically adjust the output power and repetition frequency of the laser source system based on the temperature data monitored by the temperature sensor (3); the input end of the imaging data processing system is connected to the output end of the imaging module, and is used to receive and process the image formed by the imaging module, and control the output power and repetition frequency of the laser source system based on the processing result, while simultaneously identifying the corresponding... area. 6.According to claim 5, a laser ablation device is characterized in that: the compound eye imaging component (7) includes a compound eye lens and a detector disposed on the side wall of the catheter envelope (6), the compound eye lens being a microlens array disposed along an arc surface, the microlens array including multiple microlenses arranged in a hexagonal honeycomb structure; each microlens and its six neighboring microlenses are defined as a monitoring unit, and the microlenses of each monitoring unit are respectively set with different working wavelengths to achieve large field-of-view multispectral compound eye imaging; the detector is disposed corresponding to the compound eye lens, used to detect and obtain multispectral images, and transmits the multispectral images of the ablation site to the imaging data processing system via wireless transmission; the imaging data processing system performs spectrochemical analysis on the composition of the ablation site based on the multispectral images. 7.A control method for a laser ablation device, characterized by comprising the following steps: Step 1, constructing the laser ablation device as described in claim 5; Step 2, activating the laser source system, the intelligent temperature control system, and the imaging data processing system; Step 3, the laser source system operates according to a preset output power and repetition frequency, while simultaneously injecting physiological saline into the lumen (1) of the catheter; Step 4, adjusting the output power and repetition frequency of the laser source system, as well as the injection speed of physiological saline, according to data from the intelligent temperature control system and the imaging data processing system; wherein the output power and repetition frequency are simultaneously decreased or simultaneously increased. 8.According to claim 7, a control method for a laser ablation device is characterized in that: in step 1, the imaging module of the compound eye laser ablation catheter in the laser ablation device is used to acquire multispectral images; step 4 specifically involves: detecting the temperature of the ablation site using a temperature sensor (3); adjusting the output power and repetition frequency of the laser source system and the injection speed of saline solution according to the temperature of the ablation site; and the imaging data processing system fitting the real-time spectral curve of the current region based on the multispectral image, comparing it with the spectral curve of the region under normal conditions, and reducing the output power and repetition frequency of the laser source system and increasing the injection speed of saline solution when the deviation is within a set threshold; and adjusting the repetition frequency to 0, stopping the output of the light source, and increasing the injection speed of saline solution to the maximum to achieve cooling when the deviation exceeds the set threshold. 9.According to claim 8, the control method of a laser ablation device is characterized in that, in step 4, the intelligent temperature control system adjusts the output power and repetition frequency of the laser source system according to the temperature of the ablation site, specifically: when the temperature of the ablation site and its surrounding tissue is 50-55°C, the output power is reduced to a lower level and the injection speed of physiological saline is increased; when the temperature of the ablation site and its surrounding tissue is greater than 55°C, the repetition frequency is adjusted to 0, the output of the laser source system is stopped, and physiological saline is injected at the maximum speed to cool down. | 2025 | XIAN INST OPTICS & PREC MECHANICS | Research & development | NA | http://www.opt.ac.cn | China | The patent focuses on a 'compound-eye' laser ablation catheter with a cooling-saline control mechanism and head-end imaging that doesn't reduce laser transmission efficiency. It addresses the issues that uncontrolled cooling saline can damage local vessels and tip-imaging optics traditionally reduce laser energy. The improved geometry lets the catheter image the site while maintaining laser energy for safer atherosclerosis treatment. | הפטנט מתמקד בקטטר אבלציית לייזר 'עין מורכבת' עם מנגנון בקרת תמלחת קירור והדמיית קצה ראשי שאינה מפחיתה את יעילות העברת הלייזר. הוא מתמודד עם הבעיות שתמלחת קירור לא מבוקרת יכולה לפגוע בכלי דם מקומיים ושאופטיקת הדמיית קצה מסורתית מפחיתה אנרגיית לייזר. הגיאומטריה המשופרת מאפשרת לקטטר להדמות את האתר תוך שמירה על אנרגיית לייזר לטיפול בטוח יותר באטרוסקלרוזיס. |  | 1 | [11 empty] | 1 | [16 empty] | 1 |  | XIAN INST OPTICS & PREC MECHANICS CAS | XIAN INST OPTICS & PREC MECHANICS CAS | XIAN INST OPTICS & PREC MECHANICS CAS | XIAN INST OPTICS & PREC MECHANICS CAS | XIAN INST OPTICS & PREC MECHANICS CAS | 2025-08-22 | 2025 | 2025-11-07 | 2025 | 2025-08-22 | 2025 | Application | Utility Invention | active | CN202511184569 | 324400040 | [NA] | CN
414 | CN215384559U | Safe and intelligent ablation host combined with nerve monitoring | A safe and intelligent ablation host combined with nerve monitoring comprises an ablation needle interface, a nerve monitoring interface, a main control module, a signal switching module, an ablation module and a nerve monitoring module, and the output end of a radio frequency output circuit and the output end of a stimulation circuit are connected to the ablation needle interface through the signal switching module. The main control module is respectively connected with the signal switching module, the ablation module and the nerve monitoring module, the signal switching module comprises a relay, a switching signal end of the relay is respectively connected with the ablation module and the nerve monitoring module, and a control signal of the relay is connected with the main control module. The ablation electrode of the radiofrequency ablation needle is used as a stimulation electrode for nerve monitoring, a stimulation electrode does not need to be additionally arranged independently, a special ablation needle does not need to be produced independently, and the monitoring effect can be greatly improved through time-division intermittent work of radiofrequency ablation and nerve monitoring. If nerves are about to be damaged in the puncture and ablation process, an alarm can be given in advance, and damage is prevented. | 1 . A safe and intelligent ablation host combined with nerve monitoring, characterized in that it comprises an ablation needle interface, a nerve monitoring interface, a main control module, a signal switching module, an ablation module and a nerve monitoring module, wherein the ablation module includes a radio frequency output circuit, The nerve monitoring module includes a stimulation circuit, the radio frequency output circuit and the output end of the stimulation circuit are connected to the ablation needle connector through a signal switching module, and the main control module is respectively connected to the signal switching module, the ablation module and the nerve monitoring module, so The signal switching module includes a relay, the switching signal terminals of the relay are respectively connected to the ablation module and the nerve monitoring module, and the control signal of the relay is connected to the main control module. 2 . The safe and intelligent ablation host combined with nerve monitoring according to claim 1, wherein the nerve monitoring module comprises a stimulation circuit and a signal receiving and conditioning circuit, and the stimulation circuit is connected to the ablation needle connector through the signal switching module, so that the The ablation needle interface is connected to the ablation needle ablation electrode circuit, and the signal receiving and conditioning circuit is connected to an external nerve monitoring recording electrode through the nerve monitoring receiving interface. 3 . The safe and intelligent ablation host combined with nerve monitoring according to claim 2, wherein the signal receiving and conditioning circuit comprises a signal filtering and amplifying circuit, an AD conversion circuit and an algorithm monitoring circuit connected in sequence, and the algorithm monitoring circuit outputs The main control module is connected, and the signal filtering and amplifying circuit, the AD conversion circuit and the algorithm monitoring circuit are respectively connected to the main control module. 4 . The safe and intelligent ablation host combined with nerve monitoring according to claim 1, wherein the ablation module comprises a radio frequency output circuit and a temperature detection circuit, and the radio frequency output circuit comprises an external DC power input, a DC- A DC voltage regulation conversion circuit and a DC-AC radio frequency output circuit, the DC-AC radio frequency output circuit is connected to the ablation needle interface through a signal switching module, the ablation needle interface is connected to the ablation needle ablation electrode circuit, and the temperature detection circuit does not pass through The signal switch is directly connected to the thermocouple interface of the ablation needle interface. 5 . The safe and intelligent ablation host combined with nerve monitoring according to claim 4, wherein the ablation module further comprises an impedance monitoring circuit, and the output of the impedance monitoring circuit is connected to the main control module. 6 . The safe and intelligent ablation host combined with nerve monitoring according to claim 5, wherein the main control module is further connected with an alarm circuit. | 2020 | ZHEJIANG CURAWAY MEDICAL TECH | Medical and Healthcare | Start Up | http://www.curaway.com/en/ | China | The patent focuses on a utility-model ablation host that combines neural monitoring with a signal switching module connecting RF/stimulation circuits to a needle interface. It addresses the safety burden of needing a separate nerve monitoring device during ablation procedures, where missing nerve injury risk is a real concern. The setup integrates ablation and nerve safety monitoring in one host. | הפטנט מתמקד במארח אבלציה במודל שימושיות המשלב ניטור עצבי עם מודול מיתוג אותות המחבר מעגלי RF/גירוי לממשק מחט. הוא מתמודד עם נטל הבטיחות של צורך במכשיר ניטור עצבים נפרד במהלך פרוצדורות אבלציה, שבהן החמצת סיכון לפגיעה בעצבים היא דאגה אמיתית. ההגדרה משלבת אבלציה וניטור בטיחות עצבים במארח אחד. | 1 |  |  |  | 1 | [5 empty] | 1 | [18 empty] | 1 |  |  | ZHEJIANG CURAWAY MEDICAL TECH CO LTD | ZHEJIANG CURAWAY MEDICAL TECH CO LTD | ZHEJIANG CURAWAY MEDICAL TECH CO LTD | ZHEJIANG CURAWAY MEDICAL TECH CO LTD | ZHEJIANG CURAWAY MEDICAL TECH CO LTD | 2020-11-13 | 2020 | 2022-01-04 | 2022 | 2020-11-13 | 2020 | Grant | Utility Model | active | CN202022622766U | 278977557 | [NA] | CN
415 | CN215899864U | Radio frequency ablation electrode needle with multi-dimensional nerve monitoring function | A radio frequency ablation electrode needle with a multi-dimensional nerve monitoring function comprises a nerve monitoring flexible electrode, a radio frequency thermal coagulation electrode and three knob modules used for adjusting the angle of the nerve monitoring flexible electrode, and an X-axis rotating knob, a Y-axis rotating knob and a Z-axis telescopic knob are arranged on a triangular fixing base. The X-axis rotary knob, the Y-axis rotary knob and the Z-axis telescopic knob are connected with the nerve monitoring flexible electrode through control stay wires respectively, the front end of the triangular fixing seat is fixedly connected with the radio frequency thermal coagulation electrode, and the rear end of the triangular fixing seat is in butt joint with a cable. Compared with the prior art, the electrode tip stretching in multiple dimensions penetrates out of the head end of the radio frequency thermal coagulation electrode, front-back and left-right deflection in the X/Y direction and stretching in the Z direction can be achieved, and the detection position of the flexible electrode tip can be controlled through the X/Y/Z rotary knob; the monitoring electrode disclosed by the utility model can be used for monitoring an ablated nerve focus part in all directions, so that the problem of incomplete nerve damage or excessive nerve damage is avoided to a certain extent. | 1. A radiofrequency ablation electrode needle with multi-dimensional nerve monitoring, characterized in that it includes a nerve monitoring flexible electrode, a radiofrequency thermocoagulation electrode, a knob module for adjusting the angle of the nerve monitoring flexible electrode, and a cable, and the knob module includes a triangular A fixed seat, an X-axis rotation knob, a y-axis rotation knob and a Z-axis telescopic knob, the X-axis rotation knob, the y-axis rotation knob and the Z-axis telescopic knob are respectively arranged on the triangular fixed seat, the X-axis rotation knob, the y-axis rotation knob, The shaft rotation knob and the Z-axis telescopic knob are respectively connected to the nerve monitoring flexible electrodes through control cables, the front end of the triangular fixing base is fixedly connected to the radio frequency thermocoagulation electrode, and the rear end is connected to the cable. 2. The radiofrequency ablation electrode needle with multi-dimensional nerve monitoring according to claim 1, wherein a guide rail device is provided inside the radiofrequency thermocoagulation electrode, and the nerve monitoring flexible electrode is installed on the radiofrequency thermocoagulation electrode through the guide rail device. Internally and isolated from it, the nerve monitoring flexible electrode protrudes from the front end of the radiofrequency thermocoagulation electrode. 3. The radiofrequency ablation electrode needle with multi-dimensional nerve monitoring according to claim 2, wherein the control wire comprises an X-axis wire, a y-axis wire and a Z-axis wire, and the front ends of the flexible electrodes for nerve monitoring are respectively provided with There are X-axis cable fixing points, X-axis corner buckles, y-axis cable fixing points, and y-axis corner buckles. One end of the X-axis cable is wound on the X-axis rotating knob, and the other end is fixed on the X-axis cable through the X-axis corner buckle. On the fixed point, one end of the y-axis cable is wound on the y-axis rotating knob, the other end is fixed on the y-axis cable fixing point through the y-axis corner buckle, and one end of the Z-axis cable is wound on the Z-axis telescopic knob, and the other end is fixed on the y-axis cable fixing point. One end is connected to the side of the rear end of the flexible electrode for nerve monitoring, the radio frequency thermocoagulation electrode is provided with a wire tube inside, and the X-axis pulling wire, the y-axis pulling wire and the Z-axis pulling wire are arranged through the wire tube. 4The radiofrequency ablation electrode needle with multi-dimensional nerve monitoring according to claim 3, wherein a thermocouple is arranged at the inner front end of the radiofrequency thermocoagulation electrode. 5. The radiofrequency ablation electrode needle with multi-dimensional nerve monitoring according to claim 3, wherein the bottom surface of the triangular holder is an equilateral triangle, and the triangular holder comprises an electrode communication socket, a triangular holder body and a A vertebral body with a 6% Luer taper, the X-axis rotation knob, the y-axis rotation knob and the Z-axis telescopic knob are respectively located at the top three corners of the bottom surface of the triangular fixed seat, and the electrode communication socket is located on the triangular fixed seat The center position of the triangular fixation seat, there is a cavity inside the triangular fixation seat, the top of the triangular fixation seat is provided with a vertebral body with a 6% Luer taper, and the radiofrequency thermocoagulation electrode is inserted into the vertebral body with a 6% Luer taper, connected to Assemble in a sealed box with glue. 6. The radio frequency ablation electrode needle with multi-dimensional nerve monitoring according to claim 5, wherein the two ends of the cable are respectively provided with an electrode communication plug and a radio frequency host connection plug, and the electrode communication plug is used to connect the triangle The electrode communication socket of the fixed seat, the radio frequency host connection plug is used to connect the radio frequency host. 7The radiofrequency ablation electrode needle with multi-dimensional nerve monitoring according to claim 6, wherein the electrode communication plug is in the shape of a meniscus, and the electrode communication socket is in the shape of a meniscus. | 2020 | ZHEJIANG CURAWAY MEDICAL TECH | Medical and Healthcare | Start Up | http://www.curaway.com/en/ | China | The patent focuses on a utility-model RF ablation electrode needle with multi-dimensional nerve monitoring and X/Y rotational and Z-axis telescopic knobs to fine-tune the nerve electrode angle and depth. It addresses the difficulty of precisely positioning a nerve monitoring electrode relative to the RF ablation tip during pain treatment. Operators can fine-tune the nerve electrode position relative to the RF tip. | הפטנט מתמקד במחט אלקטרודת אבלציית RF במודל שימושיות עם ניטור עצבים רב-ממדי וכפתורי סיבוב X/Y וטלסקופ Z לכוונון עדין של זווית ועומק אלקטרודת העצב. הוא מתמודד עם הקושי במיקום מדויק של אלקטרודת ניטור עצבים ביחס לקצה אבלציית ה-RF במהלך טיפול בכאב. מפעילים יכולים לכוונן בעדינות את מיקום אלקטרודת העצב ביחס לקצה ה-RF. | 1 |  |  |  | 1 | [5 empty] | 1 | [18 empty] | 1 |  |  | ZHEJIANG CURAWAY MEDICAL TECH CO LTD | ZHEJIANG CURAWAY MEDICAL TECH CO LTD | ZHEJIANG CURAWAY MEDICAL TECH CO LTD | ZHEJIANG CURAWAY MEDICAL TECH CO LTD | ZHEJIANG CURAWAY MEDICAL TECH CO LTD | 2020-12-29 | 2020 | 2022-02-25 | 2022 | 2020-12-29 | 2020 | Grant | Utility Model | active | CN202023259802U | 292546677 | [NA] | CN
416 | CN117731246A | Ablation point position and temperature field monitoring device based on thin film optical fiber sensing array | The invention discloses an ablation point position and temperature field monitoring device based on a thin-film optical fiber sensing array, and the device comprises the thin-film optical fiber sensing array which is used for being attached to the surface of a human body for monitoring a body surface temperature field; the signal demodulation module is used for demodulating the optical signals in the thin film optical fiber sensing array into temperature data; the upper computer is used for obtaining ablation point positions and temperature field distribution by utilizing an ablation point position and temperature field prediction algorithm according to the temperature data, and the ablation point position and temperature field prediction algorithm is constructed through the following processes: generating a training data set according to array arrangement information sensed by the thin film optical fiber and an ablation point temperature field prediction model; training a neural network by using the training data set; and acquiring temperature field information data of the thin-film optical fiber sensing array obtained by the calibration device, so that the trained neural network is corrected and calibrated to obtain an ablation point position and a temperature field prediction algorithm. | 1· An ablation point position and temperature field monitoring device based on a thin film optical fiber sensing array, which is characterized by including: a thin film optical fiber sensing array, the thin film optical fiber sensing array is used to attach to the surface of the human body to monitor the body surface temperature field Monitoring; a signal demodulation module, the signal demodulation module is used to demodulate the optical signal in the thin film optical fiber sensing array into temperature data; a host computer, the host computer is used to use ablation according to the temperature data The point position and temperature field prediction algorithm obtains the ablation point position and temperature field distribution. The ablation point position and temperature field prediction algorithm is constructed through the following process: According to the array arrangement information of the thin film optical fiber sensing and the ablation point temperature field prediction model Generate a training data set, and use the training data set to train the neural network; obtain the temperature field information data of the thin film optical fiber sensing array obtained through the calibration device, thereby correcting and calibrating the trained neural network to obtain the ablation point Position and temperature field prediction algorithm. 2. The device according to claim 1, wherein the thin film optical fiber sensing array includes a flexible film and an optical fiber sensing array, and the optical fiber sensing array is based on the topological design of an equally spaced parallel grid and is shaped by a mold. The optical fiber arrangement is adhered to the heat shrink film and fixed on the surface of the flexible film. 3. The device according to claim 2, wherein the optical fiber sensing array includes a temperature sensing segment optical fiber and a connecting segment optical fiber, and the temperature sensing segment optical fiber is arranged on the flexible film in an array form, It is used for temperature field monitoring, and the connecting section optical fiber is used to connect adjacent temperature sensing section optical fibers. 4. The device according to claim 1, characterized in that the signal demodulation module is connected to the thin film optical fiber sensing array through an optical fiber jumper. 5. The device according to claim 1, characterized in that the signal demodulation module is connected to the host computer through a communication optical cable. 6. The device according to claim 1, wherein the neural network is a long short-term memory network, a bidirectional long short-term memory network or a gated recurrent unit neural network. 7. The device according to claim 1, characterized in that generating a training data set based on the array arrangement information of the thin film optical fiber sensing and the ablation point temperature field prediction model includes: based on the temperature in the thin film optical fiber sensing array. Construct a temperature measuring point information matrix based on the position of the measuring point; according to the thin film optical fiber sensing array temperature measuring point position information matrix and the preset ablation point position and temperature information, a thin film optical fiber sensor is generated through the classic ablation point temperature field prediction model The array temperature field distribution data and the major diameter and short diameter of the ablation temperature affected area are then used to obtain a training data set for neural network training; wherein, the temperature field distribution data of the thin film optical fiber sensing array is the training input in the training data set. , the ablation point position, temperature and the major axis and minor axis of the ablation temperature affected area are the training outputs in the training data set. 8. The device according to claim 1, characterized in that the ablation point temperature field prediction model is: - except not^πk^\d~ +x wherein, % is the ablation point ablation temperature, k is the human body at the ablation site Thermal conductivity of tissue, d is the vertical distance from the ablation point to a position near the ablation point, quality. 9. The device according to claim 1, characterized in that, the thin film optical fiber sensing array obtained by the calibration device is obtained. A series of temperature field information data is obtained to correct and calibrate the trained neural network, including: obtaining the temperature field information data of the thin film optical fiber sensing array obtained through the calibration device, wherein the temperature field information data of the thin film optical fiber sensing array is obtained through the calibration device. The position movement and temperature adjustment of the point-like heat source inside the calibration device are obtained by simulating different ablation positions and temperature conditions; the data are used to correct and calibrate the trained neural network. 10. The device according to claim 1, wherein the calibration device includes simulated human tissue, a fiber grating array and a point heat source, and the point heat source simulates different ablation positions and temperatures by moving the position and adjusting the temperature, so The fiber grating array is arranged on the simulated human tissue in vertical and horizontal directions for temperature calibration. | 2023 | ZHEJIANG LAB | Research & development | NA | https://www.zhejianglab.com | China | The patent focuses on a thin-film fiber-optic sensing array attached to the body surface that monitors body-surface temperature field and ablation-point position, plus a signal demodulator and host computer. It addresses the safety challenge in laser/HIFU/RF thermal ablation where high-temperature exposure to surrounding tissue isn't well monitored externally. The array reads temperature and position non-invasively from the body surface. | הפטנט מתמקד במערך חישת סיב אופטי של סרט דק המוצמד למשטח הגוף המנטר שדה טמפרטורת משטח הגוף ומיקום נקודת אבלציה, בתוספת מדמודלת אותות ומחשב מארח. הוא מתמודד עם אתגר הבטיחות באבלציה תרמית בלייזר/HIFU/RF שבה חשיפת רקמה סובבת לטמפרטורה גבוהה אינה מנוטרת היטב חיצונית. המערך קורא טמפרטורה ומיקום באופן לא פולשני ממשטח הגוף. |  | 1 | [20 empty] | 1 | [7 empty] | 1 |  | ZHEJIANG LAB | ZHEJIANG LAB | ZHEJIANG LAB | ZHEJIANG LAB | ZHEJIANG LAB | 2023-12-18 | 2023 | 2024-03-22 | 2024 | 2023-12-18 | 2023 | Application | Utility Invention | active | CN202311738770 | 295086898 | [NA] | CN
417 | WO2022067146A3 | SYSTEMS, DEVICES AND METHODS FOR TREATING LUNG TUMORS WITH A ROBOTICALLY DELIVERED CATHETER | A system for treatment of a target region of lung tissue including: a flow regulator configured to be interposed between a conductive fluid source and a conductive fluid outlet positionable at or in proximity of the target region of lung tissue, the flow regulator being further configured for controlling a flow rate or a bolus quantity of conductive fluid coming from the fluid source and delivered to the conductive fluid outlet; and a controller communicatively connectable with said flow regulator and with at least one sensor, with the at least one sensor being configured for detecting values taken by at least one control parameter representative of a physical property, wherein the physical property is one of temperature (T), pressure (p), electric impedance (Z), or electric conductivity (C) of material present at or in proximity of the target region of lung tissue. | 1. CLAIMS: The invention is: 1. A system for treatment of a target region of lung tissue, the system comprising: at least one a flow regulator configured to be interposed between a conductive fluid source and a conductive fluid outlet positionable at or in proximity of the target region of lung tissue, the flow regulator being further configured for controlling a flow rate or a bolus quantity of conductive fluid coming from the fluid source and delivered to the conductive fluid outlet; a controller communicatively connectable with said flow regulator and with at least one sensor, with the at least one sensor being configured for detecting values taken by at least one control parameter representative of a physical property, wherein the physical property is one of temperature (T), pressure (p), electric impedance (Z), or electric conductivity (C) of material present at or in proximity of the target region of lung tissue; wherein the controller is configured for: receiving from said sensor signals representative of detected values of the control parameter; controlling the flow regulator based on one or more detected values of the control parameter, wherein controlling the flow regulator comprises executing a control cycle including: o controlling the flow regulator in a high delivery mode, wherein in the high delivery mode: ■ the flow rate of conductive fluid delivered to the conductive fluid outlet is equal or above a set high flow rate, or■ the bolus quantity of conductive fluid delivered to the conductive fluid outlet is equal above a set high bolus quantity, o controlling the flow regulator in a low delivery mode, wherein in the low delivery mode: ■ the flow rate of conductive fluid delivered to the conductive fluid outlet is equal or below a set low flow rate smaller than the set high flow rate, or ■ the bolus quantity of conductive fluid delivered to the conductive fluid outlet is equal or below a set low bolus quantity smaller than the set high bolus quantity. 2. A system according to claim 1, wherein in the low delivery mode: ■ the flow rate of conductive fluid delivered to the conductive fluid outlet is equal or below a set low flow rate smaller than 50% of the set high flow rate, or■ the bolus quantity of conductive fluid delivered to the conductive fluid outlet is equal or below a set low bolus quantity smaller than 50% of the set high bolus quantity. 3. A system according to claim 1 or 2, wherein in the low delivery mode the set low flow rate is between 0 and 5 ml/min or wherein the set low bolus quantity is between 0 and 10 ml. 4. A system according to claim 1 or 2 or 3, wherein in the high delivery mode the set high flow rate is between 2 and 16 ml/min or the set high bolus quantity is between 0.3 and 60 ml. 5. A system according to any one of the preceding claims, wherein controlling the flow regulator comprises repeatedly executing said control cycle. 6. A system according to any one of the preceding claims, comprising at least one ablation element positionable at the target region of the lung tissue and connectable to an ablation source. 7. A system according to claim 6, comprising at least one flexible shaft configured to advance through an airway passage of a lung and having an active portion positionable at the target region of the lung tissue and including the at least one ablation element. 8. A system according to any one of the preceding claims, comprising the at least one sensor, the sensor being configured to be positionable at the target region of the lung tissue. 9. A system according to claim 7 in combination with any one of claims 1-6 and 8, wherein the at least one sensor is carried by the active portion of said flexible shaft. 10. A system according to claim 7 in combination with any one of claims 1-6 and 8, wherein the at least one sensor is configured to be positioned in correspondence of a volume surrounding the active portion of said flexible shaft. 11. A system according to claim 9 or 10, wherein the at least one sensor is configured for sensing values taken by the at least one control parameter, and wherein the physical property is one of temperature, pressure, electric impedance, or electric conductivity of material present in a volume surrounding the active portion. 12. A system according to any one of the preceding claims comprising the conductive fluid outlet which is configured to be placed in fluid communication with the conductive fluid source. 13. A system according to claim 7 and 12, wherein the conductive fluid outlet is carried by the flexible shaft active portion. 14. A system according to claim 7 and 12, wherein the conductive fluid outlet is configured to be positioned in correspondence of said volume surrounding the active portion. 15. A system according to any one of the preceding claims 6-14, wherein the controller is connectable with said ablation source and configured for controlling the ablation energy source to deliver ablation energy to the at least one ablation element. 16. A system according to claim 15, further wherein the controller is configured for executing said steps of: receiving from said sensor signals representative of sensed values of the control parameter, and105 controlling the flow regulator based on one or more sensed values of the control parameter and executing, optionally repeatedly executing, said control cycle, while the controller commands the ablation energy source to deliver ablation energy to the at least one ablation element. 17. A system according to any one of the preceding claims 7-16, comprising an electrical connector carried by the flexible shaft and adapted to electrically connect the at least one ablation element to the ablation source. 18. A system according to any one of the preceding claims, wherein the control cycle includes: verifying if one or more sensed values of the control parameter fall below a set low threshold (T Low), and wherein said controlling the flow regulator to low delivery mode is executed if the one or more sensed values of the control parameter fall below the set low threshold (T Low). 19. A system according to any one of the preceding claims, wherein the control cycle includes: verifying if one or more sensed values of the control parameter exceed a set high threshold (T High, Z High), and wherein said controlling the flow regulator to high delivery mode is executed if the one or more sensed values of the control parameter exceed the set high threshold (T High, Z High). 20. A system according to any one of the preceding claims, wherein the control cycle includes: periodically verifying if one or more sensed values of the control parameter fall below a set low threshold (T Low),106 switching the flow regulator from high delivery mode to low delivery mode when the one or more sensed values of the control parameter fall below the set low threshold (T Low); optionally, wherein said step of periodically verifying is executed at least 10 times per second. 21. A system according to any one of the preceding claims, wherein the control cycle includes: periodically verifying if one or more sensed values of the control parameter exceed a set high threshold (T High, Z High), switching the flow regulator from low delivery mode to high delivery mode when the one or more sensed values of the control parameter exceed the set high threshold (T High, Z High); optionally, wherein said step of periodically verifying is executed at least 10 times per second. 22. A system according to any one of the preceding claims wherein the controller is configured for repeating the control cycle a plurality of times during a same treatment session. 23. A system according to claim 22, wherein the controller is configured to control the flow regulator in high delivery mode or in low delivery mode for a respective time interval, and wherein a duration of said respective time intervals is either predetermined or determined by detection of a triggering event. 24. A system according to claim 23, wherein the controller is configured to determine duration of said time intervals by detection of a triggering event, wherein detection of the triggering event comprises one or more of: detection that one or more values of the sensed parameter exceeds a set very high threshold (T Overheat),107 detection that one or more values of the sensed parameter exceeds said set high threshold (T High, Z High), detection that one or more values of the sensed parameter falls below a set low threshold (T Low). 25. A system according to claim 22 or 23 or 24, wherein the controller is configured to execute the same treatment session which includes a plurality of time intervals where the flow regulator is adjusted to low delivery mode intercalated by time intervals where the flow regulator is adjusted to high delivery mode, thereby reducing the overall amount of conductive fluid delivered over said treatment session while maintaining under control the detected values of the parameter. 26. A system according to any one of the preceding claims, wherein the step of controlling the flow regulator to low delivery mode comprises: adjusting the flow regulator to maintain the flow rate of conductive fluid to the conductive fluid outlet equal or below said set low flow rate during a low delivery time interval (Flow Low Time), in particular comprised between 1 to 10 seconds; or adjusting the flow regulator to deliver to the conductive fluid outlet the bolus quantity of conductive fluid equal or below said set low bolus quantity within a low delivery time interval (Flow Low Time), in particular comprised between 1 to 10 seconds. 27. A system according to claim 26, wherein the cycle comprises a sub-routine optionally executed after expiration of said low delivery time interval, said sub-routine including: a further step of verifying if one or more values of the sensed parameter falls below or above the set low threshold (T Low), in case one or more values of the parameter sensed in the further step of verifying falls below the set low threshold (T Low), assigning a decreased value to the set low flow rate or to the set low bolus quantity, and108 repeating controlling the flow regulator to low delivery mode using the decreased value of the set low flow rate or the decreased value of set low bolus quantity. 28. A system according to any one of the preceding claims, wherein the step of controlling the flow regulator to high delivery mode comprises: adjusting the flow regulator to maintain the flow rate of conductive fluid to the conductive fluid outlet equal or above said set high flow rate during a high delivery time interval (Flow High Time), in particular comprised between 1 to 30 seconds; or adjusting the flow regulator to deliver to the conductive fluid outlet the bolus quantity of conductive fluid equal or above said set high bolus quantity within a high delivery time interval (Flow High Time), in particular comprised between 1 to 30 seconds. 29. A system according to claim 28, wherein the cycle comprises a sub-routine optionally executed after expiration of said high delivery time interval, said sub-routine including: a further step of verifying if one or more values of the sensed parameter falls below or above the set low threshold (T Low), in case one or more values of the parameter sensed in the further step of verifying remains above the set low threshold (T Low), assigning an increased value to the set high flow rate or to the set high bolus quantity, and repeating controlling the flow regulator to high delivery mode using the increased value of the set high flow rate or the increased value of set high bolus quantity. 30. A system according to claim 29, wherein the cycle provides for repeating the subroutine of claim 29 until when the further step of verifying if one or more values of the sensed falls below the set low-threshold (T Low) is positively passed. 31. A system according to claim 16 and 30, wherein the controller is configured to interrupt or reduce delivery of ablation energy to the at least one ablation element if it has109 determined that, after a predetermined number of repetitions of the subroutine of claim 29, the step of verifying if one or more values of the sensed fall below the set low threshold (T Low) is not positively passed. 32. A system according to any one of the preceding claims wherein the cycle comprises: determining occurrence of a safety relevant event if the one or more parameter values are above a set over-high threshold (T Over High; Z Over High), which is greater than said high threshold (T High; Z High); if a safety relevant condition is determined, then: o temporarily adjust down power supplied to the ablation energy source and/or o control the flow regulator to a very high delivery mode, wherein in the very high delivery mode the flow rate of conductive fluid delivered to the conductive fluid outlet is equal or above a set very high flow rate greater than the set high flow rate, or the bolus quantity of conductive fluid delivered to the conductive fluid outlet is equal or above a very set high bolus quantity greater than the high bolus quantity. 33. A system according to any one of the preceding claims 6 to 32, wherein the controller is configured for maintaining power supplied by the ablation energy source in a range comprised between 20 and 200 W over a major portion of the treatment session, optionally over the entire treatment session. 34. A system according to any one of preceding claims 6-33, wherein the controller is configured to increase power supplied by the ablation energy source from an initial value to a regimen value, during an initial portion of the treatment session optionally lasting between 10% and 30% of the entire treatment session. 11035. A system according claim 34, wherein the controller is configured to maintain power supplied by the ablation energy source at the regimen value during a major portion of the treatment session following said initial portion of the treatment session. 36. A system according to claim 35, wherein the initial value is comprised between 20 W and 80 W and wherein the regimen value is comprised between 40 W and 200 W, further wherein the initial value smaller than 80% of the regimen value, optionally smaller than 50% of the regimen value. 37. A system according to any one of the preceding claims 22-36 wherein the treatment session has a total treatment duration comprised between 30 seconds and 30 minutes. 38. A system according to any one of preceding claims 6-37, wherein the controller is configured to automatically stop delivery of power from the ablation energy source and automatically command the flow regulator to stop delivery of conductive fluid when the treatment duration has expired. 39. A system according to any one of the preceding claims 22-38, wherein the controller is configured to control the flow regulator to impose that: a maximum volume of conductive fluid delivered during the treatment session is comprised between 0.3 ml and 60 ml, and/or an average flow rate of conductive fluid maintained during the treatment session is of 0.1 to 15 ml/min, in particular wherein the controller is configured to automatically stop delivery of power from the ablation energy source and/or automatically command the flow regulator to stop delivery of conductive fluid when said maximum of conductive fluid delivered has been reached. 40. The system of any one of the preceding claims in combination with one of claims 18- 21, wherein the set high threshold (T High) is greater than the set low threshold (T Low). 11141. The system of claim 40 in combination with claim 24, wherein the set very high threshold (T Overheat) is greater than the set high threshold (T High). 42. A system according to any one of the preceding claims, wherein the physical property is the temperature of material present at the target region, in particular when this claim also depends upon claim 11 the physical property is the temperature of material present in the volume surrounding the active portion. 43. The system of claims 40 and 42, wherein said set low threshold (T Low) is 60 to 95°C. 44. The system of claim 40 in combination with one of claims 42 or 43, wherein said set high threshold (T High) is from above 75°C to 105°C. 45. The system of claim 41 in combination with one of claims 42 or 43 or 44, wherein said set very high threshold (T Overheat) is between 85 to 115°C. 46. The system of any one of the preceding claims, wherein the ablation energy source is a radiofrequency generator and wherein the controller is configured to control the radiofrequency generator to deliver RF, having a power in a range of 1 to 200 W, in particular comprised between 20 and 200 W, for a duration of 30 seconds to 30 minutes. 47. The system of any one of the preceding claims, wherein the flow regulator comprises: - a pump, optionally a syringe pump or a peristaltic pump or an infusion pump, or- a valve. 48. The system of any one of the preceding claims 7-47, comprising: - a conductive fluid source configured to deliver an hypertonic saline solution; - a fluid port connectable to the conductive fluid source and in fluid communication with the conductive fluid outlet,112 optionally, wherein the hypertonic saline solution includes a reverse phase transition polymer and water, which may transition to higher viscosity when transitioned from below body temperature to body temperature. 49. The system of claim 48, wherein the hypertonic saline solution comprises one or more physiologically acceptable solutes and has a theoretical Osmolarity between 0.8 and 15 Osm/L, calculated according to the formulaOsmolarity = (molarity X ri) Each solute in which n is the number of particles that dissociate from each solute molecule; optionally wherein the hypertonic saline solution is a solution according to any one of claims from 123 to 146. 50. The system of claim 48 or 49, wherein the hypertonic saline solution comprises sodium chloride (NaCl) at a concentration of between 3% to 30% (w/v). 51. The system of any one of the preceding claims 7-50, wherein the flexible shaft is the flexible shaft of an ablation catheter. 52. The system of claim 51, wherein the ablation catheter has a/the fluid port is at a proximal end of the flexible shaft and is in fluid communication with the conductive fluid outlet which is located at the active portion of the flexible shaft. 53. The system of any one of the preceding claims 7-52, wherein the active portion is a distal end portion of the flexible shaft. 54. The system of any one of the preceding claims 7-53 comprising at least one space occluder operative at or proximate to the flexible shaft active portion, in particular at or proximate to the flexible shaft distal end portion. 55. The system of claim 54, wherein the space occluder is one of a deployable balloon, a deployable valve, a deployable stent, or a tapered shaft section. 56. The system of claim 54 or 55, wherein the space occluder comprises a deployable occlusion balloon having a first cross section width of 1 to 30 mm, a length in a range of 5 to 30 mm, and wherein the occlusion balloon is configured to expand to occlude a portion of the airway. 57. The system of claim 56, wherein the first cross section width is at a proximal region of the deployable occlusion balloon, a second cross section width in a range of 1 to 30 mm is at a distal region of the balloon, and a cross section width between the first and second cross section width is less than both the first and second cross section width. 58. The system of claim 56, wherein the first cross section width is at a proximal region of the deployable occlusion balloon, and a second cross section width in a range of 1 to 20 mm and less than the first cross section width is at a distal region of the balloon. 59. The system of any one of the preceding claims 7-58, comprising a tubular sheath or bronchoscope receiving said flexible shaft, wherein at least the active portion, in particular the distal end portion, of the flexible shaft is configured to emerge from the tubular sheath or bronchoscope. 60. The system of claim 59, in combination with any one of claims 54 to 58, wherein the space occluder is carried by the tubular sheath or bronchoscope. 61. The system of any one of claims 54 to 58, wherein the at least one space occluder is directly carried by the flexible shaft or by another shaft. 62. The system of any one of claims 54 to 61, further comprising an inflating lumen extending through the flexible shaft and having a proximal end connectable to a source of a fluid, optionally a source of liquid or a source of gas, and a distal end in fluid communication with inside of the balloon. 63. The system of any one of the preceding claims 7 to 62, wherein the flexible shaft comprises depth markers at least on 5 cm of the proximal region and 5 cm of the distal region. 64. The system of any one of the preceding claims 7 to 62, comprising at least one suction opening at the flexible shaft distal end portion configured to be placed in fluid communication with a vacuum source to aspirate air from a lung volume surrounding the distal end portion of the shaft. 65. The system of claim 64, in combination with any one of claims 54 to 63, wherein the at least one suction opening is positioned distal with respect to the space occluder. 66. The system of claim 64, in combination with any one of claims 54 to 63, comprising an additional space occluder operative at or proximate to the shaft distal end portion, in particular wherein the additional space occluder is one of a deployable balloon, a deployable valve, a deployable stent, and wherein the at least one suction opening is positioned between the space occluder and the additional space occluder. 67. The system of claim 65 or 66, wherein the at least one conductive fluid outlet is positioned distal with respect to the space occluder or between the space occluder and the additional space occluder. 68. The system of any one of claims 54 to 67, wherein the at least one sensor is positioned distal to the space occluder or between the space occluder and the additional space occluder. 69. The system of claim 67, wherein the at least one sensor is carried by the flexible shaft distal end portion and wherein the physical property is of one of temperature, pressure, electric impedance, or electric conductivity of material surrounding the distal end portion of the flexible shaft. 11570. The system of claim 68 or 69, wherein the at least one sensor comprises a first sensor positioned proximal to the ablation element and a second sensor positioned distal to the ablation element. 71. The system of any one of the preceding claims 7-70, wherein the at least one ablation element comprises an ablation element having a rounded distal tip positioned at a distal tip of the flexible shaft. 72. The system of claim 64, in combination with any of claims 54 to 63 and 65 to 71, comprising: a common lumen extending through the flexible shaft and having a proximal end, selectively connectable to at least one of the source of the conductive liquid and the vacuum source, and a distal end, forming a common opening defining said at least one outlet and said at least one suction opening; or a dedicated irrigation lumen and a dedicated air suction lumen, with the irrigation lumen connected to the at least one outlet and extending through the catheter flexible shaft, the irrigation lumen having an inlet port configured to be connected to the source of conductive fluid, and with the air suction lumen connected to the at least one air suction opening and extending through the catheter flexible shaft, the air suction lumen having a suction port configured to be connected to the source of vacuum. 73. The system of any one of the preceding claims 7-72, wherein the flexible shaft has an outer diameter less than or equal to 2 mm. 74. The system of any one of the preceding claims 7-73, wherein at least a portion of the flexible shaft is capable of turning such that a bend in the shaft has a radius of curvature of at least 7 mm. 75. The system of any one of the preceding claims 7-74, wherein the flexible shaft has a length of at least 50 cm. 11676. The system of any one of the preceding claims 7-75, wherein the elongated shaft has a guidewire lumen in the distal end portion configured to receive a guidewire. 77. The system of claims 72 and 76, wherein the suction lumen and guidewire lumen are formed by a common lumen. 78. The system of any one of the preceding claims 6-77, wherein the ablation element comprises at least one electrode characterized by one or more of the following features: total surface area not greater than 120 mm2; diameter in a range of 0.5 to 2 mm; length in a range of 3 to 20 mm. 79. The system of any one of the preceding claims 6-78, wherein the at least one ablation element includes at least two electrodes, and wherein a separation between electrodes is between five to fifteen mm. 80. The system of any one of the preceding claims including an interface component connectable with said at least one sensor and at least communicatively connectable with the controller to transfer to the controller the detected values of said at least one control parameter detected by the sensor. 81. The system of any one of the preceding claims, wherein the controller is configured for: processing said sensed values, and based on one or more of said sensed values, generating at least one output signal which comprises one or more of: o a user identifiable output, optionally the user identifiable output comprising an audible signal, a visual signal or a vibratory signal signaling to the user to deploy at least one space occluder operative at or proximate to the shaft distal end portion,117 o a status output, indicative of the degree of air volume reduction of a lung portion located at a/the catheter distal end portion, o an output command automatically deploying at least one space occluder operative at or proximate to a/the shaft distal end portion, o a temperature output providing an indication of the temperature of material surrounding a/the distal end portion of a/the flexible shaft, o an electric property output providing an indication of the impedance or conductivity of material surrounding a/the distal end portion of the shaft, o a pressure output providing an indication of the pressure of material surrounding a/the distal end portion of a/the flexible shaft. 82. The system of any one of the preceding claims, wherein the controller is configured to: receive signals from the at least one sensor, said sensor being a temperature sensor configured to: monitor temperature at said target region, and control the conductivity or the composition of the conductive fluid delivered through said at least one outlet based on the monitored temperature to maintain the temperature values detected by the temperature sensor within a determined temperature range or above a certain temperature threshold. 83. The system of any one of the preceding claims, wherein controller configured to: receive signals from the at least one sensor, said sensor being a temperature sensor, in particular when this claim depends upon claim 7 said sensor being configured for detecting values of temperature of material surrounding the distal end portion of the flexible shaft, monitor temperature at the target region, adjust the ablation energy power output by the energy source to maintain the temperature values detected by the temperature sensor within a determined temperature range or above a certain temperature threshold. 11884. The system of claim 82 or 83, wherein the determined temperature range is between 60 and 115 °C and the certain temperature threshold is at least 80 °C. 85. The system of any one of the preceding claims 1-84, further comprising a navigation sensor, such as a three-dimensional navigation sensor, or a shape sensor, such as a Fiber Bragg Grating sensor, on at least the distal end region, in particular wherein the navigation sensor is one or more of an electromagnetic sensor, a 3D electromagnetic sensor, shape sensor, FBG sensor, a 3D ultrasound sensor, and an impedance tracking for 3D navigation. 86. The system of any one of the preceding claims 7-85, further comprising a perforation element at a distal tip of the flexible shaft configured to advance through a tumor, wherein the perforation element is selected from a list comprising a needle, a deployable needle, and an RF perforation electrode. 87. The system of any one of the preceding claims 6-87 in combination with claim 54, wherein a distance between the space occluder and the ablation element is in a range of 1 mm to 40 mm. 88. An ablation catheter comprising: a flexible shaft configured to advance through airway passages of a lung; at least one ablation element, positioned at a distal end portion of the flexible shaft and electrically connectable to an ablation energy source; a fluid port connectable to a conductive fluid source; at least one outlet for conductive fluid, the outlet being located at said distal end portion and being in fluid communication with the fluid port. 89. The catheter of claim 88 further comprising at least one space occluder operative at or proximate to the shaft distal end portion, in particular wherein the space occluder is one of a tapered shaft section, a deployable balloon, a deployable valve, or a deployable stent. 11990. The catheter of claim 89, wherein the occluder comprises a deployable occlusion balloon having a first cross section width of 1 to 30 mm, a length in a range of 5 to 30 mm, and wherein the occlusion balloon is configured to expand to occlude a portion of the airway. 91. The catheter of claim 90, wherein the first cross section width is at a proximal region of the deployable occlusion balloon, a second cross section width in a range of 1 to 30 mm is at a distal region of the balloon, and a cross section width between the first and second cross section width is less than both the first and second cross section width. 92. The catheter of claim 90, wherein the first cross section width is at a proximal region of the deployable occlusion balloon, and a second cross section width in a range of 1 to 20 mm and less than the first cross section width is at a distal region of the balloon. 93. The catheter of claim 88 to 92, comprising a tubular sheath or bronchoscope receiving said shaft, wherein at least the distal end portion of the flexible shaft is configured to emerge from the tubular sheath or bronchoscope. 94. The catheter of claim 93 in combination with claim 90, wherein the space occluder is carried by the tubular sheath or bronchoscope or wherein the at least one space occluder is directly carried by the shaft or by a different shaft. 95. The catheter of claim 88 to 94 in combination with claim 90, further comprising an inflating lumen extending through the flexible shaft and having a proximal end connectable to a source of a fluid, optionally a source of liquid or a source of gas, and a distal end in fluid communication with inside of the balloon. 96. The catheter of any of claims 88 to 95, wherein the flexible shaft comprises depth markers at least on 5 cm of the proximal region and 5 cm of the distal region. 12097. The catheter of any of claims 88 to 96, comprising at least one suction opening at the shaft distal end portion configured to be placed in fluid communication with a vacuum source to aspirate air from a lung volume surrounding the distal end portion of the shaft. 98. The catheter of claim 97 in combination with claim 90, wherein the at least one suction opening is positioned distal with respect to the space occluder. 99. The catheter of claim 98 comprising an additional sp<TRUNCATED> | 2020 | ZIDAN MEDICAL | Medical and Healthcare | Start Up | https://www.zidanmedical.com/ | USA | The patent focuses on a robotically delivered catheter system for treating lung tumors that uses a flow regulator between a conductive-fluid source and an outlet at the target. It addresses the difficulty of safely delivering conductive fluid for lung tumor ablation when flow rate or bolus size needs precise control. Controlled fluid delivery (rate or bolus) supports safer ablation. | הפטנט מתמקד במערכת קטטר המסופק רובוטית לטיפול בגידולי ריאה המשתמשת בוסת זרימה בין מקור נוזל מוליך לפלט במטרה. הוא מתמודד עם הקושי במסירה בטוחה של נוזל מוליך לאבלציית גידול ריאתי כאשר קצב הזרימה או גודל הבולוס זקוקים לבקרה מדויקת. מסירת נוזל מבוקרת (קצב או בולוס) תומכת באבלציה בטוחה יותר. |  |  |  |  | 1 | [5 empty] | 1 | [15 empty] | 1 | [5 empty] | ZIDAN MEDICAL INC | ZIDAN MEDICAL INC | ZIDAN MEDICAL INC | ZIDAN MEDICAL INC | ZIDAN MEDICAL INC | 2021-09-27 | 2021 | 2022-05-12 | 2022 | 2020-09-28 | 2020 | Application | Utility Invention | inactive | WO2021US52121 | 291867582 | AU2021350113A1¦CA3192769A1¦CN116744872A¦EP4216857A2¦JP2023544302A¦KR20230097033A¦US20230355300A1¦WO2022067146A2 | WO
418 | CN111429432B | Thermal ablation area monitoring method and system based on radio frequency processing and fuzzy clustering | The invention provides a thermal ablation area monitoring method and system based on radio frequency processing and fuzzy clustering. The thermal ablation area monitoring method based on radio frequency processing and fuzzy clustering comprises the following steps: S1, acquiring radio frequency data in a thermal ablation process, and processing the acquired radio frequency data; s2, related to fuzzy clustering; s3, performing iterative processing on fuzzy clustering according to the processed radio frequency data and set fuzzy clustering parameters to obtain a clustering result; and S4, displaying a thermal ablation area according to the obtained clustering result. On one hand, the difference between thermal ablation and normal tissue is maximized and highlighted from the perspective of signal processing; on the other hand, an unsupervised fuzzy clustering algorithm is used for automatically identifying the thermal ablation area, the scheme can be well compatible with a current ultrasonic system, and compared with other disclosed thermal ablation monitoring methods, the method has the advantages of being reliable in result, easy to operate and easy to implement. | 1·The thermal ablation area monitoring method based on radio frequency processing and fuzzy clustering is characterized by including the steps: S1. Obtain radio frequency data during the thermal ablation process, and process the obtained radio frequency data; S2. Set parameters related to fuzzy clustering; S3. According to the processed radio frequency data and the set fuzzy clustering parameters, iterative processing of fuzzy clustering is performed to obtain the clustering results; S4. Display the thermal ablation area based on the obtained clustering results; Step S1 specifically includes: SU. Acquire backscattered RF signals during thermal ablation; S12. Perform detection and demodulation processing on the acquired radio frequency signal to obtain the envelope signal; S13. Adjust the dynamic range of the obtained envelope signal to maximize the differentiation of the thermal ablation area; S14. Rearrange the envelope signals; In step S14, the envelope signals are rearranged in a column-first manner; where the envelope signals before rearrangement are two-dimensional matrices, with the number of rows and the number of columns respectively. are Row and Col; the rearranged envelope signal is one-dimensional data reshapeEnv, and the length is Row*Col; Step S3 specifically includes: S31. Initialize the membership matrix U; where the number of rows of the membership matrix U is cluster and the number of columns is Row*Col; S32. Calculate the current cluster center; S33. Update the membership matrix; S34. Determine whether the change in two adjacent iterations of the clustering loss function is less than the preset threshold. If so, stop the iteration; if not, re-execute step S32 based on the updated membership matrix; In step S32, calculate the current clustering Class center, expressed as: ZU1 (Λ j}reshapeEm<J) Center^ = ---Koreaerj----------------j = V...cluster j=1 Among them, Centeri represents the cluster center; In step S33, the membership matrix is updated, expressed as : I¦re5/jrtpp£iiv(j) - ceuter,"work U (lj) ------------------------J = j = 1Rou^Col Z Xu M Yan Shi M% m Zhao * where, u (i, j) represents the updated membership matrix; In step S34, it is judged whether the change in two adjacent iterations of the clustering loss function is less than the preset threshold, indicating For: Five empty things £ port? Teleport corpse dead Loss = Σ Σ U^J^reslmpeEm'ij)J=1 /=1 Among them, Loss represents the clustering loss function. 2. The thermal ablation area monitoring method based on radio frequency processing and fuzzy clustering according to claim 1, characterized in that the parameters in step S2 include the number of cluster categories and the iteration termination condition of fuzzy clustering; in step S2, Iteration of fuzzy clustering The generation termination condition is the change in two adjacent iterations of the clustering loss function. 3. The thermal ablation area monitoring method based on radio frequency processing and fuzzy clustering according to claim 1, characterized in that step S4 is specifically: calculating the average signal strength of the two clusters according to the clustering results; wherein, the average Clusters with high signal intensity correspond to thermal ablation areas. | 2020 | JURONG MEDICAL TECH HANGZHOU | NA | NA | NA | China | The patent focuses on a thermal ablation area monitoring method that processes radio-frequency data and uses fuzzy clustering to identify the ablated region. It addresses the difficulty of accurately delineating the ablation zone during thermal therapy from raw RF data. Fuzzy clustering produces a clearer picture of the ablation area for procedure monitoring. | הפטנט מתמקד בשיטת ניטור אזור אבלציה תרמית המעבדת נתוני תדר רדיו ומשתמשת באשכולות עמומים לזיהוי האזור שעבר אבלציה. הוא מתמודד עם הקושי בהגדרה מדויקת של אזור האבלציה במהלך טיפול תרמי מנתוני RF גולמיים. אשכולות עמומים מייצרים תמונה ברורה יותר של אזור האבלציה לניטור פרוצדורה. |  |  | 1 | [19 empty] | 1 | [7 empty] | 1 |  | [NA] | [NA] | [NA] | [NA] | JURONG MEDICAL TECH HANGZHOU CO LTD | 2020-03-24 | 2020 | 2024-05-03 | 2024 | 2020-03-24 | 2020 | Grant | Utility Invention | active | CN202010213186 | 286648252 | CN111429432A | CN
419 | CN119564326B | Pulse electric field ablation control device | The invention relates to the technical field of catheter mapping and ablation, and discloses a pulsed electric field ablation control device. The pulsed electric field ablation control device comprises: an impedance detection part for performing fine impedance detection on a contacted tissue by using an electrode on an ablation catheter; the tissue thickness determining part is used for determining the thickness of the tissue according to the change of the detected impedance under different contact conditions; and the ablation control part is used for controlling the ablation catheter to carry out pulsed electric field ablation output according to the determined thickness of the tissue. According to the focusing electrode, the impedance of the contact tissue and the tiny impedance change under different contact conditions can be accurately detected, so that the tissue thickness can be calculated, proper ablation parameters can be matched, and complications caused by too large ablation parameters are avoided. In addition, according to the pulse electric field ablation control technology, the ablation depth can be determined, and excessive ablation is avoided. | 1· A pulse electric field ablation control device, comprising: an impedance detection part, using electrodes on an ablation catheter to perform fine impedance detection on a contacted tissue; a tissue thickness determination part, determining the thickness of the tissue according to changes in the detected impedance under different contact conditions; a tissue site determination part, determining the site of the tissue that the electrode is contacting according to electrophysiological signals collected by the electrode; an ablation control part, controlling the ablation catheter to output pulse electric field ablation according to the site of the tissue and the determined thickness of the tissue, characterized in that the contact condition includes abutment pressure, wherein: the impedance detection part is configured to: collect impedance values for the tissue under different abutment pressure conditions, the tissue thickness determination part is configured to: plot the change characteristics of the collected impedance values as a curve, and compare them with a plurality of basic curves that are drawn and saved in advance through experimental data and reflect the change characteristics of impedance values of different tissue thicknesses, when the change characteristic curve of the collected impedance values matches one of the plurality of basic curves that are drawn and saved in advance through experimental data and reflect the change characteristics of impedance values of different tissue thicknesses, the thickness of the tissue is determined as the tissue thickness reflected by the matched basic curve. 2.The device according to claim 1 is characterized in that the ablation control part further includes an ablation parameter determination unit, which determines the ablation parameters based on the location of the tissue and the determined thickness of the tissue. 3.The device according to claim 2 is characterized in that the ablation parameters include: pulse discharge voltage, single effective discharge time, pulse width, and number of superpositions. 4.The device according to claim 1 is characterized in that the collected impedance value is the difference between the real-time detected impedance value and the impedance value in the blood. 5.The device according to claim 1 is characterized in that during the pulse electric field ablation output process, when the impedance detection part detects a real-time impedance mutation, the ablation control part is further configured to immediately stop the pulse electric field ablation output. 6.The device according to claim 1 is characterized in that the ablation control part is further configured to control whether the ablation catheter continues to output pulsed electric field ablation based on the comparison result between the real-time ablation depth and the thickness of the tissue. 7.The device according to claim 6 is characterized in that the device further includes an ablation depth determination part, which calculates the real-time ablation depth according to the ablation parameters and the real-time contact pressure during the pulse electric field ablation output process. 8. The device according to claim 7, characterized in that the ablation depth determination part calculates the real-time ablation depth according to the following formula: PFADepth = C * £, rCF(t) * furnace gas beating*, where PFADepth is the real-time ablation depth, V is the pulse discharge voltage, T is the single effective discharge time, u is the pulse width, n is the number of superpositions, CF is the real-time contact pressure, t is the time variable, and C, a, and β are correction coefficients. 9.The device according to claim 6 is characterized in that the ablation control part is further configured to: when the real-time ablation depth is less than the thickness of the tissue, control the ablation catheter to continue pulsed electric field ablation output; when the real-time ablation depth is equal to the thickness of the tissue, control the ablation catheter to stop ablation. 10The device according to claim 1, characterized in that the device further comprises an ablation lesion estimation part, The ablation damage range in the tissue is estimated based on the real-time damage detection result, and the ablation control part is further configured to: control whether the ablation catheter continues to output pulsed electric field ablation based on the ablation damage range in the tissue. 11.The device according to claim 10 is characterized in that the ablation catheter includes an optical sensor arranged between electrodes, and the ablation damage estimation part is configured to estimate the ablation damage range in the tissue based on the light signal detected by the optical sensor. 12.The device according to claim 10 is characterized in that the ablation catheter includes an ultrasonic sensor arranged between electrodes, and the ablation damage estimation part is configured to estimate the ablation damage range in the tissue based on the ultrasonic imaging signal detected by the ultrasonic sensor. 13.The device according to claim 10 is characterized in that the device further includes an ablation depth determination part, which calculates the real-time ablation depth according to the ablation parameters and the real-time contact pressure during the pulse electric field ablation output process, and the ablation damage estimation part is further configured to: correct the real-time damage detection result and the calculated real-time ablation depth to each other, and estimate the ablation damage range in the tissue. 14.The device according to claim 1 is characterized in that the contact condition also includes the direction and position relationship of the abutment and the shape of the ablation catheter. 15.The device according to claim 1 is characterized in that the electrodes on the ablation catheter are bipolar electrodes and the distance between electrodes of different polarities is minimized. | 2025 | SICHUAN JINJIANG ELECTRONIC MEDICAL INSTRUMENT \ | Medical and Healthcare | NA | http://www.jjet.cn/ | China | The patent focuses on a pulsed-field ablation control device with an impedance detection part for fine impedance measurement at the catheter electrode plus a tissue-thickness determining part. It addresses the gap that conventional PFA control doesn't account for fine local impedance and thickness when shaping ablation. Combined impedance and thickness data improve catheter mapping and ablation control. | הפטנט מתמקד במכשיר בקרת אבלציה בשדה פולסי עם חלק גילוי אימפדנס למדידת אימפדנס עדינה באלקטרודת הקטטר בתוספת חלק קביעת עובי רקמה. הוא מתמודד עם הפער שבקרת PFA קונבנציונלית אינה מתחשבת באימפדנס מקומי עדין ובעובי בעיצוב האבלציה. נתוני אימפדנס ועובי משולבים משפרים את מיפוי הקטטר ואת בקרת האבלציה. | 1 | 1 | [9 empty] | 1 | [9 empty] | 1 | [8 empty] | 1 |  | [NA] | [NA] | [NA] | [NA] | SICHUAN JINJIANG ELECTRONIC MEDICAL INSTRUMENT TECH CO LTD | 2025-02-10 | 2025 | 2025-06-10 | 2025 | 2025-02-10 | 2025 | Grant | Utility Invention | active | CN202510142752 | 319594831 | CN119564326A | CN
420 | CN114788729A | Catheter integrating ultrasonic imaging and radiofrequency ablation | The invention discloses an ultrasonic imaging and radiofrequency ablation integrated catheter which comprises a head electrode, a miniature ultrasonic probe is arranged in the middle of the head electrode, and the miniature ultrasonic probe scans the area right in front of the catheter to obtain the ablation degree of the ablation catheter on tissue; the head electrode is arranged at the top end of the catheter body; ring electrodes are arranged on the catheter head of the catheter body at intervals; the catheter body comprises a controllable bending section and a main body section; wherein the controllable bending section is arranged on the front side of the main body section; a pull wire is arranged in the catheter body; ablation tissue in the front direction or the side direction of the catheter is scanned through the miniature ultrasonic probe installed on the catheter, and the ablation degree of the ablation catheter to a focus is judged in real time in an ultrasonic imaging mode; the traditional mode of radiofrequency ablation operation is changed, and doctors are helped to accurately ablate ablation tissues. | 1. a catheter integrating ultrasound imaging and radiofrequency ablation, characterized in that it comprises a head electrode (1), a miniature ultrasonic probe (3) is arranged in the middle of the head electrode (1), and the area directly in front of the catheter is scanned to obtain a The degree of tissue ablation by the ablation catheter is obtained; the head electrode (1) is arranged on the top of the catheter body (4); the catheter head of the catheter body (4) is provided with ring electrodes (2) at intervals. 2. The catheter integrating ultrasound imaging and radio frequency ablation according to claim 1, wherein the catheter body (4) comprises a controllable bending section and a main body section; wherein the controllable bending section is arranged in front of the main body section side; the catheter body (4) is provided with a traction wire inside. 3. a catheter integrating ultrasonic imaging and radio frequency ablation, is characterized in that, comprises miniature ultrasonic probe (3); Described miniature ultrasonic probe (3) is located in two ring electrodes (2) spaced part, scans catheter body (4) lateral area, to obtain ablation catheter For the degree of tissue ablation, a plurality of miniature ultrasonic probes are evenly arranged along the circumference of the catheter. 4. A catheter integrating ultrasonic imaging and radio frequency ablation according to claim 3, characterized in that, a temperature sensor (5) is also provided on the inner wall of the catheter body (4); the bottom of the temperature sensor (5) is provided with a temperature sensor (5). There are connecting wires that go through the inside of the catheter to an external controller. 5. The catheter integrating ultrasound imaging and radio frequency ablation according to claim 3, wherein the temperature sensor (5) is a type of sensor that can work normally at a temperature of less than 70°C. 6. a kind of catheter integrating ultrasonic imaging and radio frequency ablation according to claim 1, is characterized in that, described head electrode (1), ring electrode (2) and miniature ultrasonic probe (3) rear side are all provided with lead wire, so The wire is fixed inside the catheter body (4). 7. A catheter integrating ultrasound imaging and radio frequency ablation according to claim 3, wherein the temperature sensor (5) is provided with a wire on the rear side, and the wire is fixed inside the catheter body (4). | 2021 | PLA GENERAL HOSPITAL IN NORTHERN THEATER | NA | NA | NA | China | The patent focuses on an integrated ultrasound imaging and RF ablation catheter where a miniature ultrasound probe sits in the middle of the head electrode and scans the area in front of the catheter. It addresses the operator's lack of direct visualization of the area being RF-ablated, which limits real-time confirmation of treatment effect. The miniature ultrasound probe shows the ablation degree right in front of the tip. | הפטנט מתמקד בקטטר משולב להדמיית אולטרסאונד ואבלציית RF שבו גשוש אולטרסאונד מיניאטורי יושב באמצע אלקטרודת הראש וסורק את האזור שלפני הקטטר. הוא מתמודד עם החוסר של המפעיל בהדמיה ישירה של האזור שעובר אבלציית RF, המגביל אישור בזמן אמת של אפקט הטיפול. גשוש האולטרסאונד המיניאטורי מראה את דרגת האבלציה ממש לפני הקצה. |  |  | 1 |  | 1 | [5 empty] | 1 | [19 empty] | 1 |  | PLA GENERAL HOSPITAL IN NORTHERN THEATER | PLA GENERAL HOSPITAL IN NORTHERN THEATER | PLA GENERAL HOSPITAL IN NORTHERN THEATER | PLA GENERAL HOSPITAL IN NORTHERN THEATER | PLA GENERAL HOSPITAL IN NORTHERN THEATER | 2021-01-24 | 2021 | 2022-07-26 | 2022 | 2021-01-24 | 2021 | Application | Utility Invention | active | CN202110108278 | 283866771 | [NA] | CN
421 | WO2025219165A1 | SYSTEM FOR AUTOMATIC TEMPERATURE REGULATION FOR AN MRI-GUIDED MULTI-EMITTER LASER THERMAL THERAPY DEVICE | The invention relates to a method for regulating temperature in a plurality of regions of interest (R1, R2, R3) of a biological tissue (T), which regions are subjected to a plurality of laser beams of predetermined wavelengths which have adjustable laser emission powers and/or durations and which originate from a multi-emitter laser apparatus (4) within a device (S) for MRI-guided thermal therapy. The method comprises steps of generating and visualising thermal images Tn(x, y, z) in said plurality of regions of interest (R1, R2, R3) by processing complex modulus and phase images originating from a device (20) for magnetic resonance imaging (MRI). This method further comprises real-time processing of the thermal images, taking into account a set of planning parameters, in order to output set values for regulating the power and/or the duration of energy emission of all or part of said laser beams which are applied to one or more of the regions of interest (R1, R2, R3). | 1. System (1) for regulating the temperature in a plurality of regions of interest (R1, R2, R3; Ri) of a biological tissue, in each of which is inserted the emitting head of an optical fiber among a plurality of heads emitting laser beams of predetermined wavelengths and adjustable power or emission durations, produced by a multi-emitter laser source (4) within MRI-guided thermotherapy equipment (S), this system (1) comprising:##TABLE## characterized in that the real-time processing means (31) of the temperature images comprise a controller implementing a regulation algorithm (31) of the Proportional-Integral-Derivative (PID) type integrating the resolution of the heat equation using a fast Fourier transformation (FFT), this regulation algorithm being arranged to deliver for each laser beam “j” after the calculation of the “xi »at iteration “n”, an energy setpoint EI(n) of the type:with:n representing time, j representing each laser emitter, PI representing the laser power delivered by the laser transmitter “j” during the interval ,Q representing a system of equations allowing to link the contributions of the spatial distributions of thermal energy SI to temperature risesxi in each region of interest Ri Ri representing a region of interest to be controlled in temperature, xi representing the evolution of temperaturenecessary in each region Ri between time "n" and time to follow a set temperatureSI representing the spatial distribution of thermal energy delivered in the biological tissue by the emitting head FI. 2. Regulation system according to the preceding claim, characterized in that the regulation algorithm uses the following input data:##TABLE## 3. Control system according to any one of the preceding claims, characterized in that the set of planning parameters can advantageously comprise a temperature time setpoint curve (There(n)) for a given region of interest, as well as one or more parameters of the biological tissue to be treated and the heating characteristics SI of the emitting head of each optical fiber FI. 4. Control system according to any one of the preceding claims, characterized in that the set of equations Q is expressed in the following form: 5. Regulation system (1) according to any one of the preceding claims, characterized in that the means for processing the temperature images are integrated into the magnetic resonance imaging (MRI) equipment. 6. Regulation system according to any one of claims 1 to 4, characterized in that the means for processing the temperature images are integrated into an image reconstruction unit downstream of the magnetic resonance imaging (MRI) equipment. 7. MRI-guided laser thermotherapy equipment (S), comprising means (7, 8, 9) for emitting a plurality of laser beams of determined wavelengths and adjustable emission powers or durations and means (F1, F2, F3; FI) to apply said laser beams respectively in a plurality of regions of interest (R1,R2,R3;Ri) of a biological tissue, and magnetic resonance medical imaging equipment (20) arranged to deliver and display temperature images Tn(x,y,z) of one or more of said one or more regions of interest (R1,R2,R3;Ri), characterized in that it further comprises a system (1) for regulating the temperature in said one or more regions of interest (R1, R2, R3; Ri) according to any one of the preceding claims. 8. MRI-guided laser thermotherapy equipment according to the preceding claim, characterized in that the means for applying the laser beams comprise a plurality of optical fibers powered by laser emitters and having respective emitting heads arranged to be inserted into a region of interest of a biological tissue and to emit therein laser light of determined wavelength and adjustable power or emission duration. 9. MRI-guided laser thermotherapy equipment according to the preceding claim, characterized in that at least one of the optical fibers comprises an emitting head (Fa) arranged to diffuse the laser light within the biological tissue in which said at least one emitting head (Fa) is inserted. 10. MRI-guided laser thermotherapy equipment according to one of the two preceding claims, characterized in that at least one of the optical fibers comprises an emitting head (Fb) arranged to axially emit the laser light within the biological tissue in which said at least one emitting head (Fb) is inserted. 11. MRI-guided laser thermotherapy equipment according to one of the three preceding claims, characterized in that at least one of the optical fibers comprises an emitting head (Fc) arranged to radially emit the laser light within the biological tissue in which said at least one emitting head (Fc) is inserted. 12. MRI-guided laser thermotherapy equipment according to one of the four preceding claims, characterized in that at least one of the optical fibers comprises an emitting head (Fd) arranged to emit laser light at a predetermined angle relative to the axis of said optical fiber within the biological tissue in which said at least one emitting head is inserted. 13. MRI-guided laser thermotherapy equipment according to one of the five preceding claims, characterized in that the means for applying the laser beams further comprise one or more emitting heads having a source function adapted to be used as input data for the temperature regulation system according to claim 8. | 2024 | CERTIS THERAPEUTICS | Medical and Healthcare | Start up | http://www.certis-therapeutics.com/ | France | The patent focuses on a method for automatic temperature regulation across multiple regions of interest in tissue subjected to multiple laser beams from a multi-emitter laser apparatus during MRI-guided thermal therapy. It addresses the difficulty of keeping multiple tissue regions at desired temperatures simultaneously during multi-emitter laser therapy. The method automatically adjusts laser emission powers and durations to regulate temperature across regions. | הפטנט מתמקד בשיטה לוויסות טמפרטורה אוטומטי על פני מספר אזורי עניין ברקמה החשופה למספר אלומות לייזר ממכשיר לייזר רב-פולט במהלך טיפול תרמי מונחה MRI. הוא מתמודד עם הקושי בשמירה על מספר אזורי רקמה בטמפרטורות הרצויות בו זמנית במהלך טיפול לייזר רב-פולט. השיטה מתאימה אוטומטית את עוצמות הפליטה ומשכי הלייזר לוויסות טמפרטורה על פני אזורים. |  |  |  |  | 1 |  |  |  | 1 |  |  |  |  | 1 | [5 empty] | 1 | [10 empty] | 1 |  | CERTIS THERAPEUTICS¦UNIV BORDEAUX¦CENTRE NAT RECH SCIENT | CERTIS THERAPEUTICS¦University of Bordeaux¦French National Centre for Scientific Research (CNRS) | CERTIS THERAPEUTICS¦UNIV BORDEAUX¦CENTRE NAT RECH SCIENT | CERTIS THERAPEUTICS¦University of Bordeaux¦French National Centre for Scientific Research (CNRS) | CERTIS THERAPEUTICS¦University of Bordeaux¦French National Centre for Scientific Research (CNRS) | 2025-04-08 | 2025 | 2025-10-23 | 2025 | 2024-04-15 | 2024 | Application | Utility Invention | active | WO2025EP59656 | 323959114 | FR3161119A1 | WO
422 | CN116172695B | Interventional intravascular multi-mode imaging and ablation integrated catheter | The invention discloses an interventional intravascular multi-mode imaging and ablation integrated catheter, belongs to the technical field of intravascular disease diagnosis, integrates optical, acoustic and electric paths of photoacoustic/ultrasonic/OCT/temperature four-mode imaging and photothermal ablation, and overcomes the defect that a traditional interventional therapy catheter cannot realize multi-mode imaging and ablation. The catheter comprises a catheter body, a torque spring arranged on the outer side of the catheter body in a sleeving mode and an outer skin of the catheter body. A metal shell used for reinforcing and protecting internal components is arranged at the front end of the catheter body, and a multi-mode imaging component and a laser ablation component are integrated in the metal shell. When the catheter is used for interventional operation, accurate structural component information and temperature distribution information of focus tissue can be provided, accurate positioning of a treatment boundary is achieved, micron-sized high-precision photo-thermal ablation treatment is completed, and the problems of high-resolution real-time imaging and multi-mode accurate positioning in a blood vessel are effectively solved; lateral light emitting of imaging laser and treatment laser and focusing depth adjustment of the treatment laser are achieved. | 1.An interventional intravascular multi-modal imaging and ablation integrated catheter, characterized by: including a catheter body (3), a torque spring (1) set on the outside of the front end of the catheter body (3), and an outer skin (2) of the catheter body (3). ); The front end of the catheter body (3) is provided with a metal shell (13) for reinforcing and protecting the internal components, and the four-modal imaging component of photoacoustic, ultrasound, OCT and temperature and the laser ablation component are integrated inside the metal shell (13) , the photoacoustic, ultrasound, OCT and temperature four-modal imaging component includes a first optical fiber (5), a reflective prism (11) and an ultrasonic transducer (14); the first optical fiber (5) is a photoacoustic The beam path of the modal and OCT modes can provide data to the back-end computer to realize the temperature mode. The front end of the first optical fiber (5) is provided with a self-focusing lens (16), a reflective prism (11) and an ultrasonic transducer (14) in sequence. The photoacoustic mode laser and OCT mode scanning light transmitted by the first optical fiber (5) are focused by the self-focusing lens (16) and reflected by the reflection prism (11) in sequence, and then emit at a certain angle. The photoacoustic mode laser is absorbed by the imaging tissue. , generate photoacoustic signals,Detected by the ultrasonic transducer (14); after the OCT mode scanning light irradiates the tissue, coherent light is generated and propagates backward along the emitted light path to the coupler, and is finally imaged through the back-end OCT equipment. The laser ablation component includes a third Two optical fibers (7), liquid lens voltage control line (8) and liquid lens (10); the second optical fiber (7) is a path for ablation laser transmission, and a liquid lens (11) is installed between the light output of the second optical fiber (7) and the reflecting prism (11) 10), to realize the focus control of the ablation laser, by adjusting the voltage of the liquid lens voltage control line , the curvature of the droplet of the liquid lens (10) is controlled, thereby controlling the focal length, realizing the control of the focus depth of the ablation laser, and the ablation laser irradiation after focusing to the reflecting prism (11), and then illuminated on the tissue after passing through the reflecting prism (11), thereby achieving lateral light emission and adjustable focus depth of the beam. Both the imaging laser of the first optical fiber and the ablation laser of the second optical fiber can be irradiated. To the reflecting prism (11), the imaging light and the ablation light are emitted coaxially. 2. An interventional intravascular multi-modal imaging and ablation integrated catheter according to claim 1, characterized in that: the first optical fiber (5) adopts a wide-band single-mode optical fiber. 3. An interventional intravascular multi-modal imaging and ablation integrated catheter according to claim 2, characterized in that: the first optical fiber (5), the liquid lens voltage control line (8), the second optical fiber (7) The ultrasonic electric signal conduction lines (6) of the ultrasonic transducer (14) are integrated into the catheter body (3), and the inside of the catheter body (3) is filled with flame retardant and insulating materials (4). 4. A kind of interventional intravascular multi-modal imaging and ablation integrated catheter according to claim 3, characterized in that: the rear end of the catheter body (3) is equipped with a quick connection plug, and the quick connection plug is matched with the rear end. device to connect. 5. An interventional intravascular multi-modal imaging and ablation integrated catheter according to claim 4, characterized in that: the rotating part of the quick connection plug is a rotating shell part (21), and the rotating shell part (21) ) is provided with threads on its inner circumferential surface to match the rear-end interface. The plug portion (22) of the quick-connect plug is equipped with a photoacoustic and OCT imaging interface (23), an ablation laser interface (20), and an ultrasonic electrical signal interface ( 19) and the liquid lens control interface (18), and a plug fixing pin (17) is provided on the plug part (22), and a firm connection with the back-end equipment is achieved through the plug fixing pin (17). | 2023 | Harbin Institute of Technology | University | NA | http://www.hit.edu.cn | China | The patent focuses on an interventional intravascular catheter that integrates photoacoustic, ultrasound, OCT, and temperature imaging plus photothermal ablation in one device. It addresses the limitation of traditional intravascular catheters that lack the multi-modal imaging needed to guide ablation precisely. The combined modalities support real-time imaging and treatment of intravascular disease. | הפטנט מתמקד בקטטר תוך-וסקולרי התערבותי המשלב הדמיה פוטואקוסטית, אולטרסאונד, OCT וטמפרטורה בתוספת אבלציה פוטותרמית במכשיר אחד. הוא מתמודד עם המגבלה של קטטרים תוך-וסקולריים מסורתיים החסרים את ההדמיה הרב-מצבית הנדרשת להנחיית אבלציה במדויק. המודאליות המשולבות תומכות בהדמיה וטיפול בזמן אמת של מחלות תוך-וסקולריות. |  | 1 | 1 | 1 | 1 | [8 empty] | 1 | [16 empty] | 1 |  | [NA] | [NA] | [NA] | [NA] | Harbin Institute of Technology¦Chinese Academy of Sciences | 2023-03-02 | 2023 | 2024-03-01 | 2024 | 2023-03-02 | 2023 | Grant | Utility Invention | active | CN202310192378 | 289146033 | CN116172695A | CN
423 | CN120274910A | Ablation heat metering determination method based on nuclear magnetic resonance detection signal | The invention relates to the field of tumor ablation, in particular to an ablation heat metering determination method based on a nuclear magnetic resonance detection signal, which comprises the following steps: determining a function relationship between an effective heat conductivity coefficient and a relative contrast ratio of a tissue according to multiple groups of temperature data acquired by performing thermal ablation on a tumor model; determining the relative contrast according to the T1 signal and the T2 signal of the tumor tissue and the healthy tissue of the patient; determining a heat conduction coefficient according to the relative contrast and the function relationship; and according to the heat conduction coefficient, determining the ablation heat quantity required for performing tumor thermal ablation on the patient. | 1. A method for determining ablation heat dosage based on nuclear magnetic resonance detection signals, the method comprising: determining the functional relationship between the effective thermal conductivity of tissue and the relative contrast based on multiple sets of temperature data collected for thermal ablation of a tumor model; determining the relative contrast of the patient based on the T1 signals and T2 signals of the patient's tumor tissue and healthy tissue respectively; determining the thermal conductivity coefficient based on the relative contrast and the functional relationship; determining the ablation heat dosage required for thermal ablation of the patient's tumor based on the thermal conductivity coefficient. 2.The method as claimed in claim 1, wherein determining the functional relationship between the effective thermal conductivity of tissue and the relative contrast based on multiple groups of temperature data collected by thermal ablation of tumor models includes: performing nuclear magnetic resonance detection on the established multiple groups of tumor models to determine the T1 signal and T2 signal of the healthy tissue of each group of tumor models, as well as the T1 signal and T2 signal of the tumor tissue; determining the relative contrast of each group of tumor models based on the T1 signal and T2 signal of each group of tumor models; performing thermal ablation on the tumor according to the ablation probe and the temperature measurement probe set in the tumor model, and collecting tumor boundary temperature data and ablation probe needle tip temperature data during the thermal ablation process; determining the thermal conductivity of each group of tumor models using the tumor boundary temperature data and ablation probe needle tip temperature data of each group; performing function fitting based on the thermal conductivity and relative contrast of multiple groups of tumor models to determine the functional relationship. 3.The method as claimed in claim 2, wherein the tumor model is suitable for being established by inoculating a tumor into an animal tissue. 4.The method according to claim 2 or 3, wherein the T1 signal and T2 signal are suitable for being determined by performing nuclear magnetic resonance imaging on the patient. 5.The method as claimed in claim 2 or 3, wherein determining the relative contrast of each group of tumor models based on the T1 signal and T2 signal of each group of tumor models includes: determining the relative contrast based on the ratio of the tumor tissue signal ratio and the healthy tissue signal ratio, the tumor tissue signal ratio includes the ratio of the T2 signal of the tumor tissue to the T1 signal, and the healthy tissue signal ratio includes the ratio of the T2 signal of the healthy tissue to the T1 signal. 6.The method as claimed in claim 2, wherein the step of determining the thermal conductivity coefficient of each group of tumor models based on the tumor boundary temperature data and the ablation probe tip temperature data of each group comprises: substituting the tumor boundary temperature data and the ablation probe tip temperature data into the heat transfer equation to determine the thermal conductivity coefficient of each group of tumor models. 7.The method as claimed in claim 1, wherein determining the relative contrast of the patient based on the respective T1 signals and T2 signals of the patient's tumor tissue and healthy tissue comprises: performing multi-point detection on the tumor tissue and the healthy tissue to determine the T1 signals and T2 signals of multiple points of the patient's tumor tissue, as well as the T1 signals and T2 signals of multiple points of the patient's healthy tissue; determining an average value of multiple tumor tissue signal ratios based on the T1 signals and T2 signals of multiple points of the patient's tumor tissue; determining an average value of multiple healthy tissue signal ratios based on the T1 signals and T2 signals of multiple points of the patient's healthy tissue; and determining the relative contrast based on the average value of multiple tumor tissue signal ratios and the average value of multiple healthy tissue signal ratios. 8. The method as claimed in claim 1, wherein determining the amount of ablation heat required for thermal ablation of a patient according to the thermal conductivity coefficient comprises: substituting the thermal conductivity coefficient or the thermal conductivity correction coefficient into the biological heat transfer equation to determine the amount of ablation heat required for thermal ablation of the tumor. Heat of melt measurement. 9.A computing device, characterized in that it comprises: one or more processors; a memory; and one or more devices, wherein the one or more devices include instructions for executing the method according to any one of claims 1-8. 10. A computer-readable storage medium storing one or more programs, characterized in that the one or more programs include instructions, which, when executed by a computing device, cause the computing device to perform the method according to any one of claims 1-8. | 2025 | Shanghai Jiao Tong University | University | NA | https://www.sjtu.edu.cn | China | The patent focuses on an ablation heat-metering method based on nuclear magnetic resonance signals, deriving a function relationship between effective heat conductivity and tissue contrast. It addresses the difficulty of quantifying ablation heat in tissue without dedicated thermometry hardware. NMR-derived heat measurement supports tumor ablation planning and verification. | הפטנט מתמקד בשיטת מדידת חום אבלציה מבוססת אותות תהודה מגנטית גרעינית, הגוזרת יחס פונקציה בין מוליכות חום אפקטיבית לניגודיות רקמה. הוא מתמודד עם הקושי בכימות חום אבלציה ברקמה ללא חומרת תרמומטריה ייעודית. מדידת חום נגזרת מ-NMR תומכת בתכנון ואימות אבלציית גידולים. | [8 empty] | 1 | [19 empty] | 1 |  |  |  | UNIV SHANGHAI JIAOTONG | Shanghai Jiao Tong University | UNIV SHANGHAI JIAOTONG | Shanghai Jiao Tong University | Shanghai Jiao Tong University | 2025-03-21 | 2025 | 2025-07-08 | 2025 | 2025-03-21 | 2025 | Application | Utility Invention | inactive | CN202510336777 | 322338736 | CN120274910B | CN
424 | WO2022192522A1 | DEVICES FOR THE DELIVERY OF PULSED ELECTRIC FIELDS IN THE TREATMENT OF CARDIAC TISSUE | Devices, systems and methods are provided for treating conditions of the heart, particularly the occurrence of arrhythmias, more particularly atrial fibrillation, atrial flutter, ventricular tachycardia, to name a few. The devices, systems and methods deliver therapeutic pulsed electric field energy to portions the heart to provide tissue modification, such as to the entrances to the pulmonary veins in the treatment of atrial fibrillation. Such tissue modification creates a conduction block within the tissue to prevent the transmission of aberrant electrical signals. Generally, the tissue modification systems include a specialized catheter, a high voltage waveform generator and at least one distinct energy delivery algorithm. Example embodiments of specialized catheter designs are provided and include a variety of delivery types including focal delivery, "one-shot" delivery and various possible combinations. | 1. WHAT IS CLAIMED IS: 1. A device for delivering energy to cardiac tissue of a patient comprising: a shaft having a proximal end and a distal end, wherein the shaft has an outer diameter; and an energy delivery body disposed along the distal end of the shaft, wherein the energy delivery body is transitionable between a collapsed configuration and an expanded configuration, wherein the expanded configuration has an outer diameter that is less than or equal to 6 times the outer diameter of the shaft, wherein the energy delivery body is configured to be positioned against the cardiac tissue in the expanded configuration so as to deliver energy to the cardiac tissue. 2. A device as in claim 1, wherein the energy is pulsed electric field energy and wherein the device is configured to deliver the pulsed electric field energy to the cardiac tissue. 3. A device as in any of the above claims, wherein the energy delivery body comprises a plurality of wires configured to deliver the energy. 4. A device as in claim 3, wherein the plurality of wires comprises a plurality of splines. 5. A device as any of claims 3-4, wherein the plurality of wires is comprised of shape memory material so that the energy delivery body is transitionable by release from a sheath that constrains the plurality of wires so that such release allows the plurality of wires to move toward the expanded configuration. 6. A device as in claim 5, wherein the energy delivery body is free of a central shaft when in the expanded configuration. 7. A device as in claim 5, wherein the plurality of splines form a hollow rounded cage. 8. A device as in claim 3, wherein the plurality of wires comprises a mesh. 9. A device as in claim 3, wherein the plurality of wires comprises a plurality of loops. 10. A device as in any of claims 3-9, wherein the plurality of wires is energizeable in unison so as to function in a monopolar fashion. 11. A device as in any of claims 3-10, wherein the plurality of wires has a convex distal face. 12. A device as in any of claims 3-11, wherein the energy delivery body includes a distal tip configured to deliver the energy. 13. A device as in claim 12, wherein the distal tip and the plurality of wires are energizeable in unison so as to function in a monopolar fashion. 14. A device as in any of claims 3-13, wherein a proximal portion of the plurality of wires is insulated so as to direct the energy toward a distal direction. 15. A device as in any of the above claims, further comprising a plurality of irrigation ports, wherein the device is configured so as to direct fluid through the irrigation ports in a manner that creates turbulent flow of the fluid within the energy delivery body. 16. A device as in claim 15, wherein the plurality of irrigation ports is disposed near a proximal end of the energy delivery body. 17. A device as in any of claims 15-16, further comprising one or more irrigation lumens which direct the fluid through the plurality of irrigation ports. 18. A device as in claim 17, wherein the one or more irrigation lumens is less than the plurality of irrigation ports. 19. A device as in any of the above claims, wherein the expanded configuration has an outer diameter that is 3-6 times the outer diameter of the shaft. 20. A device as in any of the above claims, wherein the expanded configuration has an outer diameter that is 8-15mm. 21. A device as in any of the above claims, wherein the energy delivery body includes a sensing electrode. 22. A device as in claim 21, wherein the sensing electrode is positioned so as to avoid contact with the cardiac tissue. 23. A device as in claim 22, wherein the energy delivery body comprises a plurality of splines forming a rounded cage and wherein the sensing electrode is disposed within the rounded cage. 24. A device as in any of the above claims, wherein the shaft includes one or more ring electrodes. 25. A device as in any of the above claims, further comprising one or more electrodes that communicate with an electrophysiological mapping system. 26. A device as in any of the above claims, further comprising a steering mechanism configured to bend the energy delivery body in relation to the shaft. 27. A device as in any of the above claims, further comprising a steering mechanism configured to bend the distal end of the shaft away from its longitudinal axis. 28. A device for delivering energy to cardiac tissue of a patient comprising: a shaft having a proximal end and a distal end; and an energy delivery body disposed along the distal end of the shaft, wherein the energy delivery body is comprised of a plurality of shape-memory splines and is transitionable between a collapsed configuration and an expanded configuration, wherein in the expanded configuration the plurality of shape-memory splines form a convex distal face positionable against the cardiac tissue so as to deliver energy to the cardiac tissue. 29. A device as in claim 28, wherein the plurality of shape-memory splines form a rounded cage having the convex distal face in the expanded configuration. 30. A device as in claim 29, wherein the rounded cage is supported solely by the plurality of splines. 31. A device as in claim 30, wherein the energy delivery body includes a distal tip electrode and wherein the plurality of splines supports a tip electrode wire extending from the distal tip electrode to the shaft and is otherwise hollow. 32. A device as in any of claims 29-31, wherein the rounded cage is flexible so as to deform upon positioning against the cardiac tissue. 33. A device as in any of claims 29-32, wherein the rounded cage is flexible so as to at least partially flatten upon positioning against the cardiac tissue. 34. A device as in any of claims 28-33, wherein the convex distal face is configured to have a footprint of 8- 15mm when positioned against the cardiac tissue so as to deliver energy to the cardiac tissue. 35. A device as in any of the above claims, wherein the plurality of splines are energizeable in unison to function in a monopolar manner. 36. A device as in any of the above claims, wherein the energy delivery body includes a distal tip electrode disposed along the convex distal face. 37. A device as in claim 36, wherein the distal tip electrode is independently energizable. 38. A device as in any of the above claims, wherein at least a portion of the energy delivery body is insulated so as to direct the energy through the convex distal face. 39. A device as in any of the above claims, further comprising at least one irrigation lumen and a plurality of irrigation ports. 40. A device as in claim 39, wherein the at least one irrigation lumen comprises a number of irrigation lumens that is less than plurality of irrigation ports. 41. A device as in claim 29, wherein the energy delivery body is electrically couplable with a generator so as to deliver pulsed electric field energy to the cardiac tissue. 42. A system for delivering energy to cardiac tissue of a patient comprising: a treatment catheter comprising a shaft having a proximal end and a distal end, wherein the shaft has an outer diameter, and an energy delivery body disposed along the distal end of the shaft, wherein the energy delivery body is transitionable between a collapsed configuration and an expanded configuration, wherein the expanded configuration has an outer diameter that is less than or equal to 6 times the outer diameter of the shaft, wherein the energy delivery body is configured to be positioned against the cardiac tissue in the expanded configuration so as to deliver energy to the cardiac tissue; and a generator electrically couplable to the treatment catheter, wherein the generator includes at least one energy delivery algorithm configured to provide an electric signal of pulsed electric field energy deliverable through the energy delivery body. 43. A system as in claim 42, wherein the pulsed electric field energy is below a threshold for inducing coagulative thermal damage. 44. A system for treating cardiac tissue of a patient comprising: a treatment device comprising a shaft having a proximal end and a distal end, and an energy delivery body disposed along the distal end of the shaft, wherein the energy delivery body is configured to be positioned against the cardiac tissue, and wherein the energy delivery body is electrically couplable with a generator so as to deliver pulsed electric field energy to the cardiac tissue; and a generator electrically couplable to the treatment catheter, wherein the generator includes at least one energy delivery algorithm configured to provide an electric signal of pulsed electric field energy deliverable through the energy delivery body. 45. A system as in claim 44, wherein the pulsed electric field energy is below a threshold for inducing coagulative thermal damage. 46. A system for delivering energy to cardiac tissue of a patient comprising: A treatment catheter comprising a shaft having a proximal end and a distal end, and an energy delivery body disposed along the distal end of the shaft, wherein the energy delivery body is comprised of a plurality of shape-memory splines and is transitionable between a collapsed configuration and an expanded configuration, wherein in the expanded configuration the plurality of shape-memory splines form a convex distal face positionable against the cardiac tissue so as to deliver energy to the cardiac tissue; and a generator electrically couplable to the treatment catheter, wherein the generator includes at least one energy delivery algorithm configured to provide an electric signal of pulsed electric field energy deliverable through the energy delivery body. 47. A system as in claim 46, wherein the pulsed electric field energy is below a threshold for inducing coagulative thermal damage. | 2021 | GALARY | NA | NA | NA | NA | The patent focuses on devices for delivering pulsed electric field energy to portions of the heart, particularly the entrances to pulmonary veins for atrial fibrillation/flutter and ventricular tachycardia treatment. It addresses the need for targeted PFA tissue modification at specific cardiac sites without thermal collateral damage. The devices deliver PEF energy to selectively modify cardiac tissue at key sites. | הפטנט מתמקד במכשירים למסירת אנרגיית שדה חשמלי פולסי לחלקים בלב, במיוחד הכניסות לוורידים הריאתיים לטיפול בפרפור פרוזדורים/רפרוף וטכיקרדיה חדרית. הוא מתמודד עם הצורך בשינוי רקמה ב-PFA ממוקד באתרים לבביים ספציפיים ללא נזק נלווה תרמי. המכשירים מספקים אנרגיית PEF כדי לשנות באופן סלקטיבי רקמה לבבית באתרים מרכזיים. | 1 |  |  |  | 1 |  |  |  |  | 1 |  | 1 | [12 empty] | 1 | [7 empty] | GALARY INC | GALARY INC | GALARY INC | GALARY INC | GALARY INC | 2022-03-10 | 2022 | 2022-09-15 | 2022 | 2021-03-10 | 2021 | Application | Utility Invention | inactive | WO2022US19719 | 318666742 | AU2022234310A1¦CA3212976A1¦CN117615727A¦EP4304508A1¦EP4304508A4¦JP2024509275A¦JP2026035771A¦US20230414270A1 | WO
425 | US12329468B1 | Apparatus and methods for identification of medical features | Apparatus for identification of medical features and methods used therein include an ultrasonic imaging system configured to detect at least a query image containing ultrasound data pertaining to a subject, an intervention system configured to perform a concurrent medical procedure as a function of the detection of the at least a query image, a display device, a processor, and a memory communicatively connected to the processor, wherein the memory contains instructions configuring the processor to receive, from the ultrasonic imaging system, the at least a query image, identify the pose of an anatomical structure as a function of the at least a query image, determine at least a medical feature as a function of the pose of the anatomical structure, and suggest, using the display device, a modification to the concurrent medical procedure as a function of the at least a medical feature. | 1. An apparatus for identifying medical features, the apparatus comprising: an ultrasonic imaging system configured to detect at least a query image comprising ultrasound data pertaining to a subject; an intervention system configured to perform a concurrent medical procedure as a function of the detection of the at least a query image; a display device; a processor; and a memory communicatively connected to the processor, wherein the memory contains instructions configuring the processor to: receive, from the ultrasonic imaging system, the at least a query image; identify a pose of an anatomical structure as a function of the at least a query image, wherein identifying the pose of the anatomical structure comprises: locating at least a point of view (POV) as a function of the query image and a first subset of the plurality of image processing algorithms; and determining a view angle (VA) as a function of the at least a POV, wherein the at least a POV and the VA define at least a portion of the anatomical structure; determine at least a medical feature as a function of the pose of the anatomical structure, wherein determining the at least a medical feature comprises: receiving feature training data comprising a plurality of training images as input and a plurality of labelled features as output; training an image processing machine learning model comprising a plurality of image processing algorithms using the feature training data; and determining the at least a medical feature as a function of the at least a query image using the trained image processing machine learning model and a second subset of the plurality of image processing algorithms; and suggest, using the display device, a modification to the concurrent medical procedure as a function of the at least a medical feature. 2. The apparatus of claim 1, wherein: the intervention system comprises a catheter configured to perform an ablation procedure; determining the at least a medical feature comprises identifying a thrombus in a left atrial appendage of a heart; and suggesting the modification to the concurrent medical procedure comprises terminating the ablation procedure upon identifying the thrombus. 3. The apparatus of claim 2, wherein determining the at least a medical feature comprises identifying a case of calcification in a heart valve. 4. The apparatus of claim 3, wherein suggesting the modification to the concurrent medical procedure comprises avoiding, using the catheter, the heart valve upon identifying the case of calcification. 5. The apparatus of claim 1, wherein: the plurality of labelled features comprises a plurality of heart valve functions; and determining the at least a medical feature comprises: classifying, using a classifier, the feature training data into a plurality of classes as a function of the plurality of heart valve functions, wherein each class of the plurality of classes shares at least one similar heart valve function; and classifying an abnormal function of a heart valve as a function of the plurality of classes. 6. The apparatus of claim 1, wherein: determining the at least a medical feature comprises determining a thickness pertaining to the anatomical structure; the plurality of labelled features comprises a plurality of thicknesses; and determining the at least a medical feature comprises: labeling, using a labeling algorithm, the feature training data with a plurality of labels correlated with the plurality of thicknesses, wherein each label of the plurality of labels represents a thickness; and determining the thickness as a function of the plurality of labels. 7. The apparatus of claim 6, wherein suggesting the modification to the concurrent medical procedure comprises adjusting at least a parameter of an ablation procedure as a function of the thickness. 8. The apparatus of claim 1, wherein the at least a query image comprises an intracardiac echocardiogram (ICE). 9. The apparatus of claim 1, wherein the at least a query image comprises a point-of-care ultrasound (POCUS). 10. A method for identifying medical features, the method comprising: detecting, using an ultrasonic imaging system, at least a query image comprising ultrasound data pertaining to a subject; performing, using an intervention system, a concurrent medical procedure as a function of the detection of the at least a query image; receiving, by a processor from the ultrasonic imaging system, the at least a query image; identifying, by the processor, a pose of an anatomical structure as a function of the at least a query image, wherein: identifying the pose of the anatomical structure comprises: locating at least a POV as a function of the query image and a first subset of the plurality of image processing algorithms; and determining a VA as a function of the at least a POV, wherein the at POV and the VA define at least a portion of the anatomical structure; determining, by the processor, at least a medical feature as a function of the pose of an anatomical structure, wherein determining the at least a medical feature comprises: receiving feature training data comprising a plurality of training images as input and a plurality of labelled features as output; training an image processing machine learning model comprising a plurality of image processing algorithms using the feature training data; and determining the at least a medical feature as a function of the at least a query image using the trained image processing machine learning model and a second subset of the plurality of image processing algorithms; and suggesting, by the processor using a display device, a modification to the concurrent medical procedure as a function of the at least a medical feature. 11. The method of claim 10, wherein: the intervention system comprises a catheter configured to perform an ablation procedure; determining the at least a medical feature comprises identifying a thrombus in a left atrial appendage of a heart; and suggesting the modification to the concurrent medical procedure comprises terminating the ablation procedure upon identifying the thrombus. 12. The method of claim 11, wherein determining the at least a medical feature comprises identifying a case of calcification in a heart valve. 13. The method of claim 12, wherein suggesting the modification to the concurrent medical procedure comprises avoiding, using the catheter, the heart valve upon identifying the case of calcification. 14. The method of claim 10, wherein: the plurality of labelled features comprises a plurality of heart valve functions; and determining the at least a medical feature comprises: classifying, using a classifier, the feature training data into a plurality of classes as a function of the plurality of heart valve functions, wherein each class of the plurality of classes shares at least one similar heart valve function; and classifying an abnormal function of a heart valve as a function of the plurality of classes. 15. The method of claim 10, wherein: determining the at least a medical feature comprises determining a thickness pertaining to the anatomical structure; the plurality of labelled features comprises a plurality of thicknesses; and determining the at least a medical feature comprises: labeling, using a labeling algorithm, the feature training data with a plurality of labels correlated with the plurality of thicknesses, wherein each label of the plurality of labels represents a thickness; and determining the thickness as a function of the plurality of labels. 16. The method of claim 15, wherein suggesting the modification to the concurrent medical procedure comprises adjusting at least a parameter of an ablation procedure as a function of the thickness. 17. The method of claim 10, wherein the at least a query image comprises an intracardiac echocardiogram (ICE). 18. The method of claim 10, wherein the at least a query image comprises a point-of-care ultrasound (POCUS). 19. An apparatus for identifying medical features, the apparatus comprising: an ultrasonic imaging system configured to detect at least a query image comprising ultrasound data pertaining to a subject; an intervention system configured to perform a concurrent medical procedure as a function of the detection of the at least a query image; a display device; a processor; and a memory communicatively connected to the processor, wherein the memory contains instructions configuring the processor to: receive, from the ultrasonic imaging system, the at least a query image; identify a pose of an anatomical structure as a function of the at least a query image; determine at least a medical feature as a function of the pose of the anatomical structure, wherein determining the at least a medical feature comprises: receiving feature training data comprising a plurality of training images as input and a plurality of labelled features as output, wherein the plurality of labelled features comprises a plurality of heart valve functions; classifying, using a classifier, the feature training data into a plurality of classes as a function of the plurality of heart valve functions, wherein each class of the plurality of classes shares at least one similar heart valve function; classifying an abnormal function of a heart valve as a function of the plurality of classes; training an image processing machine learning model comprising a plurality of image processing algorithms using the feature training data; and determining the at least a medical feature as a function of the at least a query image using the trained image processing machine learning model; and suggest, using the display device, a modification to the concurrent medical procedure as a function of the at least a medical feature. 20. A method for identifying medical features, the method comprising: detecting, using an ultrasonic imaging system, at least a query image comprising ultrasound data pertaining to a subject; performing, using an intervention system, a concurrent medical procedure as a function of the detection of the at least a query image; receiving, by a processor from the ultrasonic imaging system, the at least a query image; identifying, by the processor, a pose of an anatomical structure as a function of the at least a query image; determining, by the processor, at least a medical feature as a function of the pose of an anatomical structure, wherein determining the at least a medical feature comprises: receiving feature training data comprising a plurality of training images as input and a plurality of labelled features as output wherein the plurality of labelled features comprises a plurality of heart valve functions; classifying, using a classifier, the feature training data into a plurality of classes as a function of the plurality of heart valve functions, wherein each class of the plurality of classes shares at least one similar heart valve function; classifying an abnormal function of a heart valve as a function of the plurality of classes; training an image processing machine learning model comprising a plurality of image processing algorithms using the feature training data; and determining the at least a medical feature as a function of the at least a query image using the trained image processing machine learning model; and suggesting, by the processor using a display device, a modification to the concurrent medical procedure as a function of the at least a medical feature. | 2024 | ANUMANA | AI-driven healthcare company | Start Up | https://anumana.ai/ | USA | The patent focuses on an apparatus that combines an ultrasound imaging system, an intervention system, and processor-driven feature identification to perform medical procedures based on detected features. It addresses the challenge that ultrasound features must be identified accurately before intervention can be safely performed. Feature detection in real time guides the intervention to the right location. | הפטנט מתמקד במכשיר המשלב מערכת הדמיית אולטרסאונד, מערכת התערבות וזיהוי תכונות מונע מעבד לביצוע פרוצדורות רפואיות על בסיס תכונות מזוהות. הוא מתמודד עם האתגר שיש לזהות תכונות אולטרסאונד במדויק לפני שניתן לבצע התערבות בבטחה. זיהוי תכונות בזמן אמת מנחה את ההתערבות למיקום הנכון. |  |  | 1 | [19 empty] | 1 |  | 1 | [7 empty] | ANUMANA INC | ANUMANA INC | ANUMANA, INC. | ANUMANA INC | ANUMANA INC | 2024-07-29 | 2024 | 2025-06-17 | 2025 | 2024-07-29 | 2024 | Grant | Utility Invention | active | US202418787453 | 325966200 | US20260026897A1¦WO2026030288A1 | US
426 | EP4635440A1 | LASER ABLATION METHOD, APPARATUS AND DEVICE, AND COMPUTER READABLE STORAGE MEDIUM | The present disclosure provides a laser ablation method, a laser ablation apparatus, a laser ablation device, a computer-readable storage medium, and a computer program product. The laser ablation method includes: determining a corresponding laser ablation mode based on real-time lesion region information at a first time; obtaining real-time lesion region information at a (i+1)th time; determining whether a conformity index between the real-time lesion region information at the (i+1)th time and expected lesion region information after an i-th laser ablation exceeds a first set threshold, to determine a (i+1)th laser ablation mode or end the laser ablation. In the present application, it is able to provide an appropriate ablation strategy according to an actual condition of a lesion and complete the ablation strategy in a single ablation operation, with relatively low risk. | 1. A laser ablation method, comprising: determining a corresponding laser ablation mode based on real-time lesion region information at a first time; obtaining real-time lesion region information at a (i+1)th time, wherein the (i+1)th time is later than an i-th time, and i is a positive integer; determining whether a conformity index between the real-time lesion region information at the (i+1)th time and expected lesion region information after an i-th laser ablation exceeds a first set threshold; determining, in response to that the conformity index between the real-time lesion region information at the (i+1)th time and the expected lesion region information after the i-th laser ablation exceeds the first set threshold, a (i+1)th laser ablation mode according to an expected laser ablation mode of a (i+1)th laser ablation in a laser ablation strategy and the real-time lesion region information at the (i+1)th time, and continuing to determine a next laser ablation mode through repeating the above steps of obtaining and subsequent determining; ending the laser ablation in response to that the conformity index between the real-time lesion region information at the (i+1)th time and the expected lesion region information after the i-th laser ablation does not exceed the first set threshold. 2. The laser ablation method according to claim 1, wherein, prior to determining the corresponding laser ablation mode based on the real-time lesion region information at the first time, the method further comprises:obtaining initial lesion region information of a user and a corresponding laser ablation strategy, wherein the laser ablation strategy comprises a quantity of laser ablation, an expected laser ablation mode corresponding to each laser ablation, and expected lesion region information after each laser ablation. 3. The laser ablation method according to claim 2, wherein, prior to obtaining the initial lesion region information of the user and the corresponding laser ablation strategy, the method further comprises: receiving digital image information of the user; performing multimodal three-dimensional modeling based on the digital image information to obtain a three-dimensional model of the user; receiving a marking operation for a position of a tissue in the three-dimensional model to determine the initial lesion region information of the user; wherein the tissue comprises a lesion, nerve fiber tracts, a brain partition and a blood vessel; receiving expected lesion region information after a target ablation of a lesion corresponding to the initial lesion region information; and determining the laser ablation strategy based on the initial lesion region information and the expected lesion region information. 4. The laser ablation method according to claim 3, further comprising: receiving the marking operation for the position of the tissue in the three-dimensional model, determining whether there is a non-ablation region, and marking the non-ablation region in response to that there is a non-ablation region. 5. The laser ablation method according to claim 3, wherein determining the corresponding laser ablation mode based on the real-time lesion region information at the first time comprises: obtaining the real-time lesion region information at the first time; and determining an expected laser ablation mode of a first laser ablation in the laser ablation strategy as the laser ablation mode in response to that a conformity index between the real-time lesion region information at the first time and the initial lesion region information is less than a second set threshold. 6. The laser ablation method according to claim 1, wherein determining the (i+1)th laser ablation mode according to the expected laser ablation mode of the (i+1)th laser ablation in the laser ablation strategy and the real-time lesion region information at the (i+1)th time comprises: adjusting, in response to that a conformity index between the expected lesion region information at the (i+1)th time in the laser ablation strategy and the real-time lesion region information at the (i+1)th time is less than a third set threshold, the expected laser ablation mode of the (i+1)th laser ablation according to the real-time lesion region information at the (i+1)th time, to obtain the corresponding laser ablation mode; maintaining, in response to that the conformity index between the expected lesion region information at the (i+1)th time and the real-time lesion region information at the (i+1)th time is greater than or equal to the third set threshold, a laser ablation mode at the i-th time as the laser ablation mode at the (i+1)th time. 7. The laser ablation method according to claim 2, wherein the real-time lesion region information comprises real-time temperature information of a lesion region; and the laser ablation method further comprises: controlling, in response to that the real-time temperature information is greater than a preset temperature threshold, an optical fiber assembly for outputting the laser to be cooled until subsequent real-time temperature information is not greater than the preset temperature threshold; wherein the laser ablation method further comprises:obtaining real-time temperature information, and obtaining a three-dimensional temperature distribution map of the real-time temperature information based on the three-dimensional model of the user. 8. The laser ablation method according to claim 10, wherein obtaining the real-time temperature information comprises: obtaining the real-time temperature information through parallel imaging, wherein a scanning slice thickness is not less than 3 mm and not greater than 5 mm; or, obtaining the real-time temperature information through three-dimensional scanning; or, obtaining the real-time temperature information through thin slice scanning, wherein a scanning slice thickness is less than 3 mm. 9. The laser ablation method according to claim 1, wherein the real-time lesion region information comprises: a plurality of pieces of target region information, and the method further comprises:determining the laser ablation mode corresponding to each target region information. 10. The laser ablation method according to claim 1, wherein the real-time lesion region information comprises boundary information of a lesion region, and the method further comprises:determining in real time whether the boundary information exceeds a preset boundary threshold. 11. The laser ablation method according to claim 1, further comprising:displaying, in real time, the real-time lesion region information at each time and the conformity index between the real-time lesion region information at each time and the expected lesion region information through a three-dimensional model. 12. The laser ablation method according to claim 1, further comprising:color-marking the real-time lesion region information in a three-dimensional model, wherein the real-time lesion region information corresponding to each time has a different marking color from the expected lesion region information. 13. A laser ablation apparatus, comprising: a first processing module, configured to determine a corresponding laser ablation mode based on real-time lesion region information at a first time; a second obtaining module, configured to obtain real-time lesion region information at a (i+1)th time, wherein the (i+1)th time is later than an i-th time, and i is a positive integer; wherein a lesion region at the (i+1)th time is ablated based on a laser ablation mode at the i-th time; and a second processing module, configured to determine a (i+1)th laser ablation mode according to a laser ablation strategy and the real-time lesion region information at the (i+1)th time, until a conformity index between real-time lesion region information at a subsequent time and expected lesion region information after a final laser ablation exceeds a first set threshold. 14. A laser ablation device, comprising a control host and a computer program stored on the control host, wherein the control host is configured to execute the computer program to implement the steps of the laser ablation method according to any one of claims 1 to 12. 15. A computer program product, comprising a computer program, wherein the computer program implements, when executed by a control host, the steps of the laser ablation method according to any one of claims 1 to 12. | 2024 | HANGZHOU GENLIGHT MEDTECH | Medical and Healthcare | NA | http://en.genlight.com/ | China | The patent focuses on a laser ablation method that determines a corresponding laser ablation mode from real-time lesion-region information at successive time points. It addresses the limitation of fixed-mode laser ablation that doesn't adapt to evolving lesion conditions during a procedure. The system updates the ablation mode based on the live lesion state for better outcomes. | הפטנט מתמקד בשיטת אבלציית לייזר הקובעת מצב אבלציית לייזר מתאים מתוך מידע אזור נגע בזמן אמת בנקודות זמן עוקבות. הוא מתמודד עם המגבלה של אבלציית לייזר במצב קבוע שאינה מסתגלת לתנאי נגע מתפתחים במהלך פרוצדורה. המערכת מעדכנת את מצב האבלציה על בסיס מצב הנגע החי לתוצאות טובות יותר. |  |  |  |  | 1 |  |  |  | 1 |  |  |  |  | 1 | [16 empty] | 1 |  | HANGZHOU GENLIGHT MEDTECH CO LTD | HANGZHOU GENLIGHT MEDTECH CO LTD | HANGZHOU GENLIGHT MEDTECH CO LTD | HANGZHOU GENLIGHT MEDTECH CO LTD | HANGZHOU GENLIGHT MEDTECH CO LTD | 2024-12-26 | 2024 | 2025-10-22 | 2025 | 2024-04-16 | 2024 | Application | Utility Invention | active | EP20240223313 | 323766322 | CN118177963A¦US20250318875A1 | EP
427 | EP4570205A1 | ASSISTANCE DEVICE, ASSISTANCE METHOD, AND RECORDING MEDIUM | An assistance device includes: a processor, the processor being configured to acquire electrical information related to an energy signal for driving an energy device configured to apply energy to a site to perform treatment on the site, input the electrical information to a trained model to cause the trained model to estimate intraoperative information that is related to the site indicated during an operation, determine whether or not accuracy of the intraoperative information is equal to or greater than a fixed value, and perform control according to the intraoperative information when it is determined that the accuracy is equal to or greater than the fixed value. | 1. An assistance device comprising:a processor, the processor being configured to acquire electrical information related to an energy signal for driving an energy device configured to apply energy to a site to perform treatment on the site, the site being a target for treatment in living tissue, input the electrical information to a trained model to cause the trained model to estimate intraoperative information that is related to the site indicated during an operation, the trained model having been obtained by training a plurality of pieces of training data in each of which at least one of driving information obtained when the energy device is driven, device information indicating a state of the energy device, and site information related to a size, a type, and a state of the site is used as an input parameter and is associated with the intraoperative information that is used as an output parameter, determine whether or not accuracy of the intraoperative information is equal to or greater than a fixed value, and perform control according to the intraoperative information when it is determined that the accuracy is equal to or greater than the fixed value. 2. The assistance device according to claim 1, wherein the processor is configured to calculate the accuracy of the intraoperative information, determine whether or not the accuracy is equal to or greater than the fixed value, and perform control according to the intraoperative information when it is determined that the accuracy is equal to or greater than the fixed value. 3. The assistance device according to claim 2, wherein the processor is configured to calculate the accuracy based on a plurality of pieces of intraoperative information. 4. The assistance device according to claim 1, wherein the processor is configured to output the intraoperative information externally when it is determined that the accuracy is equal to or greater than the fixed value. 5. The assistance device according to claim 1, wherein the processor is configured to change a control parameter for driving the energy device based on the intraoperative information when it is determined that the accuracy is equal to or greater than the fixed value, and maintain the control parameter for driving the energy device when it is determined that the accuracy is not equal to or greater than the fixed value. 6. The assistance device according to claim 1, wherein the electrical information is time series data related to at least one of a US signal and a HF signal for driving the energy device. 7. The assistance device according to claim 1, wherein the intraoperative information is at least one of the site information and a control parameter for controlling the energy device used when the energy device applies the energy to the site. 8. The assistance device according to claim 1, wherein the processor is configured to input the electrical information to each of a plurality of trained models to cause each of the plurality of trained models to estimate a plurality of respective pieces of intraoperative information, each of the plurality of trained model having been obtained by training the plurality of respective pieces of training data that are obtained in different situations, calculate the accuracy of each of the pieces of intraoperative information based on the plurality of respective pieces of intraoperative information, determine whether or not the accuracy is equal to or greater than the fixed value, and perform the control according to the intraoperative information when it is determined that the accuracy is equal to or greater than the fixed value. 9. The assistance device according to claim 1, wherein the processor is configured to input the electrical information to an authentication trained model to cause the authentication trained model to estimate accuracy of the electrical information as an estimation result, the authentication trained model having been obtained by training a plurality of pieces of authentication training data in each of which at least one of the driving information obtained when the energy device is driven, the device information indicating the state of the energy device, and the site information related to the size, the type, and the state of the site is used as an input parameter and is associated with accuracy of deviation, which is used as an output parameter, between the electrical information and the plurality of pieces of training data, determine whether or not the accuracy of the electrical information is equal to or greater than the fixed value, and input the electrical information to the trained model to cause the trained model to estimate the intraoperative information when the accuracy of the electrical information is equal to or greater than the fixed value. 10. The assistance device according to claim 1, wherein the processor is configured to input the electrical information to a determination trained model to cause the determination trained model to estimate association information indicating availability to the trained model, the determination trained model having been obtained by training a plurality of pieces of determination training data in each of which at least one of the driving information obtained when the energy device is driven, the device information indicating the state of the energy device, the site information related to the size, the type, and the state of the site, and the association information is used as an input parameter and is associated with the association information that is used as an output parameter, determine whether or not the accuracy of the association information is equal to or greater than the fixed value, and input the electrical information to the trained model to cause the trained model to estimate the intraoperative information when it is determined that the accuracy of the association information is equal to or greater than the fixed value. 11. The assistance device according to claim 1, wherein the processor is configured to input the electrical information to a situation identification trained model to causes the situation identification trained model to estimate situation information that is related to a situation of the site, the situation identification trained model having been obtained by training a plurality of pieces of situation training data in each of which at least one of the driving information obtained when the energy device is driven, the device information indicating the state of the energy device, and the site information related to the size, the type, and the state of the site is used as an input parameter and is associated with the situation information that is used as an output parameter, select an optimum trained model from among a plurality of trained models that have been obtained by training the plurality of pieces of training data obtained in different situations based on the situation information, and input the electrical information to the optimum trained model to cause the optimum trained model to estimate the intraoperative information. 12. The assistance device according to claim 11, wherein the processor is configured to determine whether or not accuracy of the situation information is equal to or greater than the fixed value, and input the electrical information to the optimum trained model to cause the optimum trained model to estimate the intraoperative information when it is determined that the accuracy of the situation information is equal to or greater than the fixed value. 13. The assistance device according to claim 11, wherein the processor is configured to input the electrical information to the optimum trained model to cause the optimum trained model to estimate the intraoperative information when it is determined that accuracy of the situation information is equal to or greater than the fixed value, and maintain a control parameter for driving the energy device when it is determined that the accuracy of the situation information is not equal to or greater than the fixed value. 14. The assistance device according to claim 1, wherein the processor is configured to input the electrical information to a grasp determination trained model to cause the grasp determination trained model to estimate grasping information related to a grasping state of the site grasped by the energy device, the grasp determination trained model having been obtained by training a plurality of pieces of grasp training data in each of which the grasping information and at least one of the driving information obtained when the energy device is driven, the device information indicating the state of the energy device, and the site information related to the size, the type, and the state of the site are used as input parameters and are associated with the grasping information that is used as an output parameter, and further input the grasping information to the trained model to cause the trained model to estimate the intraoperative information. 15. The assistance device according to claim 14, wherein the processor is configured to determine whether or not accuracy of the grasping information is equal to or greater than the fixed value, and further input the grasping information to the trained model to cause the trained model to estimate the intraoperative information when it is determined that the accuracy of the grasping information is equal to or greater than the fixed value. 16. The assistance device according to claim 1, wherein the processor is considered to decide, based on the electrical information, a control parameter for controlling the energy device, to output the decided control parameter, acquire the electrical information after having output the control parameter to the energy device, and input the electrical information to the trained model to cause the trained model to estimate the intraoperative information. 17. The assistance device according to claim 16, wherein the processor is configured to input the electrical information to a quadratic discriminant trained model to cause the quadratic discriminant trained model to estimate quadratic discriminant information indicating whether or not quadratic discriminant is available by using the trained model, the quadratic discriminant trained model having been obtained by training a plurality of pieces of quadratic discriminant training data in each of which the quadratic discriminant information and at least one of the driving information obtained when the energy device is driven, the device information indicating the state of the energy device, and the site information related to the size, the type, and the state of the site are used as input parameters and are associated with the quadratic discriminant information that is used as an output parameter, and input the electrical information and the quadratic discriminant information to the trained model to cause the trained model to estimate the intraoperative information. 18. The assistance device according to claim 17, wherein the quadratic discriminant information indicates either a pressure resistance value of a blood vessel at the site or a state of the treatment performed on the site by the energy device. 19. An assistance method implemented by an assistance device that comprises a processor, the processor being configured to acquire electrical information related to an energy signal for driving an energy device configured to apply energy to a site to perform treatment on the site, the site being a target for treatment in living tissue, input the electrical information to a trained model to cause the trained model to estimate intraoperative information that is related to the site indicated during an operation, the trained model having been obtained by training a plurality of pieces of training data in each of which at least one of driving information obtained when the energy device is driven, device information indicating a state of the energy device, and site information related to a size, a type, and a state of the site is used as an input parameter and is associated with the intraoperative information that is used as an output parameter, determine whether or not accuracy of the intraoperative information is equal to or greater than a fixed value, and perform control according to the intraoperative information when it is determined that the accuracy is equal to or greater than the fixed value. 20. A non-transitory computer-readable recording medium with an executable program stored thereon, the program causing a computer to execute: acquiring electrical information related to an energy signal for driving an energy device configured to apply energy to a site to perform treatment on the site, the site being a target for treatment in living tissue; inputting the electrical information to a trained model to cause the trained model to estimate intraoperative information that is related to the site indicated during an operation, the trained model having been obtained by training a plurality of pieces of training data in each of which at least one of driving information obtained when the energy device is driven, device information indicating a state of the energy device, and site information related to a size, a type, and a state of the site is used as an input parameter and is associated with the intraoperative information that is used as an output parameter; determining whether or not accuracy of the intraoperative information is equal to or greater than a fixed value; and performing control according to the intraoperative information when it is determined that the accuracy is equal to or greater than the fixed value. | 2023 | Olympus | Medical and Healthcare | Large-Size Company | https://www.olympus-global.com | Japan | The patent focuses on an assistance device whose processor takes electrical info from an energy device and feeds it to a trained model that estimates intraoperative info about the treatment site. It addresses the gap that conventional energy devices don't translate raw electrical signals into clinically meaningful intraoperative information. The trained model gives operators useful feedback during treatment. | הפטנט מתמקד במכשיר עזר שהמעבד שלו לוקח מידע חשמלי ממכשיר אנרגיה ומזין אותו למודל מאומן המעריך מידע תוך-ניתוחי על אתר הטיפול. הוא מתמודד עם הפער שמכשירי אנרגיה קונבנציונליים אינם מתרגמים אותות חשמליים גולמיים למידע תוך-ניתוחי משמעותי קלינית. המודל המאומן נותן למפעילים היזון חוזר שימושי במהלך הטיפול. | 1 | [21 empty] | 1 | [7 empty] | 1 |  | OLYMPUS MEDICAL SYSTEMS CORP | Olympus | OLYMPUS MEDICAL SYSTEMS CORP | Olympus | Olympus | 2024-11-25 | 2024 | 2025-06-18 | 2025 | 2023-11-28 | 2023 | Application | Utility Invention | active | EP20240215116 | 321463426 | US20250169876A1¦WO2025115599A1 | EP
428 | US20240156526A1 | NEGATIVE PRESSURE BASED IMAGING AND THERAPEUTIC APPARATUS AND SYSTEM FOR WELL-CONFINED ABNORMAL MUCOSAL TISSUE ABLATION AND WORKING METHOD OF THE SYSTEM | The invention relates to a method, system, and apparatus used in endoscopic medical diagnostic applications and endoscopic medical treatment applications, particularly in applications for imaging and treating lesions occurring in the biological mucosal layer with the thermal ablation (coagulation) method. | 1. A thermal therapy system operating between 0° to 360° to be used in endoscopic medical therapy applications and especially photothermal (laser) ablation/coagulation therapy applications, characterized in that it comprises: an electric vacuum pump 70 that creates a negative pressure necessary to pull the target tissue into the recessed channel of the therapeutic apparatus, a manometer 71 that is connected in series to the electric vacuum pump 70 and continuously measures the negative pressure value, a light source 72 that exposes the target tissue pulled into the recessed channel of the therapeutic apparatus to the photothermal ablation/coagulation, a 360-degree therapeutic apparatus 10 in which an optical waveguide array surrounds the recessed channel, or a 90-degree therapeutic apparatus using optical waveguides 30, or a 90-degree therapeutic apparatus using an array of optical waveguides coupled with polished ball-lensed tips 50, or a 360-degree therapeutic apparatus with a centered conical mirror 80, or a 0 to 360-degree active angle control therapeutic apparatus 110, with at least one of them being connected to the electric vacuum pump 70, manometer 71 and light source 72. 2. The thermal therapy system according to claim 1, characterized in that it comprises an electric vacuum pump 70 which provides vacuum and creates a negative pressure in the range of −150 mmHg to −760 mmHg. 3. The thermal therapy system according to claim 1, characterized in that it comprises a light source 72 that radiates in the near-infrared wavelength for the photothermal ablation/coagulation method. 4. The thermal therapy system according to claim 3, characterized in that the light source 72 that radiates in the near-infrared wavelength for the photothermal ablation/coagulation method is an amplified spontaneous laser source, or a semiconductor laser source, or a fiber laser source, or a solid-state laser source. 5. The thermal therapy system according to claim 3, characterized in that the laser beam from the light source 72 that radiates in the near-infrared wavelength for the photothermal ablation/coagulation method has an optical waveguide to be easily transmitted to the 360-degree therapeutic apparatus 10 in which an optical waveguide array surrounds the recessed channel, or the 90-degree treatment apparatus using optical waveguides 30, or a 90-degree treatment apparatus using an array of optical waveguides coupled with polished ball-lensed tips 50, or a 360-degree treatment apparatus with a centered conical mirror 80, or a 0 to 360-degree active angle control therapeutic apparatus 110. 6. The thermal therapy system according to claim 5, characterized in that the optical waveguide transmitting the laser beam is a single-mode optical fiber, or a multi-mode optical fiber. 7. The thermal therapy according to claim 1, characterized in that the 360-degree therapeutic apparatus 10 in which an optical waveguide array surrounds the recessed channel, or the 90-degree therapeutic apparatus using optical waveguides 30, or a 90-degree therapeutic apparatus using an array of optical waveguides coupled with polished ball-lensed tips 50, or a 360-degree therapeutic apparatus with a centered conical mirror 80, or a 0 to 360-degree active angle control therapeutic apparatus 110 has an outer diameter in the range of 10 mm to 20 mm for appropriate intervention in the human gastrointestinal tract. 8. The thermal therapy system according to claim 1, characterized in that the 360-degree therapeutic apparatus 10 in which an optical waveguide array surrounds the recessed channel located therein comprises: a top cover 11 that forms the top of the 360-degree therapeutic apparatus 10, a lower cover 13 that forms the bottom of the 360-degree therapeutic apparatus 10, a body with a recessed channel 12 that overlaps with the top cover 11 and lower cover 13, forms the whole of the 360-degree therapeutic apparatus 10, and has a circumferential row of holes for the negative pressure application, an optical waveguide bundle that is located so as to have a circular array on the bottom of the recessed channel and allows the therapeutic laser beam to be transmitted to the 360-degree therapeutic apparatus 10. 9. The thermal therapy system according to claim 1, characterized in that the 90-degree therapeutic apparatus using optical waveguides 30 located therein comprises: an upper body with a 90° opening 31 for the negative pressure applications, which forms the top of the 90-degree therapeutic apparatus using optical waveguides 30, a lower cover 13 that forms the bottom of the 90-degree treatment apparatus using optical waveguides (30) and forms the whole of the 90-degree therapeutic apparatus using optical waveguides 30 by the overlap of the upper body with a 90° opening 31, an optic waveguide bundle that is located on the bottom of the recessed channel in an arc shape of a 90° angle and allows the therapeutic laser beam to be transmitted to the 90-degree therapeutic apparatus using optical waveguides 30. 10. The thermal therapy system according to claim 8 or claim 9, characterized in that the optical waveguides located therein are a flat-ended optical fiber 14 or a tapered optical fiber 15. 11. An operation method of the thermal therapy system according to claim 8 or claim 9, characterized in that it comprises the following process steps: pulling the target mucosal tissue with negative pressure into the body with a recessed channel 12 of the 360-degree therapeutic apparatus 10 in which an optical waveguide array surrounds the recessed channel, or into the opening on the upper body with a 90-degree opening 31 of the 90-degree therapeutic apparatus using optical waveguides 30, transmitting the laser beam with optical waveguides located on the body to the 360-degree therapeutic apparatus 10 or 90-degree therapeutic apparatus using optical waveguides 30. 12. The thermal therapy system according to claim 1, characterized in that the 90-degree therapeutic apparatus using an array of optical waveguides coupled with polished ball-lensed tips 50 located therein comprises: a d-shaped cylinder to fix the optical waveguides in place 51, a glass ferrule 52 located on the d-shaped cylinder to fix the optical waveguides in place 51 as a protective, a glass top cover with grooves 53 that fixes the d-shaped cylinder to fix the optical waveguides in place 51 and glass ferrule 52 and forms the top of the 90-degree therapeutic apparatus using an array of optical waveguides coupled with polished ball-lensed tips 50, a lower body with a 90-degree opening 54 that forms the 90-degree therapeutic apparatus using an array of optical waveguides coupled with polished ball-lensed tips 50 by the overlap of the framework composed of a glass top cover with grooves 53, fixing the d-shaped cylinder to fix the optical waveguides in place 51 and glass ferrule 52 in place, optical fibers coupled with polished ball-lensed tips 55 that are located in the d-shaped cylinder to fix the optical waveguides in place 51 so as to be two layers along the arc of 90° and transmits the therapeutic laser beam to the 90-degree therapeutic apparatus using an array of optical waveguides coupled with polished ball-lensed tips 50. 13. The thermal therapy system according to claim 12, characterized in that the optical waveguide coupled with ball lens in the 90-degree therapeutic apparatus design using an array of optical waveguides coupled with polished ball-lensed tips 50 is angled-polished. 14. An operation method of the thermal therapy system according to claim 12, characterized in that it comprises the following process steps: pulling the target mucosal tissue with negative pressure to the opening on the lower body with a 90-degree opening 54, transmitting the laser beam to the 90-degree therapeutic apparatus using an array of optical waveguides coupled with polished ball-lensed tips 50 by optical fibers coupled with polished ball-lensed tips 55 arranged in an arc to form a double-row in the d-shaped cylinder to fix the optical waveguides in place 51 within the glass ferrule 52 that is placed in the center of the glass top cover with grooves 53, reflecting the laser beams directed by the ball lenses the glass ferrule 52 to the opening of the 90-degree therapeutic apparatus using an array of optical waveguides coupled with polished ball-lensed tips 50. 15. The thermal therapy system according to claim 1, characterized in that the 360-degree therapeutic apparatus with a centered conical mirror 80 located therein comprises: a conical mirror 85 that is fixed in place in the center of the top cover and reflects the laser beam transmitted to the 360-degree therapeutic apparatus with a centered conical mirror 80, a top cover for 360-degree reflection 81 that forms the top of the 360-degree therapeutic apparatus with a centered conical mirror 50, a body with a recessed channel 12 that comprises a circumferential row of holes for negative pressure and forms the whole of the 360-degree therapeutic apparatus with a centered conical mirror 80 by the overlap of the top and bottom parts, a lower-right cover for a single beam 83 that forms the bottom of the 360-degree therapeutic apparatus with a centered conical mirror 80 together with the lower-left cover for a single beam 82, an optical waveguide coupled collimator (84) that is fixed in place in the center of the bottom part formed by the overlap of the lower-right cover for a single beam 83 and lower-left cover for a single beam 82 and collimates the therapeutic laser beam to be transmitted to the 360-degree therapeutic apparatus with a centered conical mirror 80. 16. An operation method of the thermal therapy system according to claim 15, characterized in that it comprises the following process steps: pulling the target mucosal tissue with negative pressure to the body with a recessed channel 12, using the optical waveguide coupled collimator 84 to transmit the laser beam to the 360-degree therapeutic apparatus with a centered conical mirror 80, circumferentially reflecting the collimated laser beam by the conical mirror 85 fixed in place thereof on the top cover of the 360-degree therapeutic apparatus with a centered conical mirror 80, positioning the circular plane formed by the reflected beam to the recessed channel of the 360-degree therapeutic apparatus with a centered conical mirror 80. 17. The thermal therapy system according to claim 1, characterized in that the 0 to 360-degree active angle control therapeutic apparatus 110 comprises: an upper left cover for single-beam reflection 111 that forms the top of the 0 to 360-degree active angle control therapeutic apparatus 110 together with the upper right cover for single-beam reflection 112, right-angle mirrors 115 that are fixed in place thereof on the top part formed by the overlap of the upper left cover for single-beam reflection 111 and the upper right cover for single-beam reflection 112 and reflect the laser beam, which is transmitted to the 0 to 360-degree active angle control therapeutic apparatus 110, vertically to the shaft rod end of the stepper motor, a body with a recessed channel 12 that that comprises a circumferential row of holes for negative pressure and forms the whole of the 0 to 360-degree active angle control therapeutic apparatus 110 by the overlap of the top and bottom parts, a lower left cover for single-beam transmission and scanning 113 that forms the bottom of the 0 to 360-degree active angle control therapeutic apparatus 110 together with the lower left cover for single-beam transmission and scanning 114, an optical waveguide coupled collimator 84 that is fixed in place in the bottom part formed by the overlap of the lower left cover for single-beam transmission and scanning 113 and the lower left cover for single-beam transmission and scanning 114 and collimates the therapeutic laser beam to be transmitted to the 0 to 360-degree active angle control therapeutic apparatus 110, a stepper motor 116 that is fixed in place in the bottom part formed by the overlap of the lower left cover for single-beam transmission and scanning 113 and the lower left cover for single-beam transmission and scanning 114 and driven by a driver board 130 for circumferential or local surface scanning of the laser beam. 18. The thermal therapy system according to claim 17, characterized in that it comprises a driver board 130 that increases the step resolution of the rotation motion of the stepper motor 116 to a maximum of 0.1 degrees for circumferential or local surface scanning of the laser beam. 19. The thermal therapy and imaging system according to claim 17, characterized in that it comprises a stepper motor 116 in which the rotational motion is actively controlled clockwise and counterclockwise by the driver board 130 for the stepper motor. 20. An operation method of the thermal therapy system according to claim 17, characterized in that it comprises the following process steps: pulling the target mucosal tissue with negative pressure to the body with a recessed channel 12, transmitting the laser beam to the 0 to 360-degree active angle control therapeutic apparatus 110 using the optical waveguide coupled collimator 84, reflecting the collimated laser beam to another right-angle mirror 115 located to the tip of the shaft rod of the stepper motor 116 with the right-angle mirrors 115 fixed in place on the top cover of the 0 to 360-degree active angle control therapeutic apparatus 110, performing the circumferential or local surface scanning of the laser beam by rotating the stepper motor 116. 21. The thermal therapy system according to claim 15 or claim 17, characterized in that it comprises an optical waveguide coupled collimator 84 that transmits the therapeutic laser beam to the 360-degree therapeutic apparatus with a centered conical mirror 80 or 0 to 360-degree active angle control therapeutic apparatus 110. 22. The thermal therapy system according to claim 21, characterized in that the optical waveguide coupled collimator 84 is a graded-index (GRIN)-lensed fiber collimator, or fiber-coupled aspherical lens collimator. 23. The thermal therapy system according to claim 15 or claim 17, characterized in that the conical mirror 85 that reflects the laser beam transmitted to the 360-degree therapeutic apparatus with a centered conical mirror 80, or right-angle mirrors 115 that reflects the laser beam, which is transmitted to the 0 to 360-degree active angle control therapeutic apparatus 110, vertically to the shaft rod end of the stepper motor has a protected gold or silver or aluminum coating. 24. The thermal therapy system according to claim 15 or claim 17, characterized in that the reflective surface of the conical mirror 85 that reflects the laser beam transmitted to the 360-degree therapeutic apparatus with a centered conical mirror 80, or right-angle mirror 115 that reflects the laser beam, which is transmitted to the 0 to 360-degree active angle control therapeutic apparatus 110, vertically to the shaft rod end of the stepper motor has a protective coating layer against scratching and oxidation. 25. The thermal therapy system according to claim 24, characterized in that the protective layer is silicon monoxide (SiO) or silicon dioxide (SiO2). 26. The thermal therapy system according to claim 8, claim 12, claim 15 or claim 17, characterized in that the parts constituting the casing of the 360-degree therapeutic apparatus 10, or the 90-degree therapeutic apparatus using optical waveguides 30, or 90-degree therapeutic apparatus using an array of optical waveguides coupled with polished ball-lensed tips 50, or 360-degree therapeutic apparatus with a centered conical mirror 80, or 0 to 360-degree active angle control therapeutic apparatus 110 are made of a biocompatible material. 27. The thermal therapy system according to claim 8, claim 12, claim 15 or claim 17, characterized in that the biocompatible material used in the 360-degree therapeutic apparatus 10, or the 90-degree therapeutic apparatus using optical waveguides 30, or 90-degree therapeutic apparatus using an array of optical waveguides coupled with polished ball-lensed tips 50, or 360-degree therapeutic apparatus with a centered conical mirror 80, or body with a recessed channel 12 of the 0 to 360-degree active angle control therapeutic apparatus 110, or upper body with a 90-degree opening 31, or lower body with a 90-degree opening 54 is transparent and in particularly has an optical transparency in the range of 400 nm to 2000 nm. 28. The thermal therapy system according to claim 26 or claim 27, characterized in that the biocompatible material further has a light and workable form which has a high surface hardness, low water absorption and dimensional stability, and is resistant to scratches, chemicals, heat, ultraviolet rays, and atmospheric conditions. 29. A 0 to 360-degree active angle controlled thermal therapy and imaging system to be used in endoscopic medical diagnostic applications, endoscopic medical therapy applications and therapy photothermal (laser) ablation/coagulation therapy applications, characterized in that it comprises: an electric vacuum pump 70 that creates a negative pressure necessary to pull the target tissue into the recessed channel of the therapeutic apparatus, a manometer 71 that is connected in series to the electric vacuum pump 70 and continuously measures the negative pressure value, a light source 72 that exposes the target tissue pulled into the recessed channel of the therapeutic apparatus to photothermal ablation/coagulation, an optical (400 nm-2000 nm) imaging device 160 that performs the imaging of the target tissue pulled into the apparatus with negative pressure, a 0 to 360-degree active angle control therapeutic and imaging apparatus 140, or MEMS mirror based 0 to 90-degree active angle control therapeutic and imaging apparatus 170, at least one of which is connected to an electric vacuum pump 70, manometer 71, light source 72 and an optical (400 nm-2000 nm) imaging device 160 to provide photothermal ablation therapy and imaging. 30. The thermal therapy and imaging system according to claim 29, characterized in that it comprises an electric vacuum pump 70 that provides vacuum and creates a negative pressure in the range of −150 mmHg to −760 mmHg. 31. The thermal therapy and imaging system according to claim 29, characterized in that it comprises a light source 72 that radiates in the near-infrared wavelength for the photothermal ablation/coagulation method or optical imaging. 32. The thermal therapy and imaging system according to claim 31, characterized in that the light source 72 that radiates in the near-infrared or visible wavelength for the photothermal ablation/coagulation method or optical imaging is an amplified spontaneous laser source, or a semiconductor laser source, or a fiber laser source, or a solid-state laser source. 33. The thermal therapy and imaging system according to claim 31, characterized in that it comprises an optical waveguide for easily transmitting the laser beam from the light source 72, which radiates in the near-infrared or visible wavelength, to the 0 to 360-degree active angle control therapeutic and imaging apparatus 140, or MEMS mirror based 0 to 90-degree active angle control therapeutic and imaging apparatus 170. 34. The thermal therapy and imaging system according to claim 33, characterized in that the optical waveguide transmitting the laser beam is a single-mode optical fiber, or a multi-mode optical fiber. 35. The thermal therapy and imaging system according to claim 29, characterized in that the diameter of the 0 to 360-degree active angle control therapeutic and imaging apparatus 140, or MEMS mirror based 0 to 90-degree active angle control therapeutic and imaging apparatus 170 is in the range of 10 mm to 20 mm for appropriate intervention in the human gastrointestinal tract. 36. The thermal therapy and imaging system according to claim 29, characterized in that the 0 to 360-degree active angle control therapeutic and imaging apparatus 140 comprises: an upper left cover for double-beam reflection 141 that forms the top of the 0 to 360-degree active angle control therapeutic and imaging apparatus 140 together with the upper right cover for double-beam reflection 142, right-angle mirrors 115 that is located on the top part formed by the overlap of the upper left cover for double-beam reflection 141 with the upper right cover for double-beam reflection 142 and transmits laser beams to the 0 to 360-degree active angle control therapeutic and imaging apparatus 140, a knife-edge prism mirror 146 that is located on the top part formed by the overlap of the upper left cover for double-beam reflection 141 with the upper right cover for double-beam reflection 142 and transmits laser beams vertically to the tip of the shaft rod of the motor of the 0 to 360-degree active angle control therapeutic and imaging apparatus 140, a body with a recessed channel 12 that forms the whole of the 0 to 360-degree active angle control therapeutic and imaging apparatus 140 by the overlap of the top and bottom parts, a lower left cover for double-beam transmission and scanning 143 that forms the bottom of the 0 to 360-degree active angle control therapeutic and imaging apparatus 140 together with the lower right cover for double-beam transmission and scanning 144, an optical waveguide coupled collimator 84 that is fixed in place in the bottom part formed by the overlap of the upper left cover for double-beam reflection 141 with the upper right cover for double-beam reflection 142 and collimates the therapeutic laser beam to be transmitted to the 0 to 360-degree active angle control therapeutic and imaging apparatus 140, an optical waveguide coupled focusing lens 145 that is fixed in place in the bottom part formed by the overlap of the upper left cover for double-beam reflection 141 with the upper right cover for double-beam reflection 142 and focuses the imaging laser beam to be transmitted to the 0 to 360-degree active angle control therapeutic and imaging apparatus 140, a stepper motor 116 that is fixed in place in the bottom part formed by the overlap of the upper left cover for double-beam reflection 141 with the upper right cover for double-beam reflection 142 and driven by a driver board 130 for circumferential or local surface scanning of the laser beam. 37. The thermal therapy and imaging system according to claim 36, characterized in that it comprises an optical waveguide coupled collimator 84 that transmits the therapeutic laser beam to the 0 to 360-degree active angle control therapeutic and imaging apparatus 140. 38. The thermal therapy and imaging system according to claim 37, characterized in that the optical waveguide coupled collimator 84 is a graded-index (GRIN)-lensed fiber collimator, or fiber-coupled aspherical lens collimator. 39. The thermal therapy and imaging system according to claim 36, characterized in that it comprises a driver board 130 that increases the step resolution of the rotational motion of the stepper motor 116 to a maximum of 0.1 degrees for circumferential or local surface scanning of the laser beam. 40. The thermal therapy and imaging system according to claim 36, characterized in that it comprises a stepper motor 116 in which the rotational motion is actively controlled clockwise and counterclockwise by the driver board 130 for the stepper motor. 41. An operation method of the thermal therapy and imaging system according to claim 36, characterized in that it comprises the following process steps: pulling the target mucosal tissue with negative pressure to the body with a recessed channel 12, transmitting the laser beam to the 0 to 360-degree active angle control therapeutic and imaging apparatus 140 using the optical waveguide coupled collimator 84, transmitting the imaging laser beam to the 0 to 360-degree active angle control therapeutic and imaging apparatus 140 using the optical waveguide coupled focusing lens 145, directing the collimated and focused laser beams to the knife-edge prism mirror 146 that is fixed in place in the top cover of the 0 to 360-degree active angle control therapeutic and imaging apparatus 140 with the right-angle mirrors 115 that are fixed in place in the top cover, reflecting the collimated and focused laser beams with to the knife-edge prism mirror 146 to another right-angle mirror 115 located to the tip of the shaft rod of the stepper motor 116, performing the circumferential or local surface scanning of the laser beams by rotating the stepper motor 116. 42. The thermal therapy and imaging system according to claim 29, characterized in that the MEMS mirror based 0 to 90-degree active angle control therapeutic and imaging apparatus 170 comprises: a back cover for single-beam reflection 171 that forms the back of the MEMS mirror based 0 to 90-degree active angle control therapeutic and imaging apparatus 170, a front cover with a 90-degree opening 172 that forms the front of the MEMS mirror based 0 to 90-degree active angle control therapeutic and imaging apparatus 170 and the whole of the MEMS mirror based 0 to 90-degree active angle control therapeutic and imaging apparatus 170 by overlap with the back cover for single-beam reflection 171, a MEMS mirror with a flexible printed circuit 173 that reflects the laser beam to the opening of the MEMS mirror based 0 to 90-degree active angle control therapeutic and imaging apparatus 170 for the surface scanning of the laser beam, an optical waveguide coupled focusing lens 145 that transmits the therapeutic and imaging laser beams and the same to the MEMS mirror of the MEMS mirror based 0 to 90-degree active angle control therapeutic and imaging apparatus 170. 43. The thermal therapy and imaging system according to claim 42, characterized in that it comprises a flexible printed circuit board 190 which controls the MEMS mirror with a flexible printed circuit 173 for the laser surface scanning. 44. An operation method of the treatment and imaging system according to claim 42, characterized in that it comprises the following process steps: pulling the target mucosal tissue with negative pressure into the opening on the front cover with a 90-degree opening 172, transmitting the laser beam to the MEMS mirror based 0 to 90-degree active angle control therapeutic and imaging apparatus 170 using the optical waveguide coupled focusing lens 145, reflecting the focused laser beam from the MEMS mirror with a flexible printed circuit 173 located above the optical focusing lens 145, in particularly to the opening of the MEMS mirror based 0 to 90-degree active angle control therapeutic and imaging apparatus 170, performing 0 to 90-degree surface scanning by moving the micro-electronic mechanical system (MEMS) mirror controlled by a flexible printed circuit at certain angles. 45. The thermal therapy and imaging system according to claim 36 or claim 42, characterized in that it comprises an optical waveguide coupled focusing lens 145 that transmits the focused laser beam to the 0 to 360-degree active angle control therapeutic and imaging apparatus 140 or MEMS mirror based 0 to 90-degree active angle control therapeutic and imaging apparatus 170. 46. The thermal therapy and imaging system according to claim 45, characterized in that the optical waveguide coupled focusing lens 145 is a graded-index (GRIN)-lensed fiber collimator, or fiber-coupled aspherical lens collimator. 47. The thermal therapy and imaging system according to claim 36 or claim 42, characterized in that the right-angle mirror 115 that reflects laser beams transmitted to the 0 to 360-degree active angle control therapeutic and imaging apparatus 140, or the knife-edge prism mirror 146 that reflects laser beams, which is transmitted to the 0 to 360-degree active angle control therapeutic and imaging apparatus 140, vertically to the shaft rod end of the stepper motor, or the MEMS mirror with a flexible printed circuit 173 that reflects the laser beam, transmitted to the MEMS mirror based 0 to 90-degree active angle control therapeutic and imaging apparatus 170, to the opening of the apparatus has a protected gold or silver or aluminum coating. 48. The thermal therapy and imaging system according to claim 36 or claim 42, characterized in that the reflective surface of the right-angle mirror 115 that reflects laser beams transmitted to the 0 to 360-degree active angle control therapeutic and imaging apparatus 140, or the knife-edge prism mirror 146 that reflects laser beams, which is transmitted to the 0 to 360-degree active angle control treatment and imaging apparatus 140, vertically to the shaft rod end of the stepper motor, or the MEMS mirror with a flexible printed circuit 173 that reflects the laser beam, which is transmitted to the MEMS mirror based 0 to 90-degree active angle control therapeutic and imaging apparatus 170, to the opening of the apparatus has a protective coating layer against scratching and oxidation. 49. The thermal therapy and imaging system according to claim 48, characterized in that the protective coating is silicon monoxide (SiO) or silicon dioxide (SiO2). 50. The thermal therapy system according to claim 36 or claim 42, characterized in that the parts constituting the casing of the 0 to 360-degree active angle control therapeutic and imaging apparatus 140 or the MEMS mirror based 0 to 90-degree active angle control therapeutic and imaging apparatus 170 are made of a biocompatible material. 51. The thermal therapy system according to claim 36 or claim 42, characterized in that the biocompatible material used in the 0 to 360-degree active angle control therapeutic and imaging apparatus 140 or the body with a recessed channel 12 of the MEMS mirror based 0 to 90-degree active angle control therapeutic and imaging apparatus 170, or the front cover with a 90-degree opening 172 is transparent and in particularly has an optical transparency in the range of 400 nm to 2000 nm. 52. The thermal therapy system according to claim 50 or claim 51, characterized in that the biocompatible material further a light and workable form which has a high surface hardness, low water absorption and dimensional stability, and is resistant to scratches, chemicals, heat, ultraviolet rays, and atmospheric conditions. 53. A 0 to 360-degree active angle controlled thermal therapy and imaging system to be used in endoscopic medical diagnostic applications, endoscopic medical therapy applications and radiofrequency ablation treatment applications, characterized in that it comprises: an electric vacuum pump 70 that creates a negative pressure necessary to pull the target tissue into the recessed channel of the therapeutic and imaging apparatus, a manometer 71 that is connected in series to the electric vacuum pump 70 and continuously measures the negative pressure value, a radiofrequency generator 250 that exposes the target tissue pulled into the recessed channel of the apparatus to radiofrequency ablation, an optical (400 nm-2000 nm) imaging device 160 that performs the imaging of the target tissue pulled into the apparatus with negative pressure, a radiofrequency-based 360-degree therapeutic and optical imaging apparatus 200, or a radiofrequency-based 90-degree therapeutic and optical imaging apparatus 220, or a 90-degree therapeutic and optical imaging apparatus using the radiofrequency ablation method and a MEMS mirror 260, at least one of which is connected to an electric vacuum pump 70, manometer 71, radiofrequency generator 250, and an optical (400 nm-2000 nm) imaging device 160 to provide radiofrequency ablation therapy and imaging. 54. The thermal therapy and imaging system according to claim 53, characterized in that it comprises an electric vacuum pump 70 which provides vacuum and creates a negati<TRUNCATED> | 2021 | IZMIR BIYOTIP VE GENOM MERKEZI¦DOKUZ EYLUL UNIVERSITESI | Research & development | NA | https://ibg.edu.tr/ | Turkey | The patent focuses on a negative-pressure-based imaging and therapeutic system for endoscopic ablation of well-confined abnormal mucosal tissue, particularly using thermal coagulation. It addresses the difficulty of precisely treating mucosal lesions where loose contact and shifting tissue limit accuracy. Negative pressure helps confine and image the target so coagulation is more accurate. | הפטנט מתמקד במערכת הדמיה וטיפול מבוססת לחץ שלילי לאבלציה אנדוסקופית של רקמה רירית חריגה תחומה היטב, במיוחד באמצעות קרישה תרמית. הוא מתמודד עם הקושי בטיפול מדויק בנגעי רירית כאשר מגע רופף ורקמה זזה מגבילים את הדיוק. לחץ שלילי עוזר לתחום ולהדמות את המטרה כך שהקרישה מדויקת יותר. |  | 1 | [8 empty] | 1 |  |  | 1 | [13 empty] | 1 |  |  |  |  | IZMIR BIYOTIP VE GENOM MERKEZI¦DOKUZ EYLUL UNIV | IZMIR BIYOTIP VE GENOM MERKEZI¦DOKUZ EYLUL UNIV | IZMIR BIYOTIP VE GENOM MERKEZI¦DOKUZ EYLUL UNIVERSITESI | IZMIR BIYOTIP VE GENOM MERKEZI¦DOKUZ EYLUL UNIVERSITESI | IZMIR BIYOTIP VE GENOM MERKEZI¦DOKUZ EYLUL UNIVERSITESI | 2022-02-11 | 2022 | 2024-05-16 | 2024 | 2021-02-11 | 2021 | Application | Utility Invention | active | US202218276874 | 304697976 | TR202101949A2¦WO2022173404A2¦WO2022173404A3 | US
429 | WO2025096240A1 | SPECTRAL COLLECTION DURING CONTINUOUS WAVE LASER EMISSION | A system for spectral collection during an in-vivo medical procedure can include a processor and memory with instructions stored thereon that cause the processor to emit an illumination signal from an illumination source toward an in-vivo target, collect a first return signal from the in-vivo target in response to the illumination signal, and emit a laser signal from a continuous wave laser source toward the in-vivo target. The instructions can further cause the processor to collect a second return signal from the in-vivo target, determine a difference between the first return signal and the second return signal, and take an action based at least in part on the determined difference. | 1. CLAIMS WHAT IS CLAIMED IS: 1. A system for spectral collection during an in-vivo medical procedure, the system comprising: a processor; and memory, including instructions stored thereon that, when performed by the processor, cause the processor to: emit an illumination signal from an illumination source toward an in-vivo target; collect a first return signal from the in-vivo target in response to the illumination signal; emit a laser signal from a continuous wave laser source toward the in-vivo target; collect a second return signal from the in-vivo target; determine a difference between the first return signal and the second return signal; and take an action based at least in part on the determined difference. 2. The system of claim 1, wherein to determine the difference between the first return signal and the second return signal, the instructions cause the processor to: compare a spectral analysis of the second return signal with a spectral analysis of the first return signal. 3. The system of claim 2, wherein spectra from at least one of the first return signal or the second return signal are inputted into an artificial intelligence algorithm or a machine-learning algorithm, wherein the artificial intelligence algorithm or the machine-learning algorithm identifies a cause of the difference and/or one or more characteristics of the in-vivo target and determines the action. Attorney Docket No.5409.890WO1 28 Client Reference No. GAP24024-URKT-WO14. The system of claim 1, wherein the instructions further cause the processor to: analyze imaging data captured concurrently with the first return signal and the second return signal; detect one or more changes in the imaging data indicative of a condition of the in-vivo target; based on the detected one or more changes in the imaging data, determine whether to continue collecting spectral data; and based on a determination to stop collecting spectral data, at least one of: stop collection of the spectral data or predict an optimal time to stop collection of the spectral data. 5. The system of claim 4, wherein the instructions further cause the processor to, based on the detected one or more changes in the imaging data, stop collection of the spectral data or predict an optimal time to stop collection of the spectral data. 6. The system of claim 1, wherein the action includes stopping emission of the continuous wave laser source. 7. The system of claim 1, wherein the action include adjusting an intensity level of the continuous wave laser source. 8. The system of claim 1, wherein the instructions further cause the processor to: determine whether the difference between the first return signal and the second return signal exceeds a predetermined threshold. Attorney Docket No.5409.890WO1 29 Client Reference No. GAP24024-URKT-WO19. The system of claim 8, wherein the action is taken in real-time or near real-time upon the processor determining that the difference between the first return signal and the second return signal exceeds the predetermined threshold. 10. The system of claim 1, wherein the instructions further cause the processor to: determine a cause resulting in the difference between the first return signal and the second return signal based on spectral analysis, video analysis or a combination thereof. 11. The system of claim 10, wherein the instructions further cause the processor to: confirm the cause based at least in part on data obtained from one or more sensors connected to or communicatively coupled to a medical scope in which at least one of the illumination source or the continuous wave laser source is included. 12. The system of claim 10, wherein the cause includes at least one of a flashing event, a change between a distance in a distal end of an endoscope and the in-vivo target, a change in the in-vivo target, or emission of the laser signal. 13. The system of claim 10, wherein the instructions further cause the processor to: identify a type of the in-vivo target based at least in part on the spectral analysis of at least one of the first return signal and the second return signal; and stop emission of the continuous wave laser source when a change in a characteristic of the in-vivo target occurs, wherein the change in the in-vivo target includes a change in a tissue type or tissue layer. Attorney Docket No.5409.890WO1 30 Client Reference No. GAP24024-URKT-WO114. A method for spectral collection during an in-vivo medical procedure, the method comprising: emitting an illumination signal from an illumination source toward an in- vivo target; collecting a first return signal from the in-vivo target in response to the illumination signal; emitting a laser signal from a continuous wave laser source toward the in-vivo target; collecting a second return signal from the in-vivo target; determining a difference between the first return signal and the second return signal; and taking an action based at least in part on the determined difference. 15. The method of claim 14, wherein determining the difference between the first return signal and the second return signal includes a comparison of a spectral analysis of the second return signal and a spectral analysis of the first return signal. 16. The method of claim 14, wherein spectra from at least one of the first return signal or the second return signal are inputted to an artificial intelligence algorithm or a machine-learning algorithm, wherein the artificial intelligence algorithm or the machine-learning algorithm identifies one or more characteristics of the in-vivo target and determines the action. 17. The method of claim 14, wherein at least one of the action is taken in real-time or near real-time upon determining that difference between the first return signal and the second return signal exceeds a predetermined threshold. 18. The method of claim 14, further comprising: determining a cause resulting in the difference between the first return signal and the second return signal based on spectral analysis, video analysis or a Attorney Docket No.5409.890WO1 31 Client Reference No. GAP24024-URKT-WO1 combination thereof, wherein the video analysis includes detection of a flashing event, a change in the in-vivo target, or an unexpected visual artifact; and confirming the cause based at least in part on data obtained from one or more sensors connected to or communicatively coupled to a medical scope in which at least one of the illumination source or the continuous wave laser source is included, wherein the cause includes at least one of a flashing event, a change between a distance in a distal end of an endoscope and the in-vivo target, or emission of the laser signal. 19. A system for spectral collection during an in-vivo medical procedure, comprising: an endoscope; a continuous wave laser source included in or connected to the endoscope; an illumination source included in or connected to the endoscope; and processing circuitry to: emit an illumination signal from an illumination source toward an in-vivo target; collect a first return signal from the in-vivo target in response to the illumination signal; emit a laser signal from a continuous wave laser source toward the in-vivo target; collect a second return signal from the in-vivo target; determine a difference between the first return signal and the second return signal; and take an action based at least in part on the determined difference. 20. The system of claim 19, wherein the processing circuitry is further to: compare a spectral analysis of the second return signal with a spectral analysis of the first return signal; analyze imaging data captured concurrently with the first return signal and the second return signal; Attorney Docket No.5409.890WO1 32 Client Reference No. GAP24024-URKT-WO1 detect one or more changes in the imaging data indicative of a condition of the in-vivo target; based on the detected one or more changes in the imaging data, determine whether to continue collecting spectral data; and based on a determination to stop collecting spectral data, at least one of: stop collection of the spectral data or predict an optimal time to stop collection of the spectral data. 21. The system of claim 19, wherein the processing circuitry is further to: determine a cause resulting in the difference between the first return signal and the second return signal based on spectral analysis, video analysis or a combination thereof, wherein the video analysis includes detection of a flashing event, a change in the in-vivo target, or an unexpected visual artifact; and confirm the cause based at least in part on data obtained from one or more sensors connected to or communicatively coupled to a medical scope in which at least one of the illumination source or the continuous wave laser source is included. Attorney Docket No.5409.890WO1 33 Client Reference No. GAP24024-URKT-WO1 | 2023 | Olympus¦GYRUS | Medical and Healthcare | Large-Size Company | http://gyrusacmi.com/ | USA | The patent focuses on a system for spectral collection during in-vivo procedures that emits an illumination signal, collects a return signal, and emits a continuous-wave laser signal from the same source. It addresses the difficulty of getting clean spectroscopic data while a continuous-wave laser is operating. The coordinated emission and collection enable spectral analysis during continuous-wave laser operation. | הפטנט מתמקד במערכת לאיסוף ספקטרלי במהלך פרוצדורות in-vivo הפולטת אות תאורה, אוספת אות חזרה ופולטת אות לייזר בגל רציף מאותו מקור. הוא מתמודד עם הקושי בקבלת נתונים ספקטרוסקופיים נקיים בעוד לייזר בגל רציף פועל. הפליטה והאיסוף המתואמים מאפשרים ניתוח ספקטרלי במהלך פעולת לייזר בגל רציף. |  | 1 | [11 empty] | 1 | [8 empty] | 1 | [7 empty] | 1 |  | GYRUS ACMI INC DBA OLYMPUS SURGICAL TECHNOLOGIES AMERICA | GYRUS | GYRUS ACMI INC D/B/A OLYMPUS SURGICAL TECH AMERICA | GYRUS | Olympus | 2024-10-22 | 2024 | 2025-05-08 | 2025 | 2023-11-02 | 2023 | Application | Utility Invention | active | WO2024US52334 | 321207334 | [NA] | WO
430 | CN219461385U | Ablation needle for spinal ablation | The ablation needle comprises a handle, a needle rod, a needle head, a transmitting antenna and a cooling pipe, one end of the needle rod is fixedly connected with the handle, the other end of the needle rod is fixedly connected with the needle head, and the transmitting antenna and the cooling pipe are both arranged in a cavity structure of the needle head and a cavity structure of the needle rod; an annular groove is formed in the end, close to the needle head, of the needle rod in the circumferential direction, a thermocouple wire is arranged in the annular groove, a temperature measuring connector is arranged on the handle, and the thermocouple wire is connected with the temperature measuring connector through a wire. The outer side of the needle rod is provided with a non-metal heat-shrinkable sleeve used for fixing the thermocouple wire, and the thermocouple wire is located between the needle rod and the non-metal heat-shrinkable sleeve. The ablation needle can accurately measure the temperature of the tissue around the needle head, so that the ablated damaged tissue can be judged according to the temperature. | 1. An ablation needle for spinal ablation, comprising a handle, a needle shaft, a needle head, a transmitting antenna and a cooling tube, one end of the needle shaft is fixedly connected to the handle, and the other end of the needle shaft is fixedly connected to the needle head, the The transmitting antenna and the cooling pipe are all arranged in the cavity structure of the needle head and the needle bar; it is characterized in that: one end of the needle bar near the needle head is provided with an annular groove along the circumference, and a thermocouple wire is arranged in the annular groove, so that The handle is provided with a temperature-measuring joint, and the thermocouple wire is connected to the temperature-measuring joint through a wire; the outside of the needle bar is provided with a non-metallic heat-shrinkable sleeve for fixing the thermocouple wire, and the thermocouple wire Located between the needle shaft and the non-metallic heat shrinkable sleeve. 2. The ablation needle for spinal ablation according to claim 1, characterized in that: the needle is a blunt needle. 3. The ablation needle for spinal ablation according to claim 1 or 2, characterized in that: the tail of the transmitting antenna is sheathed with a choke, and the choke is made of a conductive material. 4. The ablation needle for spinal ablation according to claim 3, characterized in that: the choke coil is a copper sleeve or a sleeve formed by electroplating. 5. The ablation needle for spinal ablation according to claim 1 and 2, characterized in that: the needle head and the needle shaft are connected by gluing. 6. The ablation needle for spinal ablation according to claim 1 or 2, characterized in that: the needle is made of ceramic material. 7. The ablation needle for spinal ablation according to claim 1 or 2, characterized in that: the thickness of the non-metallic heat-shrinkable sleeve is 0.02-0.05mm. 8. The ablation needle for spinal ablation according to claim 1 or 2, characterized in that: the guide wire passes through the side wall of the needle shaft from the annular groove and extends to the handle from the cavity structure of the needle shaft. 9. The ablation needle for spinal ablation according to claim 1 or 2, characterized in that: the non-metallic heat-shrinkable sleeve is made of polytetrafluoroethylene or PFA plastic. | 2022 | NANJING YIGAO MEDICAL TECH | Medical and Healthcare | NA | https://www.yigaomedical.com/en/ | China | The patent focuses on a utility-model ablation needle for spinal ablation with a handle, needle rod, needle head, transmitting antenna, and cooling pipe arranged in the cavity structure. It addresses the difficulty of safely ablating spinal lesions where heat management is critical to avoid neurological injury. Integrated cooling and a contained antenna reduce thermal risk during spinal ablation. | הפטנט מתמקד במחט אבלציה במודל שימושיות לאבלציית עמוד שדרה עם ידית, מוט מחט, ראש מחט, אנטנת שידור וצינור קירור המסודרים במבנה החלל. הוא מתמודד עם הקושי באבלציה בטוחה של נגעי עמוד שדרה שבהם ניהול חום הוא קריטי כדי להימנע מפגיעה נוירולוגית. קירור משולב ואנטנה תחומה מפחיתים סיכון תרמי במהלך אבלציית עמוד שדרה. |  |  |  |  | 1 | [25 empty] | 1 |  | NANJING YIGAO MEDICAL TECH CO LTD | NANJING YIGAO MEDICAL TECH CO LTD | NANJING YIGAO MEDICAL TECH CO LTD | NANJING YIGAO MEDICAL TECH CO LTD | NANJING YIGAO MEDICAL TECH CO LTD | 2022-10-21 | 2022 | 2023-08-04 | 2023 | 2022-10-21 | 2022 | Grant | Utility Model | active | CN202222780647U | 290423318 | [NA] | CN
431 | CN118717273A | Selectable and applicable IRE and RFA mixed ablation system and control method thereof | The invention relates to the technical field of tumor ablation, in particular to a selectable and applicable IRE and RFA mixed ablation system and a control method thereof. A selectable and applicable IRE and RFA mixed ablation system comprises an electrode needle, a data acquisition module, a heart rate monitoring unit, an output control module and an upper computer system, the output control module comprises a high-voltage pulse unit, a radio frequency current unit and an output switching control unit, and the upper computer system comprises a storage module and a control module. The storage module stores data acquired by the heart rate monitoring unit and the data acquisition module, and the control module receives and processes the data of the storage module and further controls the output control module to output a high-voltage pulse and then output a radio-frequency current or to output the radio-frequency current and then output the high-voltage pulse. Compared with single RFA or IRE ablation treatment, the mixed ablation technology combining the RFA ablation mode and the IRE ablation mode has the advantages that a larger ablation range can be obtained through mixed ablation, and a larger tumor can be ablated under single-time needle insertion. | 1. A selectively applicable IRE and RFA hybrid ablation system, characterized in that it comprises: an electrode needle (1), which extends into the area to be ablated in the human body and has a temperature sensor integrated thereon; a data acquisition module (2), which comprises a voltage/current measurement unit (21) for collecting and processing the voltage/current signal of the system, a conductivity measurement unit (22) for collecting and processing the conductivity signal of biological tissue near the electrode needle (1), and a temperature measurement unit (23) for collecting and processing the temperature signal of the temperature sensor; a heart rate monitoring unit (4) for monitoring the actual heart rate of the human body; an output control module (3), which comprises a high-voltage pulse unit (31), a radio frequency current unit (32), and an output switching control unit (33) for switching the output of the high-voltage pulse unit (31) and the radio frequency current unit (32); The upper computer system (5) comprises a storage module (51) and a control module (52), wherein the storage module (51) stores data collected by the heart rate monitoring unit (4) and the data collection module (2), and the control module (52) receives the data from the storage module (51) for processing, and further controls the output control module (3) to first output a high-voltage pulse and then output a radio frequency current, or first output a radio frequency current and then output a high-voltage pulse. 2.The selectively applicable IRE and RFA hybrid ablation system according to claim 1 is characterized in that it also includes a display module (6) for displaying the voltage and current data collected and processed by the voltage/current measurement unit (21). 3.The selectively applicable IRE and RFA hybrid ablation system according to claim 2 is characterized in that: the display module (6) is an oscilloscope. 4.The selectively applicable IRE and RFA hybrid ablation system according to claim 1 is characterized in that: the high-voltage pulse unit (31) and the radio frequency current unit (32) are connected to the output control module (3), the high-voltage pulse unit (31) includes a voltage regulation control circuit (311) and a MOS tube control circuit (312), and the radio frequency current unit (32) includes a radio frequency voltage control circuit. 5.The selectively applicable IRE and RFA hybrid ablation system according to claim 4 is characterized in that: the voltage regulation control circuit is provided with a DC/DC converter, the MOS tube control circuit is a bridge circuit, and the radio frequency voltage control circuit is provided with a DC/AC converter, a high-frequency filter and a transformer. 6.The selectively applicable IRE and RFA hybrid ablation system according to claim 1 is characterized in that: the voltage/current measurement unit (21), the conductivity measurement unit (22) and the temperature measurement unit (23) all have an acquisition card and an A/D converter, and the voltage/current detection unit (21) includes a voltage sensor and a current sensor. 7.The selectively applicable IRE and RFA hybrid ablation system according to claim 1 is characterized in that: the control module has a PID controller. 8A control method for a selectively applicable IRE and RFA hybrid ablation system as claimed in any one of claims 1 to 7, characterized in that it comprises the following process steps: S 1. The electrode needle (1) is percutaneously punctured to the target position of the area to be ablated, and the tail end of the electrode needle (1) is connected to the system output port, and the heart rate monitoring unit (4) is turned on to monitor the actual heart rate; S2. According to the medical imaging data, the treatment parameters are set on the host computer system (5) and the hybrid ablation mainly based on RFA or mainly based on IRE is selected; S3. If hybrid ablation based on RFA is selected, then: (a) multiple high-voltage pulse trains are output to bombard biological tissue near the electrode needle (1), and the conductivity of the biological tissue environment is improved while IRE ablation is performed; during the discharge process, the data acquisition unit (2) collects the voltage and current data output by the system circuit through the voltage sensor and the current sensor, and transmits them to the voltage/current measurement unit (21) in real time or at intervals for data processing, and finally sends the processed voltage/current data to the storage module (51) for storage and display on the oscilloscope; At the same time, during the discharge process, the data acquisition unit (2) collects temperature data of biological tissue near the electrode needle (1) through the temperature sensor in the electrode needle (1), and transmits it to the temperature measurement unit (23) in real time or at intervals for data processing, and finally sends the processed temperature data to the storage module (51) for storage, and feeds back to the host computer system (5); At the same time, in the gap between the pulse trains, the data acquisition unit (2) measures the conductivity of biological tissue near the electrode needle (1) through the positive and negative electrodes at the discharge end of the electrode needle (1), and transmits the collected signal to the conductivity measurement unit (22) for data processing to obtain conductivity and resistance value data, and finally sends the processed data to the storage module (51) for storage, and feeds back to the host computer system (5); (b) the heart rhythm monitoring unit (4) sends the ECG signal to the storage module (51) in real time for storage, and feeds back to the host computer system (5); the host computer system (5) determines whether the ECG signal is normal, and if so, proceeds to the next step; if abnormal, stops the output immediately; the host computer system (5) determines whether the current exceeds the set current safety threshold, and if so, stops the small B ablation; if not, proceeds to the next step; (c) After receiving the data of conductivity and resistance value, the upper computer system (5) performs AND gate operation. When the conductivity increases with the output of the pulse train, that is, the resistance value decreases with the output of the pulse train; and at the same time, when the number of pulse train outputs is less than the target number of pulse train outputs, IRE ablation continues and returns to step (a), otherwise proceed to the next step;(d) After the high-voltage pulse stops outputting, that is, after the IRE treatment stops, the radio frequency current starts to output, that is, the RFA treatment starts; (e) The data acquisition unit (2) continues to collect, process and feedback data during the discharge process; (g) Determine whether the total treatment time has reached the target time. If so, all ablations are completed; if not, proceed to the next step; (h) After the upper computer system (5) receives the temperature, conductivity or resistance value data, it performs an OR gate operation. When the conductivity begins to decrease, and the maximum conductivity in the current RFA ablation minus the current conductivity is greater than the set conductivity difference, otherwise the resistance value begins to increase, and the current resistance minus the minimum resistance in the current treatment is greater than the target resistance value; or when the temperature is greater than the set target temperature, stop RFA ablation, wait for a certain period of time, switch to IRE ablation, and return to step (a); otherwise continue RFA ablation; If mixed ablation with IRE as the main method is selected, then, (=1 \* roman dFirst, output radio frequency current for RFA ablation to increase the temperature of biological tissue; During the discharge process, the data acquisition unit (2) acquires the voltage and current data output by the system circuit through the voltage sensor and the current sensor, transmits the data to the voltage/current measurement unit (21) in real time or at intervals for data processing, and finally sends the processed voltage and current data to the storage module (51) for storage and display on the oscilloscope;At the same time, during the discharge process, the data acquisition unit (2) collects temperature data of biological tissue near the electrode needle (1) through the temperature sensor in the electrode needle (1), and transmits it to the temperature measurement unit (23) in real time or at intervals for data processing, and finally sends the processed temperature data to the storage module (51) for storage, and feeds back to the host computer system (5); At the same time, in the gap between the pulse trains, the data acquisition unit (2) measures the conductivity of biological tissue near the electrode needle (1) through the positive and negative electrodes at the discharge end of the electrode needle, and transmits the collected signal to the conductivity measurement unit (22) for data processing to obtain conductivity and resistance value data, and finally sends the processed data to the storage module (51) for storage, and feeds back to the host computer system (5); (=2 \* roman ii) After receiving the temperature data, the host computer system (5) adjusts the output radio frequency current through the PID controller feedback to maintain the temperature within a certain range above and below the set temperature for a certain time. When the time is reached, the RFA ablation is stopped and immediately switched to IRE ablation; (=3 \* roman iii) Start outputting high-voltage pulse trains to perform IRE ablation; (=4 \* roman iv) During the discharge process, the data acquisition unit (2) continues to collect, process and feedback the data; (=5 \* roman v) The heart rhythm monitoring unit (4) sends the ECG signal to the storage module (51) in real time for storage and feedback To the host computer system (5); the host computer system (5) determines whether the ECG signal is normal, if it is normal, proceed to the next step; if there is an abnormality, stop the output immediately; determine whether the total treatment time reaches the target treatment time, if so, all ablations are completed; if not, proceed to the next step; (=6 \* roman vi) After the host computer system (5) receives the conductivity and resistance value data, it performs an AND gate operation. When the conductivity increases with the output of the pulse train, conversely, the resistance value decreases with the output of the pulse train; and the number of pulse train outputs is less than the target output pulse train number, continue IRE ablation, otherwise all ablations are completed; S4. Turn off the output, save the data, and remove the electrode needle (1). 9. The control method of the selectively applicable IRE and RFA hybrid ablation system according to claim 8 is characterized in that: in the step S2, the type of hybrid ablation is determined according to the medical imaging data, specifically: if the area to be ablated is near a large blood vessel, hybrid ablation based on IRE is selected to reduce the impact of heat dissipation on the ablation range; if a large ablation range is required, hybrid ablation based on RFA is selected, and ablation based on RFA is selected. | 2024 | HUIWEI MEDICAL TECH NINGBO | NA | NA | NA | China | The patent focuses on a selectable IRE-and-RFA mixed ablation system with electrode needle, data acquisition, heart-rate monitoring, output control, and a host computer. It addresses the limitation of single-modality ablation that may miss certain tumor types where mixed IRE/RFA could be more effective. The system selects the right modality and monitors heart rate for safety during tumor ablation. | הפטנט מתמקד במערכת אבלציה משולבת IRE-ו-RFA בחירה עם מחט אלקטרודה, איסוף נתונים, ניטור קצב לב, בקרת פלט ומחשב מארח. הוא מתמודד עם המגבלה של אבלציה במודאליות יחידה שעלולה להחמיץ סוגי גידולים מסוימים שבהם IRE/RFA מעורב יכול להיות יעיל יותר. המערכת בוחרת את המודאליות הנכונה ומנטרת קצב לב לבטיחות במהלך אבלציית גידולים. | 1 |  |  |  | 1 | [5 empty] | 1 | 1 | [16 empty] | 1 |  |  |  | HUIWEI MEDICAL TECH NINGBO CO LTD | HUIWEI MEDICAL TECH NINGBO CO LTD | HUIWEI MEDICAL TECH NINGBO CO LTD | HUIWEI MEDICAL TECH NINGBO CO LTD | HUIWEI MEDICAL TECH NINGBO CO LTD | 2024-07-24 | 2024 | 2024-10-01 | 2024 | 2024-07-24 | 2024 | Application | Utility Invention | active | CN202410998758 | 315299409 | [NA] | CN
432 | CN114886553A | Temperature control method and device, electronic equipment and storage medium | The invention provides a temperature control method and device, electronic equipment and a storage medium, the temperature control method and device are applied to a control center of a microwave ablation system, the microwave ablation system comprises a temperature measuring needle, an ablation needle, the control center and a water pump, and the method comprises the steps that if the first temperature of the temperature measuring needle is larger than the first preset temperature, the temperature of the temperature measuring needle is controlled according to the difference value between the first temperature and the first preset temperature; controlling the water pump rotating speed to be increased from the first rotating speed to the second rotating speed; if the second temperature of the temperature measuring needle is greater than or equal to the second preset temperature, closing the heating function of the ablation needle and reducing the rotating speed of the water pump from the second rotating speed to a third rotating speed; if the third temperature of the temperature measuring needle is smaller than the first preset temperature, the fourth temperature of the ablation needle is smaller than the first preset temperature, and the accumulated heating duration of the ablation needle in the treatment is smaller than the preset treatment duration, the heating function of the ablation needle is started. By means of the method, the temperature of the ablation needle can be automatically adjusted, and the problem that normal tissue of the body is damaged by high temperature due to untimely temperature adjustment of the ablation needle is effectively solved. | 1. A temperature control method, characterized in that it is applied to a control center of a microwave ablation system, wherein the microwave ablation system comprises a temperature measuring needle, an ablation needle, a control center and a water pump, and the control center is respectively connected with the ablation needle, the ablation needle, and the water pump. The temperature measuring needle and the water pump are communicatively connected, and the water pump is connected to the ablation needle for adjusting the speed of water circulation in the ablation needle. The method includes: acquiring a first temperature of the temperature measuring needle, and when When the first temperature is greater than the first preset temperature, according to the first difference between the first temperature and the first preset temperature, the rotation speed of the water pump is controlled to increase from the first rotation speed to the second rotation speed; the temperature measurement The needle is used to measure the edge temperature of the target site heated by the ablation needle; the second rotational speed is proportional to the first difference value; obtain the second temperature of the temperature measuring needle, if the second temperature is greater than or is equal to the second preset temperature, turn off the heating function of the ablation needle and reduce the rotation speed of the water pump from the second rotation speed to a third rotation speed; obtain the current third temperature of the temperature measuring needle, the ablation needle the fourth temperature and the cumulative heating time of the ablation needle in this treatment, if the third temperature is less than the first preset temperature, the fourth temperature is less than the first preset temperature and the If the accumulated heating duration is less than the preset treatment duration, the heating function of the ablation needle is turned on. 2. The method according to claim 1, further comprising: acquiring the temperature of the ablation needle in real time, and when the temperature of the ablation needle is greater than or equal to a third preset temperature, turning off the ablation needle Needle heating function. 3. The method according to claim 1, wherein after acquiring the second temperature of the temperature measuring needle, the method further comprises: if the second temperature is less than the first preset temperature, according to The second difference between the second temperature and the first preset temperature controls the rotational speed of the water pump to decrease from the second rotational speed to a fourth rotational speed. 4. The method according to claim 2, wherein the method for determining the cumulative heating duration comprises: determining the heating start time and the heating off time of the ablation needle each time the heating function of the ablation needle is turned off and heating duration; obtain the estimated treatment time period of the current treatment, and determine the cumulative heating duration as the sum of at least one heating duration with the heating start time between the expected treatment time periods. 5. The method according to claim 4, wherein after the accumulated heating duration is less than a preset treatment duration, the method further comprises: determining the first time between the preset treatment duration and the accumulated heating duration Three difference values; adjusting the estimated time after the current time passes the third difference time as the estimated ending time of the current treatment; adjusting the estimated treatment time period according to the starting time and the estimated ending time. 6. The method according to claim 1, wherein after acquiring the cumulative heating duration of the ablation needle in this treatment, the method further comprises: if the cumulative heating duration is greater than or equal to the preset Treatment duration, turn off the heating function of the ablation needle and end this treatment; according to the start time and actual end time of this treatment, the ablation needle and the temperature measuring needle at the start time and the actual end time generating ablation treatment information from a temperature change curve between the ablation treatment information and the patient's identity information. 7. The method according to claim 1, wherein when the first temperature is greater than a first preset temperature, the water pump is controlled according to a first difference between the first temperature and the first preset temperature The speed is increased from the first speed to the second speed speed, including: taking the time as the abscissa and the temperature as the ordinate, drawing the temperature change curve of the temperature measuring needle in this treatment; determining the first moment when the first temperature is greater than the first preset temperature, based on the The temperature value of the temperature change curve at each moment in the preset time period after the first moment, to determine the rising acceleration of the water pump rotation speed; according to the rising acceleration, starting from the first moment, the water pump The rotational speed is adjusted from the first rotational speed to the second rotational speed. 8. A temperature control device, characterized in that the device resides in a control center of a microwave ablation system, the microwave ablation system comprises a temperature measuring needle, an ablation needle, a control center and a water pump, the water pump and the ablation needle connected to the ablation needle for adjusting the speed of water circulation in the ablation needle, the device includes: a first control unit, configured to acquire a first temperature of the temperature measuring needle, when the first temperature is greater than a first preset temperature , according to the first difference between the first temperature and the first preset temperature, control the rotation speed of the water pump to increase from the first rotation speed to the second rotation speed; the temperature measuring needle is used to measure the target site heated by the ablation needle the edge temperature of temperature, the heating function of the ablation needle is turned off and the rotation speed of the water pump is reduced from the second rotation speed to the third rotation speed; a third control unit is used to obtain the current third temperature of the temperature measurement needle, the The fourth temperature of the ablation needle and the cumulative heating time of the ablation needle in this treatment, if the third temperature is less than the first preset temperature, the fourth temperature is less than the first preset temperature and If the accumulated heating duration is less than the preset treatment duration, the heating function of the ablation needle is turned on. 9. An electronic device, characterized in that, comprising: a processor, a storage medium and a bus, wherein the storage medium stores machine-readable instructions executable by the processor, and when the electronic device runs, the processor and the The storage media communicate through a bus, and the processor executes the machine-readable instructions to perform the steps of the temperature control method according to any one of claims 1 to 7. 10. A computer-readable storage medium, wherein a computer program is stored on the computer-readable storage medium, and when the computer program is run by a processor, the temperature control according to any one of claims 1 to 7 is executed. steps of the method. | 2022 | NANJING NUOYUAN MEDICAL DEVICES | NA | NA | NA | China | The patent focuses on a temperature control method for a microwave ablation system that uses a temperature-measuring needle, an ablation needle, control center, and water pump for cooling. It addresses the safety risk that microwave ablation overheating without responsive cooling can cause unwanted tissue damage. The system adjusts cooling water flow when temperature thresholds are reached. | הפטנט מתמקד בשיטת בקרת טמפרטורה למערכת אבלציית מיקרוגלים המשתמשת במחט מודדת טמפרטורה, מחט אבלציה, מרכז בקרה ומשאבת מים לקירור. הוא מתמודד עם סיכון הבטיחות שהתחממות יתר באבלציית מיקרוגלים ללא קירור מגיב יכולה לגרום לנזק לא רצוי לרקמה. המערכת מתאימה את זרימת מי הקירור כאשר ספי טמפרטורה מגיעים. |  |  |  |  | 1 | [7 empty] | 1 | [17 empty] | 1 |  | NANJING NUOYUAN MEDICAL DEVICES CO LTD | NANJING NUOYUAN MEDICAL DEVICES CO LTD | NANJING NUOYUAN MEDICAL DEVICES CO LTD | NANJING NUOYUAN MEDICAL DEVICES CO LTD | NANJING NUOYUAN MEDICAL DEVICES CO LTD | 2022-05-23 | 2022 | 2022-08-12 | 2022 | 2022-05-23 | 2022 | Application | Utility Invention | active | CN202210560151 | 284145795 | [NA] | CN
433 | CN114521954A | Ablation electrode with miniature endoscope | The ablation electrode comprises a handle, the front end of the handle is connected with a rod part, the top end of the rod part is provided with an electrode tip, the rod part is internally provided with the miniature endoscope, the miniature endoscope is connected with a push button arranged on the handle, and the miniature endoscope can penetrate out of the rod part under the action of the push button. The ablation condition can be visually observed in the operation through the micro endoscope, and the tumor ablation progress can be fed back in real time; through video function transmission, the change of a tumor in a radiofrequency ablation operation is recorded, the three-dimensional shape of the tumor is simulated through a naked eye 3D technology, and an operator is assisted to more effectively judge the specific position and the size of the tumor; the operation is safer and more convenient, the tumor treatment means are perfected, the incidence of diseases during and after the operation is reduced, and the life quality of tumor patients is practically and effectively improved. | 1. An ablation electrode with a miniature endoscope, comprising a handle (1), the front end of the handle (1) is connected with a rod (2), and the top of the rod (2) is provided with an electrode head (3), characterized in that the The rod portion (2) is provided with a miniature endoscope (4), the miniature endoscope (4) is connected with a push button (12) provided on the handle (1), and the miniature endoscope (4) is connected to the push button (12). Under the action of 12), the rod portion (2) can be passed through. 2The ablation electrode with a miniature endoscope according to claim 1, wherein the handle (1) comprises a casing (11), and a push button (12) is provided on the casing (11), The casing (11) is provided with a cable connected with the electrode head (3) and a video transmission line connected with the miniature endoscope (4). 3. The ablation electrode with a miniature endoscope according to claim 1, wherein the rod portion (2) comprises a stainless steel tube (21), and the stainless steel tube (21) is sleeved with a miniature endoscope In the mirror tube (44) of (4), one end of the stainless steel tube (21) close to the electrode head (3) is provided with a through hole (22) through which the mirror tube (44) passes. 4. The ablation electrode with a miniature endoscope according to claim 3, wherein the stainless steel tube (21) is provided with an insulating coating or an insulating sleeve. 5. An ablation electrode with a miniature endoscope according to claim 3, characterized in that, the miniature endoscope (4) is a double miniature endoscope mode, comprising an axial endoscope (41) and Radial endoscope (42). 6The ablation electrode with a miniature endoscope according to claim 4, wherein the miniature endoscope (4) comprises a lens module (43) arranged on the top of the mirror tube (44). 7. The ablation electrode with a miniature endoscope according to claim 5, wherein the lens module (43) comprises a camera head (431) arranged in the middle of the lens module (43) and a camera head (431) surrounding the camera head (431). 431) A lighting system (432) is provided, and the lighting area of the lighting system (432) covers the field of view of the camera (431). 8The ablation electrode with a miniature endoscope according to any one of claims 1-7, characterized in that, one end of the electrode head (3) in the rod part (2) is flat. | 2022 | ZHEJIANG MUSHI MEDICAL SCIENCE AND TECH | NA | NA | NA | China | The patent focuses on an ablation electrode with a built-in miniature endoscope at the rod tip, retractable via a push button on the handle. It addresses the operator's lack of direct visualization at the ablation tip during procedures. The miniature endoscope can extend out for direct visualization of the target area. | הפטנט מתמקד באלקטרודת אבלציה עם אנדוסקופ מיניאטורי מובנה בקצה המוט, נסוג באמצעות לחצן בידית. הוא מתמודד עם החוסר של המפעיל בהדמיה ישירה בקצה האבלציה במהלך פרוצדורות. האנדוסקופ המיניאטורי יכול להתפשט החוצה להדמיה ישירה של אזור המטרה. | [8 empty] | 1 |  | 1 | [17 empty] | 1 |  |  |  | ZHEJIANG MUSHI MEDICAL SCIENCE AND TECH LIMITED COMPANY | ZHEJIANG MUSHI MEDICAL SCIENCE AND TECH LIMITED COMPANY | ZHEJIANG MUSHI MEDICAL SCIENCE AND TECH LIMITED COMPANY | ZHEJIANG MUSHI MEDICAL SCIENCE AND TECH LIMITED COMPANY | ZHEJIANG MUSHI MEDICAL SCIENCE AND TECH LIMITED COMPANY | 2022-03-01 | 2022 | 2022-05-24 | 2022 | 2022-03-01 | 2022 | Application | Utility Invention | active | CN202210191026 | 282514093 | [NA] | CN
434 | US11553962B2 | Systems for cardiac ablation and associated methods | Devices, systems, and methods for treating cardiac arrhythmia are disclosed herein. In some embodiments, devices, systems, and methods disclosed herein deliver interrogating energy to tissue at a position on a wall of an anatomical structure of a patient. If the devices, systems, and methods disclosed herein detect a change in electrical activity of the anatomical structure in response to the interrogating energy, the devices, systems, and methods disclosed herein can apply irreversible therapy to the tissue. In some embodiments, the change in electrical activity corresponds to slowing or termination of a detected arrhythmia. | 1. A method, comprising: delivering interrogating energy to tissue at a position on a wall of an anatomical structure using a catheter positioned within the anatomical structure, wherein delivering the interrogating energy includes applying non-thermal therapy such that a temperature of the tissue does not change by more than 1° C.; detecting a change in electrical activity of the anatomical structure in response to delivery of the interrogating energy, wherein the change in electrical activity corresponds to slowing or termination of a detected arrhythmia; and applying irreversible therapy to the tissue in response to detecting the change in electrical activity. 2. The method of claim 1 wherein the tissue is first tissue, wherein the arrythmia is a first arrythmia, and wherein the method further comprises: inducing the first arrythmia before delivering the interrogating energy by stimulating the first tissue or second tissue of the anatomical structure; or re-inducing the first arrythmia or a second arrhythmia by stimulating the first tissue, the second tissue, or third tissue of the anatomical structure after terminating the first arrhythmia. 3. The method of claim 1 wherein applying irreversible therapy to the tissue includes applying irreversible therapy to the tissue via the catheter. 4. The method of claim 1 wherein the catheter is a first catheter, and wherein applying irreversible therapy to the tissue includes applying irreversible therapy to the tissue using a second catheter positioned within the anatomical structure. 5. The method of claim 1 wherein the irreversible therapy includes pulsed field ablation, radiofrequency ablation, cryo-ablation, ultrasound ablation, laser balloon ablation, or hot balloon ablation. 6. The method of claim 1 wherein applying the interrogating energy includes applying an interrogating electrical pulse. 7. The method of claim 6 wherein the interrogating electrical pulse is a biphasic electrical signal. 8. The method of claim 7 wherein the polarities of the biphasic electrical signal are symmetric. 9. The method of claim 7 wherein the polarities of the biphasic electrical signal are asymmetric. 10. The method of claim 7 wherein the interrogating electrical pulse is a biphasic electrical signal such that each polarity contains not less than 80% of charge delivered of the other polarity in aggregate over a period of 10 ms or less. 11. The method of claim 6 wherein the interrogating electrical pulse is a square wave, a sine function, a raised root cosine, a Gaussian function, or a trapezoid wave. 12. The method of claim 1 wherein applying the interrogating energy includes applying at least 1000 volts for at least 100 nanoseconds to the tissue at least once. 13. The method of claim 1, further comprising capturing one or more electrical signals of the anatomical structure, and wherein detecting the change in the electrical activity of the anatomical structure includes detecting a change in (i) timing of the one or more captured electrical signals or (ii) morphology of the one or more captured electrical signals. 14. The method of claim 1, further comprising performing activation mapping or entrainment mapping of the anatomical structure to identify an area of interest on a wall of the anatomical structure. 15. The method of claim 1, further comprising: determining a location of a tip section of the catheter within the anatomical structure when the interrogating energy is delivered to the tissue; and displaying a therapy annotation in a three-dimensional model of the anatomical structure at a location in the model corresponding to the determined location of the tip section of the catheter in the anatomical structure. 16. The method of claim 15 wherein the therapy annotation includes one or more visual properties, wherein the one or more visual properties includes a color or a pattern, and wherein the color or the pattern provide an indication of whether the interrogating energy delivered to the tissue terminated the detected arrhythmia. 17. The method of claim 16 wherein at least one of the one or more visual properties of the displayed therapy annotation changes as a function of time. 18. The method of claim 1, further comprising: determining a location of a tip section of the catheter within the anatomical structure when the irreversible therapy is delivered to the tissue; and displaying a therapy annotation in a three-dimensional model of the anatomical structure at a location in the model corresponding to the determined location of the tip section of the catheter in the anatomical structure. 19. The method of claim 1 wherein delivering interrogating energy to the tissue includes delivering interrogating energy to a first area of tissue, wherein applying irreversible therapy to the tissue includes applying irreversible therapy to a second area of tissue, and wherein the first area of tissue is greater than the second area. 20. A system, comprising: one or more electrodes configured to capture an electrocardiogram (ECG) of an anatomical structure of a patient or an intracardiac reference positioned within the anatomical structure and configured to capture an electrogram of the anatomical structure; a catheter configured to be inserted within the anatomical structure, wherein the catheter includes at least one electrode configured to deliver interrogating energy and irreversible therapy to tissue on a wall of the anatomical structure; a generator configured to deliver the interrogating energy and irreversible therapy to the tissue via the at least one electrode; and a memory operably coupled to a processor and storing instructions that, when executed by the processor, cause the system to perform operations comprising: detecting an arrhythmia of the anatomical structure in the ECG or in the electrogram; delivering interrogating energy to the tissue at a first location; detecting a change in the ECG or in the electrogram in response to delivery of the interrogating energy to the tissue, wherein the change corresponds to slowing or termination of the detected arrhythmia; and delivering irreversible therapy to the tissue at the first location. 21. The system of claim 20 wherein the intracardiac reference is a second catheter different from the catheter, and wherein the second catheter is a multipolar catheter or a coronary sinus catheter. 22. The system of claim 20 wherein the interrogating energy comprises an interrogating electrical pulse, and wherein the interrogating electrical pulse is a biphasic electrical signal. 23. The system of claim 20 wherein: the operations further comprise displaying a therapy annotation in a three-dimensional model of the anatomical structure; the therapy annotation includes one or more visual properties; the one or more visual properties include a color or a pattern; and the color or the pattern provide an indication of whether the interrogating energy delivered to the tissue terminated the detected arrhythmia. 24. The system of claim 20 wherein detecting an arrythmia comprises measuring a cycle length. 25. A non-transitory, computer-readable medium storing instruction thereon that, when executed by at least one processor of a cardiac ablation system, cause the cardiac ablation system to perform operations comprising: delivering interrogating energy to tissue at a position on a wall of an anatomical structure using a catheter positioned within the anatomical structure; detecting a change in electrical activity of the anatomical structure in response to delivery of the interrogating energy, wherein the change in electrical activity corresponds to slowing or termination of a detected arrhythmia; and applying irreversible therapy to the tissue in response to detecting the change in electrical activity. 26. The non-transitory, computer-readable medium of claim 25 wherein applying irreversible therapy to the tissue includes applying irreversible therapy to the tissue via the catheter. 27. The non-transitory, computer-readable medium of claim 25 wherein the catheter is a first catheter, and wherein applying irreversible therapy to the tissue includes applying irreversible therapy to the tissue using a second catheter positioned within the anatomical structure. 28. The non-transitory, computer-readable medium of claim 25 wherein detecting a change in electrical activity of the anatomical structure comprises measuring a cycle length. | 2020 | AFFERA INC | Medical and Healthcare | Small-size company | https://affera.com/ | USA | The patent focuses on a cardiac ablation system that delivers interrogating energy to tissue and detects changes in the anatomical structure's electrical activity in response. It addresses the difficulty of confirming whether a target site is appropriate before ablating in cardiac arrhythmia treatment. The interrogating step verifies the target before the system commits to ablation. | הפטנט מתמקד במערכת אבלציה לבבית המספקת אנרגיית חקירה לרקמה ומזהה שינויים בפעילות החשמלית של המבנה האנטומי בתגובה. הוא מתמודד עם הקושי באישור האם אתר מטרה מתאים לפני אבלציה בטיפול בהפרעות קצב לבביות. שלב החקירה מאמת את המטרה לפני שהמערכת מתחייבת לאבלציה. | 1 | [23 empty] | 1 | [7 empty] | AFFERA INC | AFFERA INC | AFFERA, INC. | AFFERA INC | AFFERA INC | 2021-01-20 | 2021 | 2023-01-17 | 2023 | 2020-01-24 | 2020 | Grant | Utility Invention | active | US202117600523 | 319213112 | CN115334989A¦CN115998412A¦CN115998412B¦EP4093308A1¦EP4093308A4¦JP2023511419A¦JP7628547B2¦US12575880B2¦US20220346867A1¦US20230157751A1¦WO2021150629A1 | US
435 | US12427344B2 | Multi-functional sparse phased arrays for guiding focused ultrasound therapies | Multifunctional ultrasound systems and methods for body section registration and mapping of microbubble dynamics. A system is provided that includes one or more micromachined ultrasonic transducer arrays (MUTAs) configured to capture a high-resolution image of at least a portion of a body section using ultrasound and monitor microbubble activity during ultrasound treatment. The system includes an image registration module configured to spatially register the high-resolution image with a reference image. The system includes electronics configured to control one or more of drive signal amplitude, frequency filtering, multiplexing, and DC bias voltage. The system can be configured to control ultrasound treatment based on the Monitoring Verification monitoring of the microbubble activity during treatment. | 1. A system comprising: a first micromachined ultrasonic transducer array (MUTA) comprising one or more capacitive micromachined ultrasonic transducers (CMUTs) having a collapse voltage and comprising a substrate, a substrate electrode, edges, a suspended membrane constrained at the edges and suspended over the substrate, a membrane electrode, and a gap between the suspended membrane and the substrate; and electronics configured to control a vibrational mode of the CMUTs via application of a DC bias voltage across the electrodes of the CMUTs; wherein: the MUTA is configured to: capture a high-resolution image of at least a portion of a body section using ultrasound; and monitor microbubble activity during ultrasound treatment; the vibrational mode is selected from the group consisting of a non-collapsed mode, a collapsed mode, and a deep-collapsed mode; in the non-collapsed mode, the DC bias voltage is less than the collapse voltage, and vibration of the membrane between the constrained edges is free of contact with the substrate; in the collapsed mode, the DC bias voltage is the collapse voltage, and vibration of the membrane between the constrained edges is not free of contact with the substrate, with a first central portion of the membrane spanning across the gap and coming into contact with the substrate, thereby constraining the membrane at the edges and along a length of the first central portion; and in the deep-collapsed mode, the DC bias voltage is greater than the collapse voltage, and vibration of the membrane between the constrained edges is not free of contact with the substrate, with a second central portion of the membrane, being greater than the first central portion, spanning across the gap and coming into contact with the substrate, thereby constraining the membrane at the edges and along a length of the second central portion. 2. The system of claim 1 further comprising: an image registration module configured to spatially register the high-resolution image with a reference image; wherein the electronics is further configured to control one or more of drive signal amplitude, frequency filtering, and multiplexing; and wherein the electronics are integrated with the MUTA. 3. The system of claim 1, wherein the body section is selected from the group consisting of a skull, a skull-brain interface, a brain, a liver, an organ, and combinations thereof. 4. The system of claim 1 further comprising: a second MUTA; and an on-chip switch in communication with the MUTAs; wherein the on-chip switch is configured to adapt an effective transducer area of at least one of the MUTAs for improving an effective signal-to-noise ratio (SNR) of the system. 5. The system of claim 1, wherein the reference image is selected from the group consisting of a CAT scan image, a PET scan image, an MRI image, and combinations thereof. 6. The system of claim 4, wherein the MUTAs are adaptable by one or more of frequency response, output power, and spatial distribution over each array surface. 7. A multifunctional ultrasound system for body section registration and mapping of microbubble dynamics comprising: a first micromachined ultrasonic transducer array (MUTA) and one or more second MUTAs, the MUTAs comprising ultrasonic transducers having a membrane, wherein two or more of the ultrasonic transducers comprise different membrane geometries; and an image registration module configured to spatially register the high-resolution image with a reference image; wherein: (i) at least two of the MUTAs are disposed on a common substrate to adapt target frequency responses for at least one of: ultrasound imaging of the body section for image registration; ultrasound monitoring of microbubble activity; or generating ultrasound for treatment; (ii) at least one of: the target frequency response for ultrasound monitoring of microbubble activity is adapted to be complementary to at least a portion of the target frequency response for generating the ultrasound for treatment by using at least one of electronic filtering or DC biasing; each MUTA is configured to at least one of: monitor microbubble activity in a specific body section using image reconstruction; control microbubble activity during ultrasound treatment using image reconstruction; image a specific region of the body section for image registration; or image a specific region of the body section for treatment guidance; wherein the image reconstruction comprises at least one of plane wave spectrum combinations, delay and sum, phased array, synthetic aperture, or short spatial coherence; and (iii) at least one of: a target frequency response for capturing the image of the body section for image registration is in a range of 2 MHz to 15 MHz; a target frequency response for ultrasound monitoring of microbubble activity is in a range of 50 kHz to 5 MHz; a target frequency response for generating ultrasound for treatment is in a range of 50 kHz to 1 MHz; the different membrane geometries are characterized by lateral dimensions between 5 μm and 250 μm; the different membrane geometries are characterized by lateral dimensions between 250 μm and 1000 μm; the MUTAs are configured to detect activity of 1 to 1000 microbubbles during ultrasound treatment; the MUTAs are configured to capture the high-resolution image and monitor microbubble activity at a diffraction-limited spatial resolution; each MUTA is individually adaptable to monitor microbubble activity in a body section; or each MUTA is configured to individually generate an image of a portion of the body section. 8. The system of claim 7, wherein the system is configured to spatially map at least one of the microbubble activity or image the body section by image reconstruction. 9. A multifunctional ultrasound system for body section registration and mapping of microbubble dynamics comprising: a first micromachined ultrasonic transducer array (MUTA) and one or more second MUTAs, the MUTAs comprising ultrasonic transducers having a membrane; an image registration module configured to spatially register the high-resolution image with a reference image; and a control loop configured to maintain stable cavitation using an output signal of the MUTAs; wherein at least two of the MUTAs are disposed on a common substrate to adapt target frequency responses for at least one of: ultrasound imaging of the body section for image registration; ultrasound monitoring of microbubble activity; or generating ultrasound for treatment; and wherein at least one of: the target frequency response for ultrasound monitoring of microbubble activity is adapted to be complementary to at least a portion of the target frequency response for generating the ultrasound for treatment by using at least one of electronic filtering or DC biasing; two or more of the ultrasonic transducers comprise different membrane geometries; or each MUTA is configured to at least one of: monitor microbubble activity in a specific body section using image reconstruction; control microbubble activity during ultrasound treatment using image reconstruction; image a specific region of the body section for image registration; or image a specific region of the body section for treatment guidance; wherein the image reconstruction comprises at least one of plane wave spectrum combinations, delay and sum, phased array, synthetic aperture, or short spatial coherence. 10. The system of claim 7, wherein the MUTAs are further configured to generate focused ultrasound energy for treatment. 11. The system of claim 7, wherein the MUTAs are further configured to control the microbubble activity during the ultrasound treatment. 12. A method for body section image registration, mapping of microbubble dynamics, and ultrasound treatment with the system of claim 7 comprising: capturing, with ultrasound using the MUTAs, the high-resolution image; spatially registering with the image registration module the high-resolution image to the reference image; generating, with at least one of the MUTAs, focused ultrasound energy for ultrasound treatment; monitoring and controlling, with at least one of the MUTAs, microbubble activity during the ultrasound treatment; and combining one or more of output signals and captured high-resolution images from the MUTAs to form a composite image of the body section for at least one of image registration or treatment guidance. 13. The method of claim 12, wherein one or more of the MUTAs each comprise one or more capacitive micromachined ultrasonic transducers (CMUTs); and wherein the method further comprises adjusting a DC bias voltage to cause at least one of the CMUTs to operate in at least one of a non-collapsed mode, a collapsed mode, or a deep-collapsed mode. 14. The method of claim 12 further comprising electrically connecting two or more elements of the MUTAs by an on-chip switch to adapt an effective transducer area of the MUTAs for improving a signal-to-noise ratio (SNR). 15. The method of claim 12 further comprising: disposing two or more MUTAs on a common substrate to adapt target frequency responses for at least one of ultrasound imaging of the body section for image registration, ultrasound monitoring of microbubble activity, or generating ultrasound for treatment; and adapting the target frequency response for ultrasound monitoring of microbubble activity to be complementary to at least a portion of the target frequency response for generating the ultrasound for treatment by using at least one of electronic filtering or DC biasing; wherein a target frequency response for the ultrasound imaging of the body section for image registration is in a range of 2 MHz to 15 MHz; wherein a target frequency response for ultrasound monitoring of microbubble activity is in a range of 50 kHz to 5 MHz; and wherein a target frequency response for generating ultrasound for treatment is in a range of 50 kHz to 1 MHz. 16. A multifunctional ultrasound system for body section registration, mapping of microbubble dynamics, and ultrasound treatment, the system comprising: one or more micromachined ultrasonic transducer arrays (MUTAs) comprising ultrasonic transducers, each MUTA configured to: capture a high-resolution image of at least a portion of a body section using ultrasound; monitor and control microbubble activity during ultrasound treatment; and generate focused ultrasound energy for treatment; an image registration module configured to spatially register the high-resolution image with a reference image; and electronics configured to control a DC bias voltage to cause at least of the one ultrasonic transducers to operate in a different mode from another of the one or more ultrasonic transducers, the modes selected from a group consisting of a non-collapsed mode, a collapsed mode, and a deep-collapsed mode. 17. The system of claim 16, wherein the electronics are integrated with at least one of the one or more MUTAs; and wherein the electronics are further configured to control at least one of drive signal amplitude, frequency filtering, or multiplexing. 18. The system of claim 16 further comprising at least one on-chip switch in communication with at least two elements of the MUTAs; wherein the at least one on-chip switch is configured to adapt an effective transducer area of the one or more MUTAs for improving an effective SNR of the system. 19. The system of claim 16, wherein two or more MUTAs are disposed on a common substrate to adapt target frequency responses for at least one of ultrasound imaging of the body section for image registration, ultrasound monitoring of microbubble activity, or generating ultrasound for treatment; wherein a target frequency response for capturing the image of the body section for image registration is in a range of 2 MHz to 15 MHz; wherein a target frequency response for ultrasound monitoring of microbubble activity is in a range of 50 kHz to 5 MHz; and wherein a target frequency response for generating ultrasound for treatment is in a range of 50 kHz to 1 MHz. 20. The system of claim 19, wherein the target frequency response for ultrasound monitoring of microbubble activity is adapted to be complementary to at least a portion of the target frequency response for generating the ultrasound for treatment by using at least one of electronic filtering or DC biasing. 21. The system of claim 19, wherein the two or more MUTAs comprise different membrane geometries. 22. The system of claim 21, wherein the membrane geometries are characterized by lateral dimensions between 5 μm and 250 μm. 23. The system of claim 21, wherein the membrane geometries are characterized by lateral dimensions between 250 μm and 1000 μm. 24. The system of claim 16, wherein the system is configured to spatially map at least one of the microbubble activity or image the body section by image reconstruction. 25. The system of claim 24, wherein the image reconstruction comprises plane wave spectrum combinations, delay and sum, phased array, synthetic aperture, and short spatial coherence. 26. The system of claim 1, wherein at least one of the CMUTs operates in a different vibrational mode from another of the CMUTs. 27. The system of claim 7, wherein the MUTAs further comprise one or more piezoelectric micromachined ultrasonic transducers (PMUTs). | 2020 | Georgia Tech Research | University | NA | https://research.gatech.edu/ | USA | The patent focuses on multi-functional sparse phased-array ultrasound systems that capture high-resolution body-section images and monitor microbubble activity during focused ultrasound treatment. It addresses the limitation of conventional ultrasound arrays in providing both imaging and treatment monitoring at high resolution. The arrays support body-section registration, mapping, and microbubble dynamics monitoring. | הפטנט מתמקד במערכות אולטרסאונד מערך פאזה דליל רב-תכליתיות הלוכדות תמונות חתך גוף ברזולוציה גבוהה ומנטרות פעילות מיקרובועות במהלך טיפול אולטרסאונד ממוקד. הוא מתמודד עם המגבלה של מערכי אולטרסאונד קונבנציונליים במתן הן הדמיה והן ניטור טיפול ברזולוציה גבוהה. המערכים תומכים ברישום חתך גוף, מיפוי וניטור דינמיקה של מיקרובועות. |  |  | 1 | [12 empty] | 1 | [14 empty] | 1 |  | GEORGIA TECH RES INST | Georgia Tech Research Institute | GEORGIA TECH RESEARCH CORPORATION | Georgia Tech Research | Georgia Tech Research | 2021-06-11 | 2021 | 2025-09-30 | 2025 | 2020-06-11 | 2020 | Grant | Utility Invention | active | US202118007562 | 312982131 | EP4165438A1¦EP4165438A4¦US20230211187A1¦WO2021252833A1 | US
436 | US12343071B2 | Voltage controlled pulse sequences for irreversible electroporation ablations | At least some embodiments of the present disclosure are directed to an electroporation ablation system for treating targeted tissue in a patient. The electroporation ablation system comprises an ablation catheter including catheter electrodes configured to generate electric fields in the targeted tissue in response to a plurality of electrical pulse sequences delivered in a plurality of therapy sections; a controller configured to receive a first pulse voltage of a first electrical pulse sequence measured during a first therapy section of the plurality of therapy sections; and determine a charge voltage based on the first pulse voltage; and an electroporation generator. The electroporation generator is operatively coupled to the catheter electrodes and the controller and configured to deliver a second electrical pulse sequence at a controlled pulse voltage for a second therapy section of the plurality of therapy sections. | 1. An electroporation ablation system for treating targeted tissue in a patient, the electroporation ablation system comprising: an ablation catheter including: catheter electrodes configured to generate electric fields in the targeted tissue in response to a plurality of electrical pulse sequences delivered in a plurality of therapy sections; a controller configured to: receive a first pulse voltage of a first electrical pulse sequence measured during a first therapy section of the plurality of therapy sections; determine a charge voltage based on the first pulse voltage; and an electroporation generator operatively coupled to the catheter electrodes and the controller and configured to deliver a second electrical pulse sequence at a controlled pulse voltage for a second therapy section of the plurality of therapy sections, the second therapy section being after the first therapy section, the controlled pulse voltage being associated with the charge voltage wherein the electroporation generator is further configured to deliver a scan electrical pulse sequence at a scan voltage during a scan section prior to the plurality of therapy sections; and wherein the controller is further configured to determine an initial tissue impedance based on an initial pulse voltage of the scan electrical pulse sequence and an initial pulse current of the scan electrical pulse sequence measured during the scan section, and determine an initial charge voltage based on the initial tissue impedance. 2. The electroporation ablation system of claim 1, wherein the electroporation generator comprises a capacitor bank and the electroporation generator is configured to charge the capacitor bank to a voltage level of the charge voltage before a start of the second therapy section. 3. The electroporation ablation system of claim 1, wherein the first electrical pulse sequence comprises a plurality of first electrical pulses. 4. The electroporation ablation system of claim 3, wherein the first pulse voltage comprises one or more pulse voltages of the plurality of first electrical pulses measured during the first therapy section. 5. The electroporation ablation system of claim 1, wherein the controller is further configured to receive a first pulse current of a first electrical pulse sequence delivered during the first therapy section, wherein the controller is further configured to determine the charge voltage based on the first pulse voltage and the first pulse current. 6. The electroporation ablation system of claim 5, wherein the controller is further configured to determine a first tissue impedance based on the first pulse voltage and the first pulse current. 7. The electroporation ablation system of claim 6, wherein the controlled pulse voltage is a portion of the charge voltage. 8. The electroporation ablation system of claim 7, wherein a ratio of the controlled pulse voltage and the charge voltage is associated with the first tissue impedance. 9. The electroporation ablation system of claim 1, wherein the scan voltage is less than the controlled pulse voltage. 10. The electroporation ablation system of claim 1, wherein the scan electrical pulse sequence includes a single non-ablative electrical pulse. 11. The electroporation ablation system of claim 1, wherein the electroporation generator comprises a plurality of capacitor banks, wherein the electroporation generator is configured to charge at least one of the plurality of capacitor banks to a voltage level individually. 12. The electroporation ablation system of claim 11, wherein the electroporation generator is configured to use a first capacitor bank of the plurality of capacitor banks to deliver a pulse sequence for a specific therapy section of the plurality of therapy sections and charge a second capacitor bank of the plurality of capacitor banks to a voltage level of a determined charge voltage before a start of a therapy section immediately after the specific therapy section of the plurality of therapy sections. 13. The electroporation ablation system of claim 11, wherein the catheter electrodes comprise a plurality of electrode pairs, wherein each capacitor bank of the plurality of capacitor banks is operatively coupled to one or more electrode pairs of the plurality of electrode pairs. 14. The electroporation ablation system of claim 13, wherein the controller is configured to determine a bank charge voltage for each capacitor bank of the plurality of capacitor banks. 15. A method of using an electroporation ablation device, the method comprising: disposing a catheter of the electroporation ablation device anatomically proximate to a target ablation location, the catheter comprising one or more catheter electrodes and configured to generate electric fields in response to a plurality of electrical pulse sequences delivered in a plurality of therapy sections; receiving a first pulse voltage of a first electrical pulse sequence measured during a first therapy section of the plurality of therapy sections; determining a charge voltage based on the first pulse voltage; and delivering a second electrical pulse sequence at a controlled pulse voltage for a second therapy section of the plurality of therapy sections, the second therapy section being after the first therapy section, the controlled pulse voltage being associated with the charge voltage delivering a scan electrical pulse sequence during a scan section; receiving an initial pulse voltage of the scan electrical pulse sequence measured during the scan section; receiving an initial pulse current of the scan electrical pulse sequence measured during the scan section; and determining an initial tissue impedance based on the initial pulse voltage and the initial pulse current measured, wherein the scan section is before the first therapy section, wherein the scan electrical pulse sequence is at a scan pulse voltage lower than the controlled pulse voltage. 16. The method of claim 15, further comprising: receiving a first pulse current of first electrical pulse sequence measured during the first therapy section, wherein determining a charge voltage comprises determining the charge voltage based on the first pulse voltage and the first pulse current. 17. The method of claim 15, further comprising: determining a first tissue impedance based on the first pulse voltage and the first pulse current, wherein determining a charge voltage comprises determining the charge voltage based on the first tissue impedance. 18. The method of claim 15, wherein the scan pulse voltage is at a non-ablative voltage level. | 2021 | Boston Scientific Corporation (BSC) | Medical and Healthcare | Large-Size Company | https://www.bostonscientific.com | USA | The patent focuses on an electroporation ablation system that delivers voltage-controlled pulse sequences in multiple therapy sets via catheter electrodes. It addresses the limitation of fixed-dose electroporation that may not optimize lesion creation across diverse cardiac targets. Voltage-controlled, multi-sequence pulses give finer control over irreversible electroporation effects. | הפטנט מתמקד במערכת אבלציית אלקטרופורציה המספקת רצפי פולסים מבוקרי מתח במספר ערכות טיפול דרך אלקטרודות קטטר. הוא מתמודד עם המגבלה של אלקטרופורציה במינון קבוע שעלולה לא לבצע אופטימיזציה של יצירת נגעים על פני מטרות לבביות מגוונות. פולסים מבוקרי מתח ורב-רצפיים נותנים בקרה עדינה יותר על אפקטי אלקטרופורציה בלתי הפיכה. | 1 | [10 empty] | 1 | [18 empty] | 1 |  | BOSTON SCIENT SCIMED INC | Boston Scientific Corporation (BSC) | BOSTON SCIENT SCIMED INC | Boston Scientific Corporation (BSC) | Boston Scientific Corporation (BSC) | 2022-01-24 | 2022 | 2025-07-01 | 2025 | 2021-01-27 | 2021 | Grant | Utility Invention | active | US202217582733 | 322255308 | CN116744870A¦EP4284276A1¦JP2024504184A¦JP2026009208A¦JP7764483B2¦US20220233237A1¦WO2022164750A1 | US
437 | EP4628019A1 | METHOD OF ANALYZING AN INTRACARDIAC ELECTROGRAM | The invention relates to a method of analyzing an intracardiac electrogram via a control system (1), wherein the intracardiac electrogram comprises a, in particular unipolar, channel with a post-PFA section (8) measured at a location of the human heart after an application of pulsed field ablation, wherein the control system (1) performs a filtering step, wherein in the filtering step the control system (1) applies a filter to the post-PFA section (8) of the unipolar channel (7), thereby generating a filtered channel section (26). It is proposed that the filter has a variable attenuation, thereby removing a variable amount of energy associated with a recurrent noise from a frequency band and leaving a variable amount of energy associated with a biosignal in the frequency band, that the variable amount of energy removed from the unipolar channel (7) is changed over time in a filter update routine, that the control system (1) performs at least one filter update routine resulting in a change of the variable amount of energy removed from the unipolar channel (7) within a time frame after the time of the application of pulsed field ablation and implements the update resulting in a filtering of the unipolar channel (7) within the time frame, and, that the time frame is at most 60 seconds. | 1. Method of analyzing an intracardiac electrogram via a control system (1), wherein the intracardiac electrogram comprises a, in particular unipolar, channel with a post-PFA section (8) measured at a location of the human heart after an application of pulsed field ablation, wherein the control system (1) performs a filtering step, wherein in the filtering step the control system (1) applies a filter to the post-PFA section (8) of the unipolar channel (7), thereby generating a filtered channel section (26),characterized inthat the filter has a variable attenuation, thereby removing a variable amount of energy associated with a recurrent noise from a frequency band and leaving a variable amount of energy associated with a biosignal in the frequency band, that the variable amount of energy removed from the unipolar channel (7) is changed over time in a filter update routine, that the control system (1) performs at least one filter update routine resulting in a change of the variable amount of energy removed from the unipolar channel (7) within a time frame after the time of the application of pulsed field ablation and implements the update resulting in a filtering of the unipolar channel (7) within the time frame, and, that the time frame is at most 60 seconds. 2. Method according to claim 1 or 2, characterized in that the frequency band is a power line frequency band, in particular around a power line frequency of 50 Hz or 60 Hz, preferably, that the frequency band has a width of at most 4 Hz, preferably at most 2 Hz, around the power line frequency. 3. Method according to one of the preceding claims, characterized in that the control system (1) identifies, in particular detects from the intracardiac electrogram, the application of the pulsed field ablation, preferably, that the control system (1) triggers and/reconfigures the filter update routine in the time frame in response to identifying the pulsed field ablation, in particular within 10 seconds of a time of the pulsed field ablation. 4. Method according to one of the preceding claims, characterized in that the biosignal is a local activation (20), and/or, that the location is inside an atrium, in particular at a pulmonary vein entrance (5). 5. Method according to one of the preceding claims, characterized in that in the filter update routine the control system (1) determines an update of the filter by analyzing a current base interval (28) of the channel, preferably, that the base interval (28) is at most 60 seconds long, preferably at most 10 seconds, more preferably at most 1 second, and/or, that a sequence of filter update routines uses a sequence of base intervals (28). 6. Method according to claim 5, characterized in that in the filter update routine in the time frame the control system (1) chooses the base interval (28) such that the base interval (28) lies within the time frame, preferably, that in the filter update routine in the time frame the base interval (28) lies between the application of pulsed field ablation and the fifth, preferably fourth, more preferably third, more preferably second, more preferably first, subsequent occurrence of the biosignal. 7. Method according to claim 5 or 6, characterized in that the control system (1) performs further filter update routines after the application of pulsed field ablation, in particular in the time frame, with base intervals (28) lying between further occurrences of the biosignal, in particular with base intervals (28) lying between at least two subsequent occurrences of the biosignal. 8. Method according to one of claims 5 to 7, characterized in that the filter is a template filter, that in the filter update routine the control system (1) builds or adapts a template (25) of the recurrent noise, in particular of at least one cycle of the recurrent noise, from the base interval (28), and, that the control system (1) removes, in particular subtracts, the built or adapted filter from the channel after the filter update routine. 9. Method according to one of claims 5 to 8, characterized in that the control system (1) performs the filter update routine intermittently with unevenly spaced base intervals (28), preferably, that the control system (1) starts the filter update routines in regular intervals and, based on an analysis of the signal in the base interval (28), aborts some of the filter update routines without building or adapting the template (25). 10. Method according to one of claims 5 to 9, characterized in that the control system (1) builds a template (25) for the recurrent noise based on a base interval (28) lying prior to the application of pulsed field ablation, and, that the control system (1) adapts the template (25) based on a base interval (28) lying after the application of pulsed field ablation, preferably, that the control system (1) does not perform or always aborts the filter update routine during the application of pulsed field ablation. 11. Method according to one of claims 8 to 10, characterized in that the control system (1) safes several templates (25) and updates at least, in particular exactly, one of the saved templates (31) in the filter update routine according to an update criterion, preferably, that the control system (1) changes the update criterion in response to the identification of pulsed field ablation, thereby reconfiguring the filter update routine, and/or, that the control system (1) derives the template (25) removed from the channel from the several saved templates (31), in particular as a mean value of the saved templates (31). 12. Method according to one of the preceding claims, characterized in that the channel is measured in real time, that the control system (1) filters the channel sections in real time as they are measured and strings together the filtered channel sections (26) into a filtered channel (27), preferably, that the control system (1) displays the filtered channel (27) on a user output device (32), in particular display, to a user in real time, and/or, that at least some of the channel sections are filtered with different filters after filter update routines. 13. Method according to one of the preceding claims, characterized in that the control system (1) analyzes the filtered channel section (26) and/or filtered channel (27) and preferably outputs a result of the analysis to a user, preferably, that the control system (1) derives from the filtered channel section (26) and/or filtered channel (27) a residual activation indicator. 14. Method according to one of the preceding claims, characterized in that the channel is measured with an electrode (4) of at least 6 mm2, preferably at least 7 mm2, more preferably at least 8 mm2, surface area, and/or, that the channel is measured from an electrode (4) on a catheter, and, that the pulsed field ablation was or is applied through the same electrode (4) or another electrode (4) on the same catheter. 15. Control system adapted to perform the method according to one of the preceding claims, wherein the control system (1) comprises a processor, a memory, a user output device (32), in particular a display, and a connector (6) for receiving the intracardiac electrogram, in particular for connecting a catheter to the control system (1). | 2024 | CATHVISION | Medical and Healthcare | Small-size company | https://www.cathvision.com | Denmark | The patent focuses on a method for analyzing intracardiac electrograms with a unipolar post-PFA section, where the control system performs filtering on the post-ablation segment. It addresses the difficulty of interpreting post-PFA electrograms cluttered with noise and ablation artifacts. The filtering step extracts clean signal information for assessment of pulsed-field ablation effects. | הפטנט מתמקד בשיטה לניתוח אלקטרוקרדיוגרמות תוך-לבביות עם מקטע מונופולרי לאחר PFA, שבה מערכת הבקרה מבצעת סינון על מקטע לאחר האבלציה. הוא מתמודד עם הקושי בפרשנות אלקטרוקרדיוגרמות לאחר PFA המעורבבות ברעש וארטיפקטים של אבלציה. שלב הסינון מחלץ מידע אות נקי להערכת אפקטי אבלציה בשדה פולסי. | 1 | [10 empty] | 1 | [9 empty] | 1 | [8 empty] | 1 |  | CATHVISION APS | CATHVISION APS | CATHVISION APS | CATHVISION APS | CATHVISION APS | 2024-04-04 | 2024 | 2025-10-08 | 2025 | 2024-04-04 | 2024 | Application | Utility Invention | active | EP20240168561 | 323662772 | WO2025210143A1 | EP
438 | US11903633B2 | Radiofrequency probe, system, and method for adaptive multivariate control of an ablation procedure | The present invention is directed to a system and method for adaptive multivariate control for performing a radiofrequency (RF) ablation procedure with an energy delivery device. The system includes an energy source for delivering energy to a patient's body; one or more energy delivery devices; two or more sensors for measuring at least two factors related to an ablation procedure, respectively; and at least one processor. The method includes steps of: measuring at least two factors related to an ablation procedure; determining a first operating threshold based, at least in part, on a first factor; controlling an energy delivery device based on the first operating threshold to create a lesion at the target site within the patient; determining a second operating threshold based, at least in part, on a second factor; switching control of the energy delivery device from the first factor to the second factor; and controlling the energy delivery device based on the second operating threshold to create a lesion at the target site within the patient. The present invention is also directed to a RF probe configured to be used with the adaptive multivariate control system and method to perform RF ablation procedures. | 1. A radiofrequency (RF) ablation system for performing an RF ablation procedure, the RF ablation system comprising: an energy source for delivering energy to a target site of a patient's body; an energy delivery device electrically coupled to the energy source; two or more sensors for measuring at least two factors, respectively, related to the RF ablation procedure; and at least one processor configured to perform a plurality of operations, the plurality of operations comprising: operating the energy delivery device to deliver energy to create a lesion at the target site within the patient based on a first factor of the at least two factors; while delivering energy to the target site, monitoring the at least two factors related to the RF ablation procedure based on measurements from the two or more sensors; determining whether the first factor of the at least two factors has reached a first operating threshold for the energy delivery device responsive to a determination that measurements of the first factor have stabilized; increasing or decreasing the energy delivered to the target site responsive to a determination that the first factor has not reached the first operating threshold; switching control of the energy delivery device from the first factor to a second factor of the at least two factors once the first operating threshold has been met by the first factor to further develop the lesion at the target site within the patient; determining whether the second factor of the at least two factors has reached a second operating threshold for the energy delivery device responsive to a determination that measurements of the second factor have stabilized; increasing the energy delivered to the target site responsive to a determination that the second factor has not reached the second operating threshold; and terminating the delivery of energy to the target site once the second operating threshold has been met by the second factor. 2. The system of claim 1, wherein the at least two factors comprises at least two of: power, current, voltage, impedance, temperature of tissue at the target site, temperature of a coolant associated with a cooling device for the RF ablation procedure, duration (time) of stimulation, rate of current change, rate of voltage change, rate of power change, rate of temperature change of the tissue at the target site, rate of temperature change of the coolant, pressure at the target site, gas formation at the target site, rate of change of gas formation at the target site, or a combination thereof. 3. The system of claim 2, wherein one of the at least two factors is rate of change of tissue impedance. 4. The system of claim 1, wherein the system is configured to continuously monitor and record the at least two factors. 5. The system of claim 1, wherein the first factor is a tissue temperature at the target site. 6. The system of claim 1, wherein the second factor is tissue impedance or rate of change of tissue impedance at the target site within the patient. 7. The system of claim 1, wherein the system is configured to: determine a third operating threshold for the energy delivery device based, at least in part, on a third factor; and switch control of the energy delivery device from the first operating threshold or the second operating threshold to the third operating threshold to further develop the lesion at the target site. 8. The system of claim 1, wherein at least one sensor of the two or more sensors is configured to measure a respective factor related to the RF ablation procedure at the target site. 9. The system of claim 1, wherein the energy delivery device comprises an RF probe comprising: an elongate member having a proximal end and a distal end, the distal end comprising an electrically and thermally-conductive energy delivery device for delivering one of electrical and RF energy to a patient's body, the RF probe having at least one electrode and an electrically and thermally-conductive portion having a temperature sensing element; and one or more ablation enhancing features comprising: a cooling device extending within the elongate member, a suction mechanism having a suction path, a fluid injection mechanism having a fluid path, and/or a directional energy delivery mechanism. 10. The system of claim 9, wherein the two or more sensors comprises, at least, a thermocouple positioned on the RF probe. 11. The system of claim 1, wherein control of the energy delivery device is switched from the first factor to the second factor responsive further to a determination at that the measurements of the second factor have stabilized. 12. The system of claim 1, the plurality of operations further comprising ending the RF ablation procedure by terminating the delivery of output power to the one or more energy delivery devices once the second operating threshold has been met. 13. A method for performing a radiofrequency (RF) ablation procedure with an RF probe, the method comprising: operating the RF probe to deliver energy to create a lesion at a target site within a patient based on a first operational factor related to an ablation procedure performed with the RF probe; while delivering energy to the target site, monitoring the first operational factor and a second operational factor related to the ablation procedure that is different from the first operational factor, wherein the first and second operational factors are monitored using measurements from sensors configured to measure the first operational factor and the second operational factor, respectively; determining whether the first operational factor has reached a first operating threshold for the RF probe responsive to a determination that measurements of the first operational factor have stabilized; increasing or decreasing the energy delivered by the RF probe responsive to a determination that the first operational factor has not reached the first operating threshold; switching control of the RF probe from the first operational factor to the second operational factor once the first operating threshold has been met by the first operational factor to further develop the lesion at the target site within the patient; determining whether the second operational factor has reached a second operating threshold responsive to a determination that measurements of the second operational factor have stabilized; increasing the energy delivered to the target site responsive to a determination that the second operational factor has not reached the second operating threshold; and terminating the delivery of energy to the target site once the second operating threshold has been met by the second operational factor. 14. The method of claim 13, further comprising steps of: collecting data relating to operation of the RF probe during use; inputting the collected data into a deep learning network to train the deep learning network to control the energy delivered to the target site based on the first operational factor and the second operational factor; and controlling the RF probe based on the deep learning network. | 2020 | Philips¦Avent | Medical and Healthcare | Large-Size Company | https://www.philips.com | Netherlands | The patent focuses on an RF probe with adaptive multivariate control, using two or more sensors to measure multiple ablation factors and an energy source feeding the energy delivery device. It addresses the limitation of single-variable RF ablation control that can't simultaneously balance multiple procedure factors. Multivariate adaptive control improves ablation precision and safety. | הפטנט מתמקד בגשוש RF עם בקרה רב-משתנית מסתגלת, המשתמש בשני חיישנים או יותר למדידת מספר גורמי אבלציה ובמקור אנרגיה המזין את מכשיר מסירת האנרגיה. הוא מתמודד עם המגבלה של בקרת אבלציית RF משתנית-בודדת שאינה יכולה לאזן בו זמנית מספר גורמי פרוצדורה. בקרה רב-משתנית מסתגלת משפרת דיוק ובטיחות אבלציה. | 1 |  |  |  | 1 | [5 empty] | 1 | [10 empty] | 1 | [8 empty] | 1 |  | AVENT INC | Avent | AVENT, INC. | Avent | Philips | 2020-12-30 | 2020 | 2024-02-20 | 2024 | 2020-12-30 | 2020 | Grant | Utility Invention | active | US202017137810 | 283145034 | AU2021413865A1¦AU2021413865A9¦EP4271293A1¦MX2023007929A¦US20220202471A1¦WO2022147198A1 | US
439 | WO2025200114A1 | RADIOFREQUENCY ABLATION APPARATUS AND RADIOFREQUENCY ABLATION CATHETER | Disclosed are a radiofrequency ablation apparatus and a radiofrequency ablation catheter. The apparatus comprises: a catheter body, a handle, a connection cable, and a connector which are sequentially connected from the distal end to the proximal end, wherein the catheter body is, from the distal end, sequentially provided with a rubber head, a heating section, and a main body catheter. The surface of the heating section is provided with an insulation outer sleeve with insulating and smoothing effects, and the heating section is internally provided with coils that are used for heating and are wound from alloy wires. The coils comprise a proximal end ablation coil and a distal end ablation coil. The ablation coils formed by winding different alloy wires are wound from the same start point or wound in an "end-to-end connection" manner; alloy wires with the same start point have opposite alloy wire winding directions, and alloy wires with "connected start and end points" have the same alloy wire winding direction. The radiofrequency ablation apparatus of the present disclosure can achieve radiofrequency closure of varicose veins at different diseased regions without the need to change instruments, avoiding the problem of uneven heating caused by uneven distribution of alloy wires between different coils and thereby preventing related surgical risks. | 1. A radiofrequency ablation catheter, characterized by comprising: Coil inner cannula, ablation coil; The ablation coil includes a proximal ablation coil and a distal ablation coil, wherein the proximal ablation coil and the distal ablation coil are respectively located at the proximal end and the distal end of the coil inner sleeve; The proximal ablation coil and the distal ablation coil are both formed by spirally winding an alloy wire on the outer surface of the inner sleeve of the coil; The proximal ablation coil and the distal ablation coil are formed by different alloy wires being wound from their respective starting points in opposite winding directions toward the proximal end and the distal end respectively, and the distance between the two starting points is 0-3mm; or the proximal ablation coil and the distal ablation coil are formed by different alloy wires being wound from their respective starting points in the same winding direction toward the proximal end, and the distance between the midpoint of the line connecting the two winding end points of the distal ablation coil and the winding starting point of the proximal ablation coil is 0-3mm. 2. The radiofrequency ablation catheter according to claim 1 is characterized in that the alloy wires wound on the outer surface of the coil inner sleeve are evenly distributed. 3. The radiofrequency ablation catheter according to claim 1 is characterized in that the two starting points are arranged on the same cross-section or the starting point and the midpoint of the proximal ablation coil are arranged on the same cross-section, and the cross-section is the cross-section of the coil inner sheath. 4. A radiofrequency ablation catheter as described in claim 1 is characterized in that it also includes a wire 1, and a threading hole is provided on the surface of the coil inner sleeve near the winding end point of the proximal ablation coil. The wire 1 is connected to the alloy wire of the proximal ablation coil and then enters the interior of the coil inner sleeve through the threading hole. 5. The radiofrequency ablation catheter according to claim 1 is characterized in that it also includes a second wire, and a threading hole is provided on the surface of the coil inner sleeve near the winding end point of the distal ablation coil. The second wire is connected to the alloy wire of the distal ablation coil and then enters the interior of the coil inner sleeve through the threading hole. 6. The radiofrequency ablation catheter according to claim 1, wherein the proximal ablation coil is formed by winding an alloy wire, and the distal ablation coil is formed by winding an alloy wire. 7. The radiofrequency ablation catheter according to claim 1 is characterized in that the proximal ablation coil is formed by winding two alloy wires, and the two alloy wires respectively form a first proximal ablation coil and a second proximal ablation coil. 8. A radiofrequency ablation catheter as described in claim 7 is characterized in that a threading hole is provided on the surface of the coil inner sleeve at the end position of the alloy wire winding corresponding to the first proximal ablation coil and the second proximal ablation coil, and the alloy wire is connected to the wire before entering the threading hole to prevent the connection point between the alloy wire and the wire from being located inside the coil inner sleeve. 9. The radiofrequency ablation catheter according to claim 1, characterized in that at least two temperature measuring devices are provided on the proximal ablation coil. 10. The radiofrequency ablation catheter according to claim 1 is characterized in that an insulating outer sleeve is provided on the outer surface of the ablation coil, and the insulating outer sleeve is tightly attached to the surface of the ablation coil by high-temperature heat shrinkage. 11. A radiofrequency ablation device, characterized by comprising the radiofrequency ablation catheter according to any one of claims 1 to 10. 12. A radiofrequency ablation device as described in claim 11 is characterized in that it also includes an ablation start and stop button, a radiofrequency closing generator, and a temperature control controller; the radiofrequency ablation device includes multiple working modes; different working modes correspond to the working modes when one or more alloy wires are turned on and off. 13. A radiofrequency ablation device, characterized in that it comprises the radiofrequency ablation catheter according to claim 6, an ablation start/stop button, a radiofrequency closure generator, and a temperature control controller; the radiofrequency ablation device comprises a first operating mode and a second operating mode; In the first operating mode, the RF closure generator starts working by pressing the ablation start/stop button, and simultaneously outputs high-frequency current to the proximal and distal ablation coils. The two coils heat up under the action of the high-frequency current. At the same time, the first, second, and third thermocouples all monitor the temperature changes of the vascular tissue. The three thermocouples transmit the temperature signal to the temperature controller, which controls the RF closure generator to control the current of the two coils respectively. The temperature value will also be displayed in real time on the RF closure generator screen. In the second working mode, press the ablation button and the RF closure generator starts working, while outputting high-frequency current to the distal ablation coil. The distal ablation coil generates heat under the action of the high-frequency current. At the same time, the third thermocouple monitors the temperature changes of the vascular tissue. The third thermocouple transmits the temperature signal to the temperature controller and controls the current of the distal coil by controlling the RF closure generator. The temperature value will also be displayed in real time on the RF closure generator screen. 14. A radiofrequency ablation device according to claim 13, characterized in that it further comprises a third operating mode; In the third working mode, press the ablation start and stop button, the RF closure generator starts working, and outputs high-frequency current to the proximal ablation coil. The proximal ablation coil generates heat under the action of the high-frequency current. At the same time, the first thermocouple and the second thermocouple monitor the changes in the temperature of the vascular tissue. The first thermocouple and the second thermocouple transmit the temperature signal to the temperature controller and control the current of the proximal coil by controlling the RF closure generator. The temperature value will also be displayed in real time on the RF closure generator screen. | 2024 | Jiangsu Hengrui Medicine | Medical and Healthcare | NA | https://www.hengrui.com/en/about.html | China | The patent focuses on an RF ablation apparatus and catheter with a catheter body, handle, connection cable, and connector, plus a rubber head, heating section, and main body catheter. It addresses the design challenges in RF ablation catheters that need integrated heating and reliable mechanical handling. The unified construction enables effective RF ablation in clinical settings. | הפטנט מתמקד במכשיר אבלציית RF וקטטר עם גוף קטטר, ידית, כבל חיבור ומחבר, בתוספת ראש גומי, מקטע חימום וקטטר גוף ראשי. הוא מתמודד עם אתגרי העיצוב בקטטרי אבלציית RF הזקוקים לחימום משולב וטיפול מכני אמין. הבנייה המאוחדת מאפשרת אבלציית RF יעילה בהגדרות קליניות. |  |  |  |  | 1 | [5 empty] | 1 | [10 empty] | 1 | [8 empty] | 1 |  | SUZHOU HENGRUI HONGYUAN MEDICAL TECH CO LTD | Jiangsu Hengrui Medicine | SUZHOU HENGRUI HONGYUAN MEDICAL TECH CO LTD | Jiangsu Hengrui Medicine | Jiangsu Hengrui Medicine | 2024-05-27 | 2024 | 2025-10-02 | 2025 | 2024-03-29 | 2024 | Application | Utility Invention | active | WO2024CN95542 | 310823435 | CN118078418A¦CN306171770S | WO
440 | WO2025146190A1 | ULTRASONIC ABLATION DEVICE AND SYSTEM | Disclosed in the present invention are an ultrasonic ablation device and system. The ultrasonic ablation device comprises: a control part, used for automatic control of ultrasonic ablation; a catheter, connected to the control part and at least comprises a water inlet cavity, a water outlet cavity, and a wire cavity; an ultrasonic transducer, arranged at the distal end of the catheter and connected to the control part so as to be switched between an ablation mode and a measurement mode, wherein the ultrasonic transducer is used for ultrasonic ablation in the ablation mode, and is used for temperature measurement in the measurement mode; a flexible ablation balloon, sleeved outside the ultrasonic transducer; a flexible electrode, mounted on the outer side of the balloon and used for measuring impedance; and a pressure sensor, mounted on the inner side of the balloon and used for measuring water pressure, wherein the control part is used for receiving the impedance, the water pressure, and the temperature so as to determine, on the basis of a first determination logic, whether the flexible ablation balloon is completely attached to a vessel wall; on the basis of a switching logic, controlling the ultrasonic transducer to perform mode switching; and on the basis of a second determination logic, determining whether an ablation end point is reached. | 1. An ultrasonic ablation device, characterized by comprising: A control unit, used for automatic control of ultrasonic ablation; A catheter, the proximal end of which is connected to the control unit, and the catheter at least comprises a water inlet cavity, a water outlet cavity, and a wire cavity through which a wire is passed; An ultrasonic transducer is disposed at the distal end of the catheter and is electrically connected to the control unit through the wire, so as to switch between ; wherein, in the ablation mode, the ultrasonic transducer is used to perform ultrasonic ablation on the area to be ablated; and in the measurement mode, the ultrasonic transducer is used to measure the temperature and tissue morphology of the area to be ablated; A flexible ablation balloon, which is sleeved on the outside of the ultrasonic transducer and is connected to the water inlet cavity and the water outlet cavity; A flexible electrode, mounted on the outside of the flexible ablation balloon and electrically connected to the control unit, for detecting impedance; A pressure sensor, installed on the inner side of the flexible ablation balloon and electrically connected to the control unit, for detecting water pressure; Wherein, the control unit is used to receive the impedance and water pressure, and judge whether the flexible ablation balloon is completely attached to the wall according to the first judgment logic; The control unit is also used to receive the impedance, water pressure, and the temperature and tissue morphology of the area to be ablated, and control the ultrasonic transducer to switch between the ablation mode and the measurement mode according to the switching logic; The control unit is also used to receive the impedance, water pressure, and the temperature and tissue morphology of the area to be ablated, and determine whether the ablation endpoint is reached according to the second judgment logic. 2. The ultrasonic ablation device according to claim 1, characterized in that the first judgment logic specifically includes: If the impedance value reaches a preset impedance range, and the water pressure reaches a preset water pressure range, it is determined that the flexible ablation balloon is completely attached to the wall; If the impedance value does not reach a preset impedance range, and/or the water pressure does not reach a preset water pressure range, it is determined that the flexible ablation balloon is not completely attached to the wall. 3. The ultrasonic ablation device according to claim 1, characterized in that the switching logic specifically includes: The preset duration is used as the switching condition, and the switching is performed every preset duration; Alternatively, the control unit receives the impedance, water pressure, and temperature and tissue morphology of the area to be ablated, and outputs different powers according to the difference between the temperature of the area to be ablated and a preset temperature; wherein, when it is necessary to calculate the output power, the ultrasonic transducer switches to a measurement mode to measure the temperature of the area to be ablated, thereby using the control unit to calculate the output power; when the output power calculation is completed, the ultrasonic transducer automatically switches to an ablation mode and continues for a preset time until an ultrasonic ablation is completed, and automatically switches to a measurement mode after an ultrasonic ablation is completed to calculate the next output power; the ultrasonic transducer repeatedly switches between the measurement mode and the ablation mode until the ablation end point is reached. 4. The ultrasonic ablation device according to claim 3, characterized in that the second judgment logic specifically includes: When the temperature of the area to be ablated reaches the preset temperature and remains stable for a preset time, and the water pressure and impedance drop to the theoretical value range, it is determined that the ablation end point has been reached. 5. The ultrasonic ablation device according to claim 1, characterized in that the control unit at least comprises: The main control module is used to process signals and data and control automated ablation; A signal generator, connected to the main control module, for receiving and sending ultrasonic signals; A power amplifier connected to the signal generator to amplify the ultrasonic signal; A gating module connected to the power amplifier to select the number of signal channels; A phase voltage and current module, connected to the main control module, for detecting parameters and outputting a matching ultrasonic frequency; The impedance matching module is connected to the main control module and is used to detect the matching impedance and output an ultrasonic signal of matching power. 6. The ultrasonic ablation device according to claim 5, characterized in that: The ultrasonic transducer is connected to the gating module to switch between the ablation mode and the measurement mode according to the number of signal channels; In the ablation mode, the ultrasonic transducer converts the electrical signal into mechanical vibration to continuously generate ultrasonic waves and heats the area to be ablated within the energy focus range of the sound wave transmission; In the measurement mode, the ultrasonic transducer converts the pulsed electrical signal into a mechanical signal to generate a single ultrasonic wave, and calculates the tissue morphology change and temperature change by analyzing the echo signal. 7. The ultrasonic ablation device according to claim 1, characterized in that it also includes: A display connected to the control unit for displaying data; The alarm is connected to the control unit to sound an alarm when the data is abnormal. 8. The ultrasonic ablation device according to claim 1, characterized in that: The catheter also has a guidewire cavity, in which a guidewire is passed to control the bending direction of the catheter. 9. The ultrasonic ablation device according to claim 1, characterized in that: The flexible ablation balloon is made of silicone, and the flexible electrode is a flexible circuit board. 10. An ultrasonic ablation system, characterized in that it comprises the ultrasonic ablation device as described in any one of claims 1 to 9. | 2024 | SHANGHAI GOLDEN LEAF MED TEC | Medical and Healthcare | NA | https://www.brattea.com/en/aboutUsEn.html | China | The patent focuses on an ultrasonic ablation device with a control part for automatic ultrasound control, a multi-cavity catheter (water inlet, water outlet, wire), and a distal-end ultrasonic transducer. It addresses the workflow inefficiency of manual ultrasonic ablation and limited integration of fluid management. Automatic control combined with cooling/wire pathways simplifies and improves ultrasound ablation. | הפטנט מתמקד במכשיר אבלציה אולטרסונית עם חלק בקרה לבקרת אולטרסאונד אוטומטית, קטטר רב-חללי (כניסת מים, יציאת מים, חוט) ומתמר אולטרסוני בקצה הדיסטלי. הוא מתמודד עם חוסר היעילות של זרימת עבודה ידנית של אבלציה אולטרסונית והאינטגרציה המוגבלת של ניהול נוזלים. בקרה אוטומטית בשילוב עם נתיבי קירור/חוט מפשטת ומשפרת אבלציה אולטרסאונדית. | 1 |  | 1 | [12 empty] | 1 | [14 empty] | 1 |  | SHANGHAI GOLDEN LEAF MED TEC CO LTD | SHANGHAI GOLDEN LEAF MED TEC CO LTD | SHANGHAI GOLDEN LEAF MED TEC CO LTD | SHANGHAI GOLDEN LEAF MED TEC CO LTD | SHANGHAI GOLDEN LEAF MED TEC CO LTD | 2025-01-05 | 2025 | 2025-07-10 | 2025 | 2024-01-05 | 2024 | Application | Utility Invention | active | WO2025CN70629 | 322382071 | CN117883723A | WO
441 | WO2022125933A1 | ROBOTIC SURGERY | A laser or ultrasonic instrument is used to remove tissue during a surgery, such as to form one or more pilot holes in a vertebra or a window in bone. Where a laser is used, interrogative laser pulses can be used to obtain information, such as detecting depth or tissue type. | 1. CLAIMSWe claim: 1. A method for performing spinal surgery, the method comprising: exposing (302) a vertebra (10) of a patient; disposing (304) a distal end (204) of a laser instrument (200) proximate the vertebra (10); conducting (306) a laser-based topographical analysis of the vertebra using the laser instrument (200); registering (310) a patient’s anatomy using the laser-based topographical analysis; and creating (320) a pilot hole (12) in the vertebra (10) with the laser instrument (200). 2. The method of claim 1, further comprising: determining (312) that a spine of the patient shifted since a prior non- laser-based registration; and wherein the registering of the patient’s anatomy includes: updating (314) an existing non- laser-based registration using the laser-based topographical analysis. 3. The method of claim 1 or claim 2, wherein creating (320) a pilot hole (12) in the vertebra (10) with the laser instrument (200) includes creating a multi-diameter pilot hole (12). 4. The method of claim 3, wherein creating the multi-diameter pilot hole (12) includes: creating (322) a multiple diameter pilot hole (12) having a countersink diameter and a minor diameter interference. 5. The method of any one of claims claim 1-4, further comprising: placing (331) a navigated instrument (170) or an implant (180) engaged with the navigated instrument (170) into the pilot hole (12). 6. The method of any one of claims 1-5, further comprising: conducting (332) a spectral analysis of laser reflection received through the laser instrument to determine laser characteristics to be used to complete a surgical plan. 7. The method of any one of claims 1-6, further comprising: registering a patient’s anatomy using a non- laser registration technique to produce a non- laser registration; conducting (310) a laser-based topographical analysis of the vertebra; and calculating a registration confidence value for the registration using the laser-based topographical analysis. 8. A method comprising: removing (410) a portion of facet joint tissue of a vertebra (10) using laser energy transmitted from a distal end (204) of a laser instrument (200), wherein the removing includes starting (412) a laser cut on an outside of the facet and angling the transmission of the laser energy back into bone. 9. The method of claim 8, wherein the laser instrument (200) comprises an elongate shaft (210) that defines a longitudinal axis (206); wherein the laser instrument (200) defines a laser pulse axis (208) along which the laser instrument (200) is configured to direct laser pulses; and wherein the laser pulse axis (208) is non-parallel with respect to the longitudinal axis (206). 10. The method of claim 8 or claim 9, further comprising: using (402) a tissue dilator to expose spine tissue (10). 11. The method of any one of claims 8-10, further comprising: using (422) laser topography, laser reflection, laser refraction, or optical coherence tomography to confirm when a cut is complete. 12. The method of any one of claims 8-11, wherein starting (412) the laser cut on an outside of the facet and angling the transmission of laser energy back into bone includes using a mirror or angled laser fiber to perform the angling. 13. The method of any one of claims 8-12, wherein starting the laser cut on an outside of the facet and angling the transmission of laser energy back into bone includes angling the transmission of laser energy in a direction other than toward a disc or nerve. 14. The method of any one of claims 8-13, further comprising: extending (414) the laser cut to undermine a spinous process proximate the facet joint and continuing the laser cut along a contralateral foraminal recess. 15. The method of any one of claims 8-14, further comprising: checking (404) for range of motion limits that may dictate a patient re-orientation or a different tip with more angulation. 16. The method of any one of claims 8-15, further comprising tuning (406) a laser generator (112) of the laser instrument (200) such that generated laser pulses (30) reach a frequency selected to cut a particular kind of tissue. 17. The method of any one of claims 8-16, further comprising tracking (416) the removing through laser topography, reflected laser energy, refracted laser energy, or optical coherence tomography to determine removed areas. 18. A method comprising: exposing (302) a vertebra (10) of a patient; disposing (304) a distal portion of a laser instrument (200) proximate the vertebra; and creating (530) a pilot hole (12) in the vertebra (10), wherein the creating (530) of the pilot hole (12) includes repeatedly (534) changing a direction of laser pulses (20, 30) emitted from the laser instrument (200). 19. The method of claim 18, wherein changing the direction of the laser pulses (20, 30) emitted from the laser instrument (200) includes: changing an angle at which laser pulses (20, 30) leave the laser instrument (200) by changing an angle of a light director (290) of the laser instrument (200); or changing an angle of a longitudinal axis (206) of the laser instrument (200) from a first angle to a second angle. The method of claim 18 or claim 19, wherein creating the pilot hole (12) includes: advancing the laser instrument (200) toward the vertebra (10) after removing a first amount of tissue from the vertebra (10) and before removing a second amount of tissue from the vertebra (10). The method of any one of claims 18-20, further comprising: after creating the pilot hole (12), measuring (536) an impedance related to tissue defining a bottom of the pilot hole (12) to determine whether the tissue is bone. An apparatus comprising: a laser instrument (200) configured to selectively emit: a first laser pulses (20) at a wavelength selected to affect disc tissue or cartilage tissue without ablating bone; and a second laser pulses (30) configured to interrogate a target region; a light detector (114) configured to detect the second laser pulses (30); an irrigator (116) configured to provide a fluid proximate tissue affected by the laser instrument (200) to provide cooling and a medium to evacuate debris; and a suction generator (118) configured to remove the fluid. The apparatus of claim 22, further comprising: a computer (120) coupled to the light detector (114) and having one or more processors (122) configured to: perform topographical analysis or spectral analysis based on the detected second laser pulses (30). The apparatus of claim 23, further comprising: a laser generator (112) configured to generate the first laser pulses (20) according to parameters, wherein the computer (120) is further configured to modify the parameters based on the topographical analysis or spectral analysis. 25. The apparatus of any one of claims 22-24, further comprising: a reference array (252) coupled to the laser instrument (200) and configured to permit the tracking of a location of the laser instrument (200). 26. The apparatus of claim 25, wherein the reference array includes two or more tracking fiducials (254). 27. The apparatus of any one of claims 22-26, further comprising a coupling (134) configured to couple the laser instrument to a robot (130). 28. A method comprising: receiving (610) a selection of a tissue type over a user interface (150); providing (620) an interrogation laser pulse (30) to tissue; receiving (630) the interrogation laser pulse (30) from the tissue; determining (640) an interrogated tissue type based on the received interrogation laser pulse (30); and providing (650) a removal laser pulse (20) to the tissue responsive to the interrogated tissue type matching the selected tissue type. 29. The method of claim 28, further comprising: determining (660) an interrogated tissue thickness based on the received interrogation laser pulse; and setting (670) one or more parameters based on the determined interrogated tissue thickness; and generating (680) the removal laser pulse according to the one or more parameters. 30. The method of claim 28, wherein the tissue type is ligament tissue, disc tissue, tumor tissue, or bone tissue. 31. A method comprising: disposing (710) a distal portion of a laser instrument (200) proximate a target region of bone of a patient where a pilot hole is to be created; and forming (750) a pilot hole in the target region, wherein the forming includes applying laser energy from the laser instrument (200) to the target region. 32. The method of claim 31, wherein applying the laser energy includes to apply the laser energy in a sub-ablative manner with respect to the bone. 33. The method of claim 31 or claim 32, further comprising: planning (702) a procedure, wherein planning the procedure includes: receiving an indication of a location of the target region; receiving one or more characteristics of a screw to be disposed in the pilot hole; determining a characteristic of the pilot hole based on the one or more characteristics of the screw; and receiving confirmation of the determined characteristic of the pilot hole. 34. The method of claim 33, wherein the target region is a region of a vertebra (10); and wherein the method further comprises: removing (704) tissue proximate the target region such that a screw driver can reach a starting point of the screw and a posterior cortex of a spinal vertebra. 35. The method of any one of claims 31-34, further comprising: obtaining (720) a screw (810) defining a screw cannula (824) and a screw distal end (814); obtaining (722) a screw driver (820) defining a driver cannula (852); placing (724) the laser instrument (200) through the driver cannula (852); and placing (726) the laser instrument (200) into the screw cannula (824) such that the distal portion of the laser instrument is proximate the screw distal end (814). The method of claim 35, further comprising: inserting (728) the screw driver through a guide (136) of a robot arm (132); and with a robot arm (132), disposing (730) the guide (136) proximate the target region. The method of claim 36, further comprising: after forming (750) the pilot hole in the target region, driving (790) the screw into the pilot hole using the screw driver. The method of any one of claims 35-37, further comprising: while forming a pilot hole in the target region: moving a robot arm under a surgeon’s control or automatically according to a plan. The method of any one of claims 31-38, wherein the forming includes: calculating (756) a depth of the pilot hole using the laser instrument. The method of any one of claims 31-39, wherein the laser energy is in a spectrum associated with Er:YAG; wherein the method further comprises applying cooling using a cooling lumen (216); wherein the method further comprises applying suction with a suction lumen (218); or wherein the method further comprises tracking a location with a navigation lumen (212). A method (900) comprising: forming (902) percutaneous path through tissue to a target region of a vertebra (10); inserting (904) a laser instrument (200) into the percutaneous path such that a distal end (204) of the laser instrument (200) is disposed proximate the target region; and forming (750) a pilot hole (12) at the target region, wherein the forming includes: applying (752) laser energy from the laser instrument (200) to tissue of the target region. The method of claim 41, further comprising: applying (764) laser energy from the laser instrument (200) to remove soft tissue proximate the target region. 43. The method of claim 41 or claim 42, wherein forming the pilot hole (12) at the target region includes: detecting (766) a tissue type of tissue to which the laser energy is to be applied; and modifying (768) a laser energy level based on the tissue type. 44. The method of claim 43, wherein detecting (766) the tissue type includes: performing optical coherence tomography; or performing spectroscopy. 45. The method of any one of claims 41-44, wherein applying (752) the laser energy includes applying the laser energy in a manner that is sub-ablative with respect to bone. 46. The method of any one of claims 41-45, wherein applying (752) the laser energy includes applying the laser energy in a manner that results in vaporizing water. 47. The method of any one of claims 41-46, further comprising: after forming the pilot hole, driving (790) a screw into the pilot hole. 48. The method of any one of claims 41-47, further comprising: performing (906) a topographical scan of the target region; determine (908) whether the topographical scan sufficiently matches an expected result; and responsive to the topographical scan fails to sufficiently match the expected result, determining (910) an appropriate offset transformation. 49. An apparatus (800) comprising: a screw (810) comprising a screw cannula (834), a screw head (816), and a distal tip (814); a screw driver (820) comprising a drive tip (826) and a driver cannula (852), wherein the drive tip (826) is engaged with the screw head (816); and a laser instrument (200) extending through the screw cannula (834) and the driver cannula (852), the laser instrument having a distal end (204) disposed proximate the distal tip (814) of the screw (810). 50. The apparatus of claim 49, wherein the laser instrument (200) comprises an adapter (224) configured to optically couple a first laser conduit (222) comprising silica glass to a second laser conduit (226) and shift a wavelength of light entering the adapter (224) from the first laser conduit (222) from a wavelength of less than 2700 nanometers to a wavelength of greater than 2900 nanometers. 51. A method comprising: determining (1210) a desired spinal implant geometry for a patient; obtaining (1220) an implant blank (1102); intraoperatively shaping (1230) the implant blank with a laser (200); and implanting (1240) the shaped implant in the patient. 52. The method of claim 51 , wherein the implant blank (1102) is bone of the patient; and wherein obtaining (1220) the implant blank includes cutting the implant blank from patient with the laser. 53. The method of claim 51 or claim 52, further comprising: creating a pilot hole in a vertebra of the patient with the laser. 54. The method of any one of claims 51-53, wherein intraoperatively shaping (1230) the implant blank with the laser includes applying laser energy is in a spectrum associated with Er:YAG from the laser to the implant blank. 55. The method of any one of claims 51-54, wherein implanting (1240) the shaped implant in the patient includes implanting the shaped implant in an intervertebral space of the patient. A method comprising: obtaining (720) a screw (810) defining a screw cannula (824) and a screw distal end (814); obtaining (722) a screw driver (820) defining a driver cannula (852); placing (724) a distal end of an ultrasonic cutting probe (1312) through the driver cannula (852); placing (726) the distal end of the ultrasonic cutting probe (1312) into the screw cannula (824) such that the distal end of the ultrasonic cutting probe (1312) is proximate the screw distal end (814); disposing (710) the distal end of an ultrasonic cutting probe (1312) proximate a target region of bone of a patient where a pilot hole is to be created; and forming (750) a pilot hole in the target region, wherein the forming includes applying ultrasonic cutting energy from the ultrasonic cutting probe (1312) to the target region while the ultrasonic cutting probe (1312) extends through the screw cannula (824). The method of claim 56, wherein the target region is a region of a vertebra (10); and wherein the method further comprises: removing (704) tissue proximate the target region such that the screw driver can reach a starting point of the screw and a posterior cortex of a spinal vertebra. The method of claim 56 or claim 57, further comprising: inserting (728) the screw driver through a guide (136) of a robot arm (132); and with a robot arm (132), disposing (730) the guide (136) proximate the target region. The method of claim 58, further comprising: after forming (750) the pilot hole in the target region, driving (790) the screw into the pilot hole using the screw driver. The method of claim 58 or claim 59, further comprising: while forming a pilot hole in the target region: moving the robot arm under a surgeon’s control or automatically according to a plan. 61. A method (900) comprising: forming (902) a path through tissue to a target region of a bone (1410); disposing (1510) a cutting apparatus proximate the target region; cutting (1520) a window (1420) in the bone (1410) with the cutting apparatus; removing (1530) a boney plug (1430) from the window (1420); performing (1540) a procedure through the window (1420); and replacing (1550) the boney plug (1430) in the window (1420). 62. The method of claim 61, wherein the cutting apparatus is a laser instrument (200); and wherein cutting (1520) the window (1420) includes applying laser energy to the bone (1410). 63. The method of claim 62, wherein the laser energy is in a spectrum associated with Er:YAG. 64. The method of claim 62 or claim 63, wherein applying laser energy to the bone (1410) is sub-ablative with respect to bone. 65. The method of any one of claims 62-64, wherein applying (752) the laser energy includes applying the laser energy in a manner that results in vaporizing water. 66. The method of any one of claims 61-65, wherein the cutting apparatus is an ultrasonic cutting probe. 67. The method of any one of claims 61-66, wherein performing (1540) a procedure through the window (1420) includes: decompressing tissue through the window. 68. The method of any one of claims 61-67,60 wherein the path is an anterior path; and wherein the bone (1410) is a vertebral body of a vertebra. The method of any one of claims 61-68, wherein the path is a lateral path; and wherein the bone (1410) is an iliac crest. 61 | 2020 | NuVasive | Medical and Healthcare | Large-Size Company | https://www.nuvasive.com/ | USA | The patent focuses on robotic surgery using a laser or ultrasonic instrument to remove tissue, including forming pilot holes in vertebra or bone windows, with interrogative laser pulses for depth and tissue-type detection. It addresses the manual variability of bone-removal surgery and the lack of integrated tissue sensing. The robot plus interrogative pulses enables more precise and informed bone procedures. | הפטנט מתמקד בניתוח רובוטי המשתמש בכלי לייזר או אולטרסוני להסרת רקמה, כולל יצירת חורי הנחיה בחוליה או חלונות עצם, עם פולסי לייזר חוקרים לזיהוי עומק וסוג רקמה. הוא מתמודד עם השונות הידנית של ניתוח להסרת עצם והחוסר בחישת רקמה משולבת. הרובוט בתוספת פולסים חוקרים מאפשר פרוצדורות עצם מדויקות ומיודעות יותר. | 1 | 1 | [11 empty] | 1 | [9 empty] | 1 | [7 empty] | 1 | NUVASIVE INC | NuVasive | NUVASIVE INC | NuVasive | NuVasive | 2021-12-10 | 2021 | 2022-06-16 | 2022 | 2020-12-11 | 2020 | Application | Utility Invention | inactive | WO2021US62862 | 314434859 | EP4259028A1¦EP4259028A4¦US11980415B2¦US20220183755A1¦US20240277409A1 | WO
442 | WO2025157043A1 | METHOD FOR DETERMINING ABLATION ENDPOINT OF INTRAVASCULAR NERVE ABLATION AND ABLATION DEVICE | A method for determining an ablation endpoint of an intravascular nerve ablation, and an ablation device. The determination method comprises the following steps: setting target values and target change ranges of operating parameters, wherein the operating parameters at least include power, temperature, impedance, related operating parameters, ablation duration and the like; sending an ablation catheter to a target ablation position and confirming good electrode contact; periodically collecting operating parameter data at each electrode at set time intervals; and performing corresponding calculation processing on each round of collected data to obtain related processing data such as power, temperature, impedance difference, growth rate, maximum value, minimum value, maximum fluctuation range, and standard deviation; and comprehensively determining an ablation endpoint on the basis of collected data of the power, temperature, impedance, and ablation duration, and processed data of the power, temperature, and impedance. | 1. A method for determining the ablation endpoint of intravascular nerve ablation, characterized by comprising the following steps: (1) Setting target values and target variation ranges of operating parameters, the operating parameters including at least power, temperature, impedance, and ablation duration, as well as related operating parameters of the power, temperature, and impedance, and weights of the operating parameters, as well as an overall effective value threshold, a pass number threshold, and a continuous pass number threshold; (2) Send the ablation catheter to the target ablation site and confirm that the electrode is well attached to the wall; (3) cyclically collecting working parameter data at each electrode at set time intervals, including the power, temperature, impedance, and ablation duration data; (4) Starting from the second round of data collection, corresponding calculation processing is performed on each round of data collection to obtain relevant processed data of the power, temperature, and impedance at each electrode; the relevant processed data includes the difference, maximum value, minimum value, maximum fluctuation range and standard deviation of the power, temperature, and impedance, as well as the growth rate when the power, temperature, and impedance all reach the set target variation range for the first time; (5) Determine the ablation endpoint based on the collected data of the power, the temperature, the impedance, and the ablation duration, and the processed data of the power, the temperature, and the impedance. 2. The method for determining the ablation endpoint of intravascular nerve ablation as described in claim 1 is characterized in that in step (1), the relevant working parameters include the difference and growth rate between the radio frequency power, the ablation temperature, and the tissue impedance in two consecutive acquisitions of data, and the maximum value, minimum value, and standard deviation value of the radio frequency power, the ablation temperature, and the tissue impedance in the current acquisition data. 3. The method for determining the ablation endpoint of intravascular nerve ablation according to claim 1, wherein step (3) comprises the following sub-steps: (31) In the monitoring mode, collecting the impedance data between each electrode and the neutral electrode on the body surface, collecting the impedance data between every two electrodes, and collecting the temperature data at each electrode; (32) In radio frequency mode, the power data and the ablation duration data output to each electrode are collected. 4. The method for determining the ablation endpoint of intravascular nerve ablation according to claim 1 is characterized in that in step (4), the power difference Pd between each round of collected data and the previous round of collected data isij, temperature difference Tdij, impedance difference Zdijand ZdiijThe calculation is as follows:Pdij=Pij-Pi(j-1) Tdij=Tij-Ti(j-1) Zdij=Zij-Zi(j-1) Zdiij=Ziij-Zii(j-1) Wherein, j represents the jth round of data collection, j-1 represents the previous round of data collection; Pij and Tij are the power data and temperature data in the jth round of data collection, Zij and Ziij are the impedance data in the jth round of data collection; Pi(j-1),Ti(j-1)is the power data and temperature data in the j-1th round of data collection, Zi(j-1)and Zii(j-1)The impedance data in the j-1th round of data collection; i represents the electrode number, and ii represents the number between two electrodes. 5. The method for determining the ablation endpoint of intravascular nerve ablation according to claim 1 is characterized in that in step (4), after each round of data collection, it is determined whether the power data, temperature data, and impedance data all reach the set target variation range for the first time; if the three all reach the set target variation range for the first time, the growth rates of the power data, temperature data, and impedance data are calculated; if the three do not reach the set target variation range at the same time or not for the first time, then the process proceeds to the next step; When the power, temperature and impedance all reach the target range after the jth round of data collection, the power growth rate Pinij, temperature growth rate Tinij, impedance growth rate Zinijand ZiniijThe calculation is as follows: Where, j represents the jth round of data collection, Pij and Tij are the power data and temperature data in the jth round of data collection, Zij and Ziij are the impedance data in the jth round of data collection, tjis the ablation duration in the jth round of data acquisition; i represents the electrode number, ii represents the number between two electrodes; Pis,Tisare the initial power and temperature values obtained from the first data collection, Zisand Ziisis the initial impedance value obtained from the first data acquisition. 6. The method for determining the ablation endpoint of intravascular nerve ablation according to claim 1 is characterized in that in step (4), after each round of data acquisition, the maximum power Pmax in the current ablation process isij,Minimum value Pminijand maximum fluctuation range PDijThe calculation is as follows: PDij=Pmaxij-Pminij After each round of data collection, the maximum temperature Tmax in the current ablation processij, minimum value Tminijand maximum fluctuation range TDijThe calculation is as follows: TDij=Tmaxij-Tminij After each round of data collection, the maximum impedance Zmax in the current ablation processijand Zmaxiij,Minimum value Zminijand ZminiijAnd the maximum fluctuation range ZDijand ZDiijThe calculation is as follows: ZDij=Zmaxij-Zminij ZDiij=Zmaxiij-Zminiij Among them, j represents the jth round of data collection, n represents the order number of each round of data collection, PinTo show the power value in each data collection, TinTo show the temperature value in each data collection, Zinand ZiinIndicates the impedance value collected each time, i represents the electrode number, and ii represents the number between the two electrodes. 7. The method for determining the ablation endpoint of intravascular nerve ablation as claimed in claim 1 is characterized in that in step (4), after each round of data acquisition, the power standard deviation Pstd in the current ablation process isij, temperature standard deviation Tstdij, impedance standard deviation Zstdijand ZstdiijThe calculation is as follows: Among them, j represents the jth round of data collection, n represents the order number of each round of data collection, PinTo show the power value in each data collection, TinTo show the temperature value in each data collection, Zinand ZiinIndicates the impedance value collected each time, i represents the electrode number, and ii represents the number between the two electrodes. 8. The method for determining the ablation endpoint of intravascular nerve ablation according to claim 1 is characterized in that step (5) includes the following sub-steps: (51) After each round of data collection, calculate the effective value r of each parameter datakj: Where j represents the jth round of data collection, k represents the number of the collected data and processed data obtained in a round of data collection, k = 1, 2, 3, ...; XkjRepresents specific collected data or processed data, ωkRepresents the set weight corresponding to specific collected data or processed data, YkRepresents a set target value corresponding to specific collected data or processed data; (52) Calculate the overall effective value Nj: Among them, 232 is the total number of collected data and processed data obtained in one round of data collection; (53) Accumulation of the total number of qualified times Q and the maximum number of consecutive qualified times Q1: Compare the overall effective value after each round of data collection with the set overall effective value threshold. When the overall effective value after data collection is greater than or equal to the overall effective value threshold, this data collection is recorded as a qualified number and added to the total qualified number (Q); at the same time, record the maximum value of the consecutive qualified number (Q1). 9. The method for determining the ablation endpoint of intravascular nerve ablation according to claim 8 is characterized in that: When the total number of qualified times obtained after a round of data collection reaches the set qualified times threshold, or when the maximum number of consecutive qualified times reaches the set consecutive qualified times threshold, it is determined that the ablation has reached the end point. 10. A radiofrequency ablation device, characterized by comprising an ablation catheter, an operating handle, and a radiofrequency ablation host, wherein the ablation catheter, the operating handle, and the radiofrequency ablation host are connected in sequence; wherein the front end of the ablation catheter is provided with a plurality of electrodes and a plurality of temperature sensing elements; The radiofrequency ablation host includes at least a radiofrequency signal generating module and an ablation control module; The ablation control module is used to control the entire ablation process and the ablation endpoint, and is composed of a processor and a memory, and the processor and the memory are coupled; wherein the memory is used to store a computer program, and the processor is used to run the computer program stored in the memory to execute the method for determining the ablation endpoint of intravascular nerve ablation as described in any one of claims 1 to 9. | 2024 | SHANGHAI GOLDEN LEAF MED TEC | Medical and Healthcare | NA | https://www.brattea.com/en/aboutUsEn.html | China | The patent focuses on a method for determining intravascular nerve ablation endpoints by setting target values and ranges for power, temperature, impedance, related operating parameters, and ablation duration. It addresses the difficulty of confidently knowing when to stop intravascular nerve ablation, which can lead to incomplete or excessive treatment. The multi-parameter target system gives a clearer ablation endpoint. | הפטנט מתמקד בשיטה לקביעת נקודות סיום של אבלציית עצבים תוך-וסקולרית על ידי קביעת ערכי מטרה וטווחים להספק, טמפרטורה, אימפדנס, פרמטרי הפעלה קשורים ומשך אבלציה. הוא מתמודד עם הקושי לדעת בביטחון מתי לעצור אבלציית עצבים תוך-וסקולרית, שיכול להוביל לטיפול לא שלם או עודף. מערכת המטרה הרב-פרמטרית נותנת נקודת סיום אבלציה ברורה יותר. | 1 |  |  |  | 1 | [5 empty] | 1 | [10 empty] | 1 | [7 empty] | 1 |  |  | SHANGHAI GOLDEN LEAF MED TEC CO LTD | SHANGHAI GOLDEN LEAF MED TEC CO LTD | SHANGHAI GOLDEN LEAF MED TEC CO LTD | SHANGHAI GOLDEN LEAF MED TEC CO LTD | SHANGHAI GOLDEN LEAF MED TEC CO LTD | 2025-01-15 | 2025 | 2025-07-31 | 2025 | 2024-01-26 | 2024 | Application | Utility Invention | active | WO2025CN72405 | 304660881 | CN117958953A | WO
443 | WO2022155334A1 | ABLATION DEVICES AND METHODS OF MANUFACTURING THE SAME | An ablation device includes a probe body, a distal face covering a distal end of the probe body and defining a central aperture, a distal tip extending through the central aperture in sealing relation therewith, and a sensing probe. The distal tip extends from a proximal end within an interior volume of the probe body to a distal end distally-spaced therefrom and defines a lumen. The sensing probe is sealed within the lumen and extends from within the interior volume distally through a portion of the lumen to a position proximally-spaced from the distal end of the distal tip. A portion of the sensing probe is exposed within an unoccupied volume of the lumen defined between the sensing probe and the distal end of the distal tip to enable direct contact between the exposed portion of the sensing probe and tissue extending into the unoccupied volume. | 1. WHAT IS CLAIMED IS: 1. An ablation device, comprising: a probe body defining an interior volume; a distal face covering a distal end of the probe body and defining a central aperture; a distal tip extending through the central aperture in sealing relation therewith, the distal tip including a proximal end disposed within the interior volume, a distal end distally-spaced from the probe body, and a lumen extending longitudinally between the proximal and distal ends of the distal tip; and a sensing probe disposed in sealing relation within the lumen, the sensing probe extending from within the interior volume distally through a portion of the lumen to a position proximally-spaced from the distal end of the distal tip, wherein a portion of the sensing probe is exposed within an unoccupied volume of the lumen defined between the sensing probe and the distal end of the distal tip to enable direct contact between the exposed portion of the sensing probe and tissue extending into the unoccupied volume. 2. The ablation device according to claim 1, wherein the distal end of the distal tip defines at least one cutting edge annularly surrounding the lumen. 3. The ablation device according to claim 2, wherein the at least one cutting edge includes a serrated annular cutting edge. 4. The ablation device according to claim 2, wherein the at least one cutting edge includes at least one beveled cutting edge and at least one peak cutting edge. 5. The ablation device according to claim 1, wherein the sensing probe includes a thermocouple configured to sense temperature of tissue in direct contact with the exposed portion of the sensing probe. 6. The ablation device according to claim 1, further comprising at least one fluid conduit extending at least partially through the interior volume and configured to deliver fluid to the interior volume or receive fluid from the interior volume to facilitate fluid circulation through the interior volume. 7. The ablation device according to claim 1, wherein the probe body includes or defines at least one electrode configured to be energized with RF energy for ablating tissue. 8. The ablation device according to claim 7, wherein the at least one electrode includes an active electrode configured to conduct RF energy through tissue to a remote return pad in a monopolar configuration. 9. The ablation device according to claim 7, wherein the at least one electrode includes first and second electrodes configured to conduct RF energy therebetween and through tissue in a bipolar configuration. 10. The ablation device according to claim 1, further comprising a connection hub supporting a proximal end portion of the probe body. 11. The ablation device according to claim 10, further comprising a cable coupled to the connection hub and housing at least one electrical lead for connection to the probe body. 12. The ablation device according to claim 10, further comprising inflow and outflow tubing coupled to the connection hub for providing fluid inflow and fluid outflow to and from the probe body. 13. The ablation device according to claim 1, further comprising an outer insulative jacket disposed about a portion of the probe body. 14. The ablation device according to claim 1, wherein the distal tip is welded to the distal face about the central aperture to seal the distal tip within the central aperture. 15. The ablation device according to claim 1, wherein the sensing probe is welded to the distal tip about the lumen to seal the sensing probe within the lumen. 16. The ablation device according to claim 1, wherein the distal face tapers distally and radially inwardly from an outer diameter of the probe body to the central aperture. 17. The ablation device according to claim 1, wherein the distal tip includes a proximal portion and a distal portion, the distal portion tapering distally and radially inwardly from an outer diameter of the proximal portion to the lumen at the distal end of the distal tip. 18. A method of manufacturing an ablation device, comprising: inserting a distal tip partially through a central aperture defined within a distal face covering a distal end of a probe body such that a proximal end of the distal tip is disposed within an interior volume of the probe body while a distal end of the distal tip remains distally- spaced from the probe body; welding the distal tip to the distal face about the central aperture to seal the distal tip within the central aperture of the distal face; inserting a sensing probe partially through a lumen defined within the distal tip such that the sensing probe extends from within the interior volume distally through a portion of the lumen to a position proximally-spaced from a distal end of the distal tip; and welding the sensing probe to the distal tip about the lumen to seal the sensing probe within the lumen of the distal tip. 19. The method according to claim 18, wherein the probe body includes or defines at least one electrode, and wherein the method further comprises attaching an electrical lead to each of the at least one electrodes. 20. The method according to claim 18, further comprising attaching a connection hub to a proximal portion of the probe body, the connection hub including inflow and outflow conduits extending therethrough for circulating fluid through the interior volume of the probe body. 15 | 2021 | Medtronic | Medical and Healthcare | Large-Size Company | https://www.medtronic.com | USA | The patent focuses on an ablation device with a probe body, a distal face with a central aperture, a distal tip extending through it in sealing relation, and a sensing probe with a lumen. It addresses the limitations of conventional ablation probes that lack sealed integration of distal tip and sensing probe. The integrated, sealed structure improves device reliability and sensing accuracy. | הפטנט מתמקד במכשיר אבלציה עם גוף גשוש, פנים דיסטלי עם פתח מרכזי, קצה דיסטלי המשתרע דרכו ביחס איטום, וגשוש חישה עם לומן. הוא מתמודד עם המגבלות של גששי אבלציה קונבנציונליים החסרים אינטגרציה אטומה של קצה דיסטלי וגשוש חישה. המבנה המשולב והאטום משפר את אמינות המכשיר ואת דיוק החישה. |  |  |  |  | 1 | 1 |  |  |  |  | 1 | [19 empty] | 1 |  | MEDTRONIC HOLDING CO SARL | Medtronic | MEDTRONIC HOLDING CO SARL | Medtronic | Medtronic | 2022-01-13 | 2022 | 2022-07-21 | 2022 | 2021-01-14 | 2021 | Application | Utility Invention | inactive | WO2022US12310 | 283837269 | EP4277555A1 | WO
444 | US20220233235A1 | Ablation Catheter for Pulsed-Field Ablation and Method for Electrode Position Assessment for Such Catheter | A system for treatment of patient tissue by delivery of high-voltage pulses comprising an ablation catheter, a measurement unit and an electronic control unit (ECU). The measurement unit is configured to perform measurements using an energy source, whereby the impedance and/or current measurement values are determined as response to an alternating voltage and/or at least one voltage pulse. The ECU is configured to receive and analyze said measurement values provided by the measurement unit and determine arcing risk (AR) indexes for said electrode pairs and/or a contact uniformity (CU) value based on said impedance measurement values and/or impedances for said electrodes and/or an impedance uniformity (IU) value based on said current measurement values. | 1. A system for treatment of patient tissue by delivery of high-voltage pulses, comprising: an ablation catheter, a measurement unit, and an electronic control unit (ECU), wherein the catheter comprises a catheter shaft, and an ablation portion being arranged at a distal end of the catheter shaft, with a plurality of electrodes accommodated along the ablation portion, wherein each of the plurality of electrodes is electrically connected to the measurement unit through the catheter shaft, wherein the measurement unit is configured to perform measurements using an energy source thereby determining measurement values of a subgroup of the plurality of electrodes, wherein said subgroup is formed by all or a part of the plurality of electrodes, wherein the ECU is configured to receive and analyze said measurement values provided by the measurement unit and determine arcing risk (AR) and/or a contact uniformity (CU) and/or impedance uniformity (IU) value indexes for said subgroup of the plurality of electrodes. 2. The system of claim 1, wherein said measurement values are bipolar impedance measurement values of electrode pairs of a subgroup of the plurality of electrodes and/or quasi-unipolar impedance measurement values of a subgroup of the plurality of electrodes and/or current measurement values of a subgroup of the plurality of electrodes. 3. The system of claim 2, wherein the impedance and/or current measurement values are determined as response to an alternating voltage and/or at least one voltage pulse. 4. The system of claim 1, wherein the determined arcing risk (AR) and/or a contact uniformity (CU) indexes are based on said impedance measurement values 5. The system of claim 1, wherein the impedance uniformity (IU) indexes are based on said current measurement values. 6. The system of claim 1, wherein the electronic control unit is arranged proximal to or at the proximal end of the catheter, and wherein the measurement unit is connected to or integrated within the ECU 7. The system of claim 1, wherein the measurement unit is configured to determine at least one current measurement value for each of the subgroup of the plurality of electrodes by measuring the respective current value of one or several of rectangular, sinusoidal, tooth or similar shaped voltage pulses, wherein one impedance value is determined from said determined current measurement values for each of the subgroup of electrodes. 8. The system of claim 2, wherein the ECU is configured to determine an impedance uniformity (IU) of two groups of the subgroup of electrodes, wherein ##EQU## wherein σ({Zd,p}) is the standard deviation and μ({Zd,p}) the mean value of the determined impedances of the electrodes of the respective group. 9. The system of claim 1, wherein the ECU is configured to determine the AR index for a particular electrode pair x,y from the bipolar impedance measurement values of the particular electrode pair x,y from the subgroup of electrodes scaled by the minimum of bipolar impedance measurement values of the respective electrodes with their adjoining electrodes of the subgroup. 10. The system of claim 1, wherein the ECU is configured to determine the CU value for the subgroup of electrodes based on the standard deviation of the bipolar impedance measurement values of the pairs of adjoining electrodes of said subgroup or based on the minimum and the maximum of the bipolar impedance measurement values of the pairs adjoining electrodes of said subgroup and/or to determine the CU value for the subgroup of electrodes based on the standard deviation of a quasi-unipolar impedance measurement values of all electrodes of said subgroup or based on the minimum and the maximum of the quasi-unipolar impedance measurement values of all electrodes of said subgroup. 11. The system of claim 1, wherein the ECU is configured to determine an overall risk for arcing for all electrodes of the subgroup based on a maximum of the AR index of all electrode pairs of the subgroup. 12. The system of claim 1, wherein the measurement unit is configured such that the frequency for determination of quasi-unipolar or bipolar impedance measurement values of the subgroup of electrodes is between 1 kHz and 1 MHz and/or such that the voltage amplitude of the pulses is between 1V and 1 kV, in particular between 10V and 700V, in particular between 100V and 500V 13. A method for assessment of positions and/or configuration of a plurality of electrodes of an ablation catheter for treatment of patient tissue by delivery of high-voltage pulses comprising a catheter shaft and an ablation portion, wherein the ablation portion is arranged at a distal end of the catheter shaft with the plurality of electrodes accommodated along the ablation portion, wherein each of the plurality of electrodes is electrically connected to a measurement unit through the catheter shaft, wherein the measurement unit performs measurements using an energy source thereby determining measurement values of a subgroup of the plurality of electrodes, wherein the subgroup is formed by all electrodes or a part of the plurality of electrodes, respectively, wherein said measurement values are transmitted to the ECU which receives and analyzes said measurement values as well as determines arcing risk (AR) and/or a contact uniformity (CU) and/or an impedance uniformity (IU) indexes based on said current measurement values. 14. The method of claim 13, wherein said measurement values are bipolar impedance measurement values of electrode pairs of a subgroup of the plurality of electrodes and/or quasi-unipolar impedance measurement values of a subgroup of the plurality of electrodes and/or current measurement values of a subgroup of the plurality of electrodes. 15. The method of claim 14, wherein the impedance and/or current measurement values are determined as response to an alternating voltage and/or at least one voltage pulse. 16. The method of claim 13, wherein the determined arcing risk (AR) and/or a contact uniformity (CU) indexes are based on said impedance measurement values 17. The method of claim 13, wherein the impedance uniformity (IU) indexes are based on said current measurement values. 18. The method of claim 13, wherein the electronic control unit is arranged proximal to or at the proximal end of the catheter, and wherein the measurement unit is connected to or integrated within the ECU 19. The method of claim 13, wherein the electronic control unit is arranged separate from catheter, and wherein the measurement unit is connected to or integrated within the ECU 20. The method of claim 13, wherein the measurement unit determines at least one current measurement value for each of the subgroup of the plurality of electrodes by measuring the respective current value of one or several of rectangular, sinusoidal, tooth or similar shaped voltage pulses, wherein one impedance value is determined from said determined current measurement values for each electrode of the subgroup of electrodes. 21. The method of claim 20, wherein the ECU determines an impedance uniformity (IU) of two groups of the subgroup of electrodes, wherein ##EQU## wherein ({Zd,p}) is the standard deviation and μ({Zd,p}) the mean value of the determined impedances of the electrodes of the respective group. 22. The method of claim 13, wherein the ECU determines the AR index for a particular electrode pair x,y from the bipolar impedance measurement values of the particular electrode pair x,y from the subgroup of electrodes scaled by the minimum of bipolar impedance measurement values of the respective electrodes with their adjoining electrodes of the subgroup. 23. The method of claim 13, wherein the ECU determines the CU value for the subgroup of electrodes based on the standard deviation of the bipolar impedance measurement values of the pairs of adjoining electrodes of said subgroup or based on the minimum and the maximum of the bipolar impedance measurement values of the pairs adjoining electrodes of said subgroup and/or determines the CU value for the subgroup of electrodes based on the standard deviation of the quasi-unipolar impedance measurement values of all electrodes of said subgroup or based on the minimum and the maximum of the quasi-unipolar impedance measurement values of all electrodes of said subgroup. 24. The method of claim 13, wherein ECU determines an overall risk for arcing for all electrodes of the subgroup based on a maximum of the AR index of all electrode pairs of the subgroup. 25. A computer program product comprising instructions which, when executed by a processor, cause the processor to perform the steps of the method according to claim 13. 26. A computer readable data carrier storing a computer program product according to claim 25. | 2021 | CRC EP | NA | NA | NA | USA | The patent focuses on a PFA ablation catheter system with a measurement unit determining impedance and/or current measurement values from an alternating voltage and at least one voltage pulse. It addresses the challenge of assessing electrode position during PFA, which is needed for correct lesion creation. Impedance and current measurements help the ECU evaluate where the electrodes are during PFA. | הפטנט מתמקד במערכת קטטר אבלציית PFA עם יחידת מדידה הקובעת ערכי מדידת אימפדנס ו/או זרם ממתח חילופין ולפחות פולס מתח אחד. הוא מתמודד עם האתגר בהערכת מיקום אלקטרודה במהלך PFA, הנחוץ ליצירת נגע נכונה. מדידות אימפדנס וזרם עוזרות ל-ECU להעריך היכן האלקטרודות במהלך PFA. | 1 | [10 empty] | 1 | [9 empty] | 1 |  |  | 1 | [7 empty] | CRC EP INC | CRC EP INC | CRC EP, INC. | CRC EP INC | CRC EP INC | 2021-12-28 | 2021 | 2022-07-28 | 2022 | 2021-01-22 | 2021 | Application | Utility Invention | active | US202117563620 | 327111074 | CN116669643A¦EP4280985A1¦EP4280985A4¦EP4280988A1¦EP4280988A4¦EP4280989A1¦EP4280989A4¦JP2024503564A¦US12082877B2¦US12193732B2¦US20220233234A1¦US20220233236A1¦US20240261018A1¦US20250143783A1¦WO2022159595A1¦WO2022159659A1¦WO2022159665A1 | US
445 | WO2023175478A1 | PROBE, SYSTEM, AND METHOD FOR FORMING A LESION IN A TARGET TISSUE | A probe for forming a lesion in a target tissue includes an elongate member extending longitudinally between a proximal end and a distal end. The elongate member has an active electrode proximate the distal end for delivering ablative energy to the target tissue, and an electrically insulated section proximal of the active electrode. The active electrode includes a cooled section that is coolable by receipt of a cooling fluid delivered through the elongate member. At least a first temperature sensor is positioned external of the elongate member to measure a temperature in a heat affected zone of the target tissue. The first temperature sensor is positioned proximally of the distal end and is spaced from the cooled section. | 1. CLAIMS: 1 . A probe for forming a lesion in a target tissue, the probe comprising: an elongate member extending longitudinally between a proximal end and a distal end, wherein the elongate member has an active electrode proximate the distal end for delivering ablative energy to the target tissue, and an electrically insulated section proximal of the active electrode, and wherein the active electrode has a cooled section that is coolable by receipt of a cooling fluid delivered through the elongate member; and at least a first temperature sensor positioned external of the elongate member to measure a temperature in a heat affected zone of the target tissue, wherein the first temperature sensor is positioned proximally of the distal end and is spaced from the cooled section. 2. The probe of claim 1 , wherein the elongate member comprises at least a first fluid delivery lumen for delivering the cooling fluid through the elongate member to the cooled section. 3. The probe of claim 2, wherein the elongate member further comprises at least a first fluid return lumen for delivering the cooling fluid proximally through the elongate member. 4. The probe of any one of claims 1 to 3, wherein the elongate member is configured to restrict the cooling fluid to the cooled section on delivery of the cooling fluid. 5. The probe of any one of claims 1 to 4, further comprising a barrier within the elongate member for restricting the cooling fluid to the cooled section on delivery of the cooling fluid, wherein the barrier is positioned distally of a proximal boundary of the active electrode, to further provide the active electrode with a non-cooled section. 6. The probe of claim 5, wherein the first temperature sensor is secured to the active electrode in the non-cooled section. 7. The probe of any one of claims 1 to 4 further comprising a barrier within the elongate member for restricting the cooling fluid to the cooled section on delivery of the cooling fluid, wherein the barrier is positioned at a proximal boundary of the active electrode so that the cooled section makes up an entirety of the active electrode. 8. The probe of any one of claims 1 to 7, wherein the first temperature sensor is positioned proximally of the active electrode. 9. The probe of any one of claims 1 to 8, wherein the probe comprises an electrically insulative material forming the electrically insulated section or received on the electrically insulated section, and wherein the first temperature sensor is secured to the electrically insulative material. 10. The probe of any one of claims 1 to 9, wherein the first temperature sensor is positioned to measure the temperature in a proximal projection of a thermal ablation zone. 11 .The probe of claim 1 , wherein the active electrode further comprises a non-cooled section, and wherein the first temperature sensor is secured to the non-cooled section. 12. The probe of claim 1 , wherein the first temperature sensor is spaced at least proximally from the cooled section. 13. The probe of claim 12, wherein the first temperature sensor is further spaced radially from the cooled section. The probe of any one of claims 1 to 8, wherein the elongate member comprises a metallic hypotube, wherein the metallic hypotube comprises a first section that is electrically exposed to form the active electrode, and a second section on which an electrically insulative material is received to form the electrically insulated section. The probe of claim 14, further comprising at least one electrical conductor for electrically connecting the first temperature sensor to a temperature control system, wherein the electrical conductor is positioned between the second section of the metallic hypotube and the electrically insulative material. The probe of claim 14 or 15, wherein the metallic hypotube and the electrically insulative material are permanently secured together. The probe of claims 14 or 15, wherein the metallic hypotube is removably received in the electrically insulative material. The probe of any one of claims 1 to 14, further comprising at least one electrical conductor for electrically connecting the first temperature sensor to a temperature control system, wherein the electrical conductor extends through a lumen of the elongate member. The probe of any one of claims 1 to 8, wherein the electrically insulated section comprises a hypotube fabricated from an electrically insulative material, and the active electrode comprises a metallic member secured to the electrically insulative material. The probe of any one of claims 1 to 19, wherein the first temperature sensor is further configured to deliver ablative energy to the target tissue. The probe of any one of claims 1 to 20, wherein the first temperature sensor comprises a thermocouple junction. A system for forming a lesion in a target tissue, comprising: a probe comprising an elongate member and a first temperature sensor, wherein the elongate member extends longitudinally between a proximal end and a distal end, wherein the elongate member has an active electrode proximate the distal end for delivering ablative energy to the target tissue, and an electrically insulated section proximal of the active electrode, wherein the active electrode comprises a cooled section that is coolable by receipt of a cooling fluid delivered through the elongate member, wherein the first temperature sensor is positioned external of the elongate member to measure a temperature in a heat affected zone of the target tissue, and wherein the first temperature sensor is positioned proximally of the distal end and is spaced from the cooled section; a return electrode spaced from the active electrode; a radiofrequency generator electrically connected to the active electrode for delivering the ablative energy to the active electrode and electrically connected to the return electrode for returning a return current to the radiofrequency generator; and a cooling fluid circulation system for delivering the cooling fluid to the probe. A method for forming a lesion in a target tissue, comprising: positioning an active electrode of a probe at the target tissue; delivering ablative energy from the active electrode to form the lesion in the target tissue, while delivering a cooling fluid to a cooled section of the active electrode; and using a first temperature sensor of the probe to monitor a temperature in a heat affected zone of the target tissue at a position proximal of a distal end of the probe and spaced from the cooled section. 24. The method of claim 23, wherein delivering the cooling fluid to the cooled section of the active electrode comprises delivering the cooling fluid to an entirety of the active electrode. 25. The method of claim 23, wherein delivering the cooling fluid to the cooled section of the active electrode comprises, on delivery of the cooling fluid, restricting the cooling fluid to the cooled section. 26. The method of any one of claims 23 to 25, wherein using the first temperature sensor comprises monitoring the temperature proximally of the active electrode. 27. The method of any one of claims 23 to 26, wherein the lesion has a proximal projection, and the first temperature sensor monitors the temperature in the proximal projection. 28. The method of claim 25, wherein restricting the cooling fluid to the cooled section comprising using a barrier positioned at a proximal boundary of the active electrode to restrict the cooling fluid to the active electrode. 29. The method of claim 25, wherein restricting the cooling fluid to the cooled section comprising using a barrier positioned distally of a proximal boundary of the active electrode, to further provide the active electrode with a non-cooled section. 30. The method of claim 29, wherein using the first temperature sensor comprises monitoring the temperature adjacent the non-cooled section. 31. The method of any one of claims 23 to 30, wherein the position is radially spaced from the cooled section. 32. The method of any one of claims 21 to 29, further comprising delivering the ablative energy to the target tissue from the first temperature sensor. | 2022 | BAYLIS MEDICAL TECH | Medical and Healthcare | Mid-Size Company | https://baylismedtech.com/ | Canada | The patent focuses on a probe for forming a lesion in target tissue, with an active electrode at the distal end for ablation and an electrically insulated section proximal of it, where the active electrode includes a cooled section. It addresses the unwanted heating of non-target tissue around the active electrode during ablation. Cooling at the active section limits collateral heating while still forming the intended lesion. | הפטנט מתמקד בגשוש ליצירת נגע ברקמת מטרה, עם אלקטרודה פעילה בקצה הדיסטלי לאבלציה ומקטע מבודד חשמלית פרוקסימלי לה, שבה האלקטרודה הפעילה כוללת מקטע מקורר. הוא מתמודד עם החימום הלא רצוי של רקמה לא-מטרתית סביב האלקטרודה הפעילה במהלך אבלציה. קירור במקטע הפעיל מגביל חימום נלווה תוך יצירת הנגע המיועד. |  |  |  |  | 1 | [5 empty] | 1 | [19 empty] | 1 |  | BAYLIS MEDICAL TECH INC¦BAYLIS MEDICAL TECH USA INC | BAYLIS MEDICAL TECH INC¦BAYLIS MEDICAL TECH USA INC | BAYLIS MEDICAL TECH INC¦BAYLIS MEDICAL TECH USA INC | BAYLIS MEDICAL TECH INC¦BAYLIS MEDICAL TECH USA INC | BAYLIS MEDICAL TECH INC¦BAYLIS MEDICAL TECH USA INC | 2023-03-13 | 2023 | 2023-09-21 | 2023 | 2022-03-15 | 2022 | Application | Utility Invention | inactive | WO2023IB52412 | 321683617 | CA3254408A1¦EP4493091A1¦EP4493091A4¦US20250186111A1 | WO
446 | WO2025214052A1 | LASER ABLATION CATHETER, LASER ABLATION SYSTEM, AND REGULATION METHOD FOR LASER ABLATION SYSTEM | The present invention relates to the technical field of medical instruments and regulation methods therefor. Disclosed are a laser ablation catheter, a laser ablation system, and a regulation method for the laser ablation system. The laser ablation catheter comprises a tubular structure and an optical fiber bundle. The tubular structure is provided with an optical fiber channel. The optical fiber bundle is accommodated in the optical fiber channel. The optical fiber bundle is provided with an output end, and the output end is located at the far end of the tubular structure. The laser ablation system comprises the described laser ablation catheter. The present invention further provides a regulation method for a renal artery laser ablation system for regulating the described laser ablation system. The laser ablation catheter, the laser ablation system, and the regulation method for the laser ablation system provided by the present invention have good application effects. | 1. A laser ablation catheter is characterized by comprising a tube structure and an optical fiber bundle, wherein the tube structure has an optical fiber channel, the optical fiber bundle is accommodated in the optical fiber channel, and the optical fiber bundle has an output end, which is located at the distal end of the tube structure. 2. The laser ablation catheter according to claim 1 is characterized in that the laser ablation catheter is a renal artery laser ablation catheter, the optical fiber bundle includes multiple first optical fibers with a core diameter less than 200 μm, the first optical fibers are used to transmit a first laser for irradiating the sympathetic nerves, and the first optical fibers are placed in the extension direction of the optical fiber channel. 3. The laser ablation catheter according to claim 2, characterized in that the bending radius of the first optical fiber is less than 10 mm; and/or the first optical fiber is a multimode optical fiber. 4. The laser ablation catheter according to claim 2 is characterized in that the laser ablation catheter is provided with a temperature sensor device for obtaining the temperature of the target part. 5. The laser ablation catheter according to claim 4, wherein the temperature sensor device includes a second optical fiber, the second optical fiber being used to transmit a second laser light having a central reflection wavelength for sensing the temperature of the target site, the second optical fiber being disposed among the plurality of first optical fibers and being accommodated in the optical fiber channel simultaneously with the first optical fibers; Alternatively, the temperature sensor device is a resistance temperature device or a thermocouple. 6. The laser ablation catheter according to claim 5, characterized in that the bending radius of the second optical fiber is greater than or equal to 10 mm and less than or equal to 25 mm; and/or the core diameter of the second optical fiber is greater than or equal to 8 μm and less than or equal to 12 μm. 7. The laser ablation catheter according to claim 2, characterized in that the output end of the optical fiber bundle is provided with a structure for releasing laser energy in a set direction; Alternatively, the output end of the optical fiber bundle is provided with a structure for releasing the laser energy in a circumferential direction. 8. The laser ablation catheter according to claim 2 is characterized in that the laser ablation catheter further comprises a balloon assembly, wherein the balloon assembly is sealingly sleeved on the distal end of the tube structure so that the distal ends of the tube structure and the optical fiber bundle are disposed within the balloon assembly. 9. The laser ablation catheter according to claim 5, characterized in that the tube structure comprises an inner tube and an outer tube; The outer tube is sleeved on the periphery of the inner tube and is eccentrically arranged with the inner tube to enclose a first optical fiber placement channel, and the first optical fiber is placed in the first optical fiber placement channel; the inner tube has a second optical fiber placement channel, and the second optical fiber is placed in the second optical fiber placement channel; The optical fiber channel includes the first optical fiber placement channel and the second optical fiber placement channel. 10. The laser ablation catheter according to claim 1, characterized in that the tube structure is a slender tube structure, comprising an inner tube and an outer tube; the outer tube is sleeved on the periphery of the inner tube and is eccentrically arranged with the inner tube to enclose a first optical fiber placement channel, and the inner tube is used for passing a guide wire; The output end of the optical fiber bundle is located at the distal end of the tube structure and has an inclined surface that is angled with the extended axis of the tube structure; The optical fiber bundle includes multiple first optical fibers, the core diameter of the first optical fibers is smaller than the core diameter of the optical fiber bundle; the multiple first optical fibers are arranged in the first optical fiber placement channel. 11. The laser ablation catheter as described in claim 10 is characterized in that the laser ablation catheter further includes a rotating member, which is connected to the proximal end of the tube structure and is used to drive the tube structure and the optical fiber bundle to rotate around the axial direction of the laser ablation catheter in a preset manner, so that the output laser of the optical fiber bundle can irradiate the area to be ablated at an angle of 360°. 12. The laser ablation catheter according to claim 11, characterized in that the rotating member is connected to the proximal end of the tube structure and is used to drive the tube structure and the optical fiber bundle to rotate around the axial direction of the laser ablation catheter in a preset manner, so that the output laser of the optical fiber bundle can irradiate the area to be ablated at an angle of 360°, specifically comprising: The rotating part is a Luer adapter, and the preset mode is to emit laser light once every 45 degrees or 90 degrees for 10 seconds or 15 seconds. 13. The laser ablation catheter as described in claim 10 is characterized in that the laser ablation catheter also includes a balloon assembly, which is sealed and sleeved on the distal end of the tube structure along the extension direction of the catheter, so that the inclined surface is arranged inside the balloon assembly. 14. The laser ablation catheter according to claim 9 or 13, characterized in that the balloon assembly includes a first balloon and/or a second balloon; The first balloon and/or the second balloon are both sealed and sleeved on the distal end of the tube structure or the optical fiber bundle, and the second balloon is sleeved on the periphery of the first balloon. 15. The laser ablation catheter according to claim 7 or 10, characterized in that the output end of the optical fiber bundle is an inclined surface; the inclined surface is arranged at an angle with the central axis of the laser ablation catheter, and the angle is greater than or equal to 20° and less than or equal to 40°; Alternatively, the output end of the optical fiber bundle is a tapered tip or a tapered groove, the tapered angle of the tapered tip or the tapered groove is greater than or equal to 30° and less than or equal to 80°, or the output end of the optical fiber bundle is a truncated cone tip. 16. The laser ablation catheter according to claim 11, wherein the optical fiber bundle further comprises a second optical fiber, the second optical fiber being located within the first optical fiber placement channel and sharing the optical fiber bundle input end and the optical fiber bundle output end with the first optical fiber, or the second optical fiber being located within the inner tube; the second optical fiber being used to sense the temperature of the distal end of the optical fiber bundle at the site to be ablated; Alternatively, the laser ablation catheter is provided with a temperature sensor for acquiring the temperature of the target site. 17. The laser ablation catheter according to claim 10, characterized in that the bending radius of the first optical fiber is less than 10 mm; and/or the core diameter of the first optical fiber is less than 200 μm. 18. The laser ablation catheter according to claim 16, characterized in that the bending radius of the second optical fiber is greater than or equal to 10 mm and less than or equal to 25 mm; and/or the core diameter of the second optical fiber is greater than or equal to 8 μm and less than or equal to 12 μm. 19. The laser ablation catheter according to claim 13 is characterized in that it further includes: a catheter seat, the catheter seat is connected to the proximal end of the tube structure, and the catheter seat is provided with a liquid injection port for injecting liquid into the balloon assembly. 20. A laser ablation system, characterized by comprising the laser ablation catheter according to any one of claims 1 to 19. 21. The laser ablation system according to claim 20 is characterized in that it also includes a control system, wherein the control system is at least used to control the laser ablation catheter to output a laser with a wavelength of 1064 nm. 22. A laser ablation system, characterized in that the laser ablation system is a renal artery laser ablation system, comprising a laser ablation catheter according to any one of claims 1 to 19, wherein the optical fiber bundle of the laser ablation catheter comprises a first optical fiber and a second optical fiber, the first optical fiber being used to deliver a first laser and release the first laser to a sympathetic nerve at a target site, and the second optical fiber being used to deliver a second laser that senses the temperature of the target site and generates a central reflection wavelength; A control unit, the control unit being connected to at least the optical path of the laser ablation catheter and configured to perform the following steps: emitting the first laser light toward the first optical fiber, where the first laser light is used to irradiate a target area; After emitting the second laser light toward the second optical fiber, receiving a central reflection wavelength from the second optical fiber, and obtaining a temperature of a target portion based on the central reflection wavelength; The energy of the first laser beam emitted next time is determined according to the temperature of the target part. 23. The laser ablation system according to claim 22 is characterized in that the receiving of the central reflection wavelength from the second optical fiber and the acquisition of the temperature of the target part according to the central reflection wavelength specifically include: the central reflection wavelength and the temperature of the target part are in a mapping relationship. 24. A method for controlling a laser ablation system, characterized in that it is used to control the laser ablation system according to any one of claims 22 to 23, comprising the following steps: emitting a first laser and a second laser, wherein the first laser is used to irradiate the sympathetic nerves of the target area, and the second laser is used to trigger the temperature sensing of the target area; Obtain the temperature of the target area; The energy of the first laser beam emitted next time is determined according to the temperature of the target site. 25. The control method of the laser ablation system according to claim 24 is characterized in that determining the emission energy of the first laser next time based on the temperature of the target site specifically comprises the following steps: The first laser uses a continuous laser, and when the temperature of the target part is less than or equal to a preset maximum temperature value and greater than or equal to a preset minimum temperature value, the output power of the first laser is stopped from being adjusted for the next time; When the temperature of the target part is greater than the preset maximum temperature value, the output power of the first laser is reduced next time; When the temperature of the target part is lower than the preset minimum temperature value, the output power of the first laser is increased next time; Alternatively, the first laser is a pulsed laser, and when the temperature of the target portion is less than or equal to a preset maximum temperature value and greater than or equal to a preset minimum temperature value, the adjustment of the output power or frequency of the first laser for the next time is stopped; When the temperature of the target part is greater than the preset maximum temperature value, the output power or frequency of the first laser is reduced next time; When the temperature of the target part is lower than the preset minimum temperature value, the output power or frequency of the next first laser is increased. 26. The control method of the laser ablation system according to claim 25 is characterized in that the preset maximum temperature value is 60°C and the preset minimum temperature value is 55°C. 27. The control method of a laser ablation system according to any one of claims 24 to 26, wherein emitting the first laser is specifically: The wavelength of the first laser is between 1034 nm and 1094 nm. 1. [Edited] A laser ablation catheter, characterized in that it includes a tube structure, an optical fiber bundle and a temperature sensor for obtaining the temperature of a target part; the tube structure includes an inner tube and an outer tube sleeved around the inner tube, the outer tube and the inner tube are eccentrically arranged and enclose a first optical fiber placement channel, and the inner tube has a second optical fiber placement channel; the optical fiber bundle includes a plurality of first optical fibers for transmitting a first laser for irradiating sympathetic nerves, the first optical fibers are accommodated in the first optical fiber placement channel, and the temperature sensor is accommodated in the second optical fiber placement channel; the optical fiber bundle has an output end, and the output end is located at the distal end of the tube structure. 2. [Modified] The laser ablation catheter according to claim 1 is characterized in that the laser ablation catheter is a renal artery laser ablation catheter, and the core diameter of the first optical fiber is less than 200 μm. 3. The laser ablation catheter according to claim 2, characterized in that the bending radius of the first optical fiber is less than 10 mm; and/or the first optical fiber is a multimode optical fiber. 4. [Modified] The laser ablation catheter according to claim 2, characterized in that the temperature sensor device includes a second optical fiber, which is used to transmit a second laser that senses the temperature of the target area and generates a central reflection wavelength; or, the temperature sensor device is a resistance temperature device or a thermocouple. 5. [Modified] The laser ablation catheter according to claim 4, characterized in that the bending radius of the second optical fiber is greater than or equal to 10 mm and less than or equal to 25 mm; and/or, the core diameter of the second optical fiber is greater than or equal to 8 μm and less than or equal to 12 μm. 6. [Modified] The laser ablation catheter according to claim 2 is characterized in that the output end of the optical fiber bundle is provided with a structure for releasing the laser energy in a set direction; or, the output end of the optical fiber bundle is provided with a structure for releasing the laser energy in a circumferential direction. 7. [Modified] The laser ablation catheter according to claim 2, characterized in that the laser ablation catheter further comprises a balloon assembly, wherein the balloon assembly is sealingly sleeved on the distal end of the tube structure so that the distal ends of the tube structure and the optical fiber bundle are disposed within the balloon assembly. 8. [Modified] The laser ablation catheter according to claim 1 is characterized in that the second optical fiber placement channel of the inner tube is also used for allowing a guide wire to pass through; the output end of the optical fiber bundle is located at the distal end of the tube structure and has an inclined surface that is angled with the extended axis of the tube structure. 9. [Modified] The laser ablation catheter according to claim 8 is characterized in that the laser ablation catheter further comprises a rotating member, which is connected to the proximal end of the tube structure and is used to drive the tube structure and the optical fiber bundle to rotate around the axial direction of the laser ablation catheter in a preset manner, so that the output laser of the optical fiber bundle can irradiate the area to be ablated at an angle of 360°. 10. [Modified] The laser ablation catheter according to claim 9 is characterized in that the rotating member is connected to the proximal end of the tube structure and is used to drive the tube structure and the optical fiber bundle to rotate around the axis of the laser ablation catheter in a preset manner, so that the output laser of the optical fiber bundle can irradiate the area to be ablated at an angle of 360°, specifically including: the rotating member is a Luer adapter, and the preset manner is to emit laser light once every 45 degrees or 90 degrees for 10 seconds or 15 seconds. 11. [Modified] The laser ablation catheter according to claim 8 is characterized in that the laser ablation catheter further includes a balloon assembly, which is sealed and sleeved on the distal end of the tube structure along the extension direction of the catheter, so that the inclined surface is arranged inside the balloon assembly. 12. [Modified] The laser ablation catheter according to claim 1 or 11, characterized in that the balloon assembly includes a first balloon and/or a second balloon; the first balloon and/or the second balloon are both sealed and sleeved on the distal end of the tube structure or the optical fiber bundle, and the second balloon is sleeved on the periphery of the first balloon. 13. [Modified] The laser ablation catheter according to claim 6 or 8, characterized in that the output end of the optical fiber bundle is an inclined surface; the inclined surface is set at an angle with the central axis of the laser ablation catheter, and the angle is greater than or equal to 20° and less than or equal to 40°; or, the output end of the optical fiber bundle is a tapered tip or a tapered groove, and the tapered angle of the tapered tip or the tapered groove is greater than or equal to 30° and less than or equal to 80°, or, the output end of the optical fiber bundle is a truncated cone tip. 14. [Modified] The laser ablation catheter according to claim 9, characterized in that the temperature sensor device includes a second optical fiber, and the second optical fiber is placed in the second optical fiber placement channel of the inner tube; the second optical fiber is used to sense the temperature of the distal end of the optical fiber bundle at the site to be ablated. 15. [Modified] The laser ablation catheter according to claim 8, characterized in that the bending radius of the first optical fiber is less than 10 mm; and/or the core diameter of the first optical fiber is less than 200 μm. 16. [Modified] The laser ablation catheter according to claim 14, characterized in that the bending radius of the second optical fiber is greater than or equal to 10 mm and less than or equal to 25 mm; and/or the core diameter of the second optical fiber is greater than or equal to 8 μm and less than or equal to 12 μm. 17. [Modified] The laser ablation catheter according to claim 11 is characterized in that it also includes: a catheter seat, the catheter seat is connected to the proximal end of the tube structure, and the catheter seat is provided with a liquid injection port for injecting liquid into the balloon assembly. 18. [Modified] A laser ablation system, characterized by comprising the laser ablation catheter according to any one of claims 1 to 17. 19. [Modified] The laser ablation system according to claim 18 is characterized in that it also includes a control system, wherein the control system is at least used to control the laser ablation catheter to output a laser with a wavelength of 1064 nm. 20. [Edited] A laser ablation system, characterized in that the laser ablation system is a renal artery laser ablation system, comprising a laser ablation catheter as described in any one of claims 1 to 17, the temperature sensing device comprising a second optical fiber, the second optical fiber being used to transmit a second laser for sensing the temperature of a target site and generating a central reflection wavelength; a control unit, the control unit being connected to at least an optical path of the laser ablation catheter, and being used to perform the following steps: emitting the first laser toward the first optical fiber, the first laser being used to irradiate the target site; after emitting the second laser toward the second optical fiber, receiving the central reflection wavelength from the second optical fiber, and obtaining the temperature of the target site based on the central reflection wavelength; and determining the emission energy of the first laser for the next time based on the temperature of the target site. 21. [Modified] The laser ablation system according to claim 20, wherein the receiving of the central reflection wavelength from the second optical fiber and obtaining the temperature of the target part according to the central reflection wavelength specifically includes: the central reflection wavelength and the temperature of the target part are in a mapping relationship. 22. [Edited] A method for controlling a laser ablation system, characterized in that it is used to control the laser ablation system as described in any one of claims 20-21, comprising the following steps: emitting a first laser and a second laser, the first laser being used to irradiate the sympathetic nerves of a target area, and the second laser being used to trigger the sensing of the temperature of the target area; obtaining the temperature of the target area; and determining the emission energy of the first laser next time based on the temperature of the target area. 23. [Modified] The control method of the laser ablation system according to claim 22 is characterized in that the determining the emission energy of the first laser for the next time according to the temperature of the target part specifically includes the following steps: the first laser uses a continuous laser, and when the acquired temperature of the target part is less than or equal to a preset maximum temperature value and greater than or equal to a preset minimum temperature value, the adjustment of the output power of the first laser for the next time is stopped; when the acquired temperature of the target part is greater than the preset maximum temperature value, the output power of the first laser for the next time is reduced; when the acquired temperature of the target part is less than the preset minimum temperature value, the output power of the first laser for the next time is increased; or, the first laser uses a pulsed laser, and when the acquired temperature of the target part is less than or equal to the preset maximum temperature value and greater than or equal to the preset minimum temperature value, the adjustment of the output power or frequency of the first laser for the next time is stopped; when the acquired temperature of the target part is greater than the preset maximum temperature value, the output power or frequency of the first laser for the next time is reduced; when the acquired temperature of the target part is less than the preset minimum temperature value, the output power or frequency of the first laser for the next time is increased. 24. [Modified] The control method of the laser ablation system according to claim 23, characterized in that the preset maximum temperature value is 60°C and the preset minimum temperature value is 55°C. 25. [Modified] The control method of a laser ablation system according to any one of claims 22 to 24, characterized in that the emission of the first laser is specifically: the wavelength of the first laser is between 1034nm and 1094nm. | 2024 | MICRO ENERGY INTERVENTION MEDICAL | Medical and Healthcare | NA | https://www.micro-energy.com/en/ | China | The patent focuses on a laser ablation catheter with a tubular structure containing an optical fiber bundle and an optical fiber channel, plus a regulation method for the system. It addresses the difficulty of regulating laser power and distribution in catheter-based laser ablation. The fiber arrangement and regulation method improve laser delivery consistency during ablation. | הפטנט מתמקד בקטטר אבלציית לייזר עם מבנה צינורי המכיל צרור סיבים אופטיים וערוץ סיב אופטי, בתוספת שיטת ויסות למערכת. הוא מתמודד עם הקושי בויסות הספק והתפלגות לייזר באבלציית לייזר מבוססת קטטר. סידור הסיבים ושיטת הויסות משפרים את עקביות מסירת הלייזר במהלך האבלציה. |  | 1 |  |  | 1 | [8 empty] | 1 | [7 empty] | 1 |  |  |  | 1 | [6 empty] | MICRO ENERGY INTERVENTION MEDICAL CO LTD | MICRO ENERGY INTERVENTION MEDICAL CO LTD | MICRO ENERGY INTERVENTION MEDICAL CO LTD | MICRO ENERGY INTERVENTION MEDICAL CO LTD | MICRO ENERGY INTERVENTION MEDICAL CO LTD | 2025-03-12 | 2025 | 2025-10-16 | 2025 | 2024-04-07 | 2024 | Application | Utility Invention | active | WO2025CN82121 | 310702761 | CN118058830A¦CN118105164A¦CN221023056U¦CN222217931U | WO
447 | WO2023205066A1 | METHODS AND SYSTEMS TO IDENTIFY ACTUAL ESOPHAGEAL TISSUE CHANGES DURING CARDIAC ABLATION | Methods and systems to identify actual esophageal tissue changes (e.g., during cardiac ablation) are disclosed. An example system includes a probe and a data analysis unit. The probe is configured for positioning in a patient's esophagus and includes a sheath, and a plurality of tendrils extendible by a user from an end of the sheath within the esophagus. Each tendril has a plurality of nodes spaced apart along the tendril, wherein each node of the plurality of nodes is configured to be used to introduce signals into esophageal tissue and/or receive signals affected by esophageal tissue. The data analysis unit is configured to determine a value of a characteristic indicative of actual change to esophageal tissue between two nodes based on a signal received using at least one of the nodes, and to provide an indication of actual esophageal tissue change based on the determined value. | 1. What Is Claimed Is: 1 . A system for detecting changes in esophageal tissue properties, the system comprising: a probe configured for positioning in a patient’s esophagus, the probe including one or more tendrils that can be manipulated be a user within the esophagus, wherein the one or more tendrils are configured to contact respective esophageal tissue at different locations in the esophagus, wherein each tendril includes a plurality of nodes spaced apart along the tendril, and wherein each node of the plurality of nodes is configured to be used to introduce signals into esophageal tissue and/or receive signals affected by esophageal tissue; and a data analysis unit configured to: introduce a first signal into esophageal tissue using a first node; determine a value of a characteristic indicative of changes to esophageal tissue at the first node or between a second node and a third node on the same tendril or different tendrils; and provide an indication of changing tissue properties from a determined baseline value. 2. The system of claim 1 , wherein determining the value of the characteristic based on the second signal received using the second node is also based on a third signal received using the third node. 3. The system of claim 1 , wherein the first, second and third nodes are of a first tendril of the one or more tendrils. 4. The system of claim 1 , wherein the third node is the first node. 5. The system of claim 1 , wherein the first node is of a first tendril of the one or more tendrils, and the second node is of a second tendril of the one or more tendrils. 6. The system of claim 1 , wherein the data analysis unit is configured to provide the indication of actual esophageal tissue changes when at least one of the determined values, or a change in the at least one of the determined values, passes a threshold. 7. The system of claim 1 , further comprising an optical fiber having a plurality of temperature sensors, wherein the data analysis unit is further configured to: determine a plurality of temperature values of respective esophageal tissue at respective ones of the plurality of temperature sensors; and provide an indication of potential esophageal tissue changes based on at least one of the determined temperature values. 8. The system of claim 7, wherein the temperature sensors comprise fiber Bragg gratings spaced apart along the optical fiber, with each of the fiber Bragg gratings being configured to reflect a wavelength of light. 9. The system of claim 8, wherein the data analysis unit comprises: a light source configured to emit light into the optical fiber; and a wavelength interrogator configured to determine the wavelengths of light reflected by the plurality of fiber Bragg gratings. 10. The system of claim 7, wherein the temperature sensors utilize Rayleigh scatter to measure temperature along the length of the optical fiber. 11 . The system of claim 7, wherein the temperature sensors comprise thermocouples spaced apart along the length of the tendril. 12. The system of claim 7, wherein a first tendril of the one or more tendrils includes the optical fiber. 13. The system of claim 7, wherein the data analysis unit is configured to provide the indication of potential esophageal tissue changes when at least one of the determined temperature values exceeds a threshold. 14. The system of claim 1 , wherein the data analysis unit is configured to determine the value of the characteristic of esophageal tissue between the second and third nodes based upon one or more differences between the first and second signals. 15. The system of claim 1 , wherein the data analysis unit is configured to determine the value of the characteristic of esophageal tissue between the second and third nodes based upon one or more differences between the second signal and a third signal received using the third node. 16. The system of claim 15, wherein the first signal is a current-controlled signal, the one or more differences include a difference in voltage, and the characteristic is at least one of a bioimpedance, a resistance, a reactance, or a conductance. 17. The system of claim 15, wherein the first signal is a sound wave, the one or more differences includes a propagation time, and the characteristic is a shear-wave velocity. 18. The system of claim 17, wherein the sound wave is an ultrasonic wave. 19. The system of claim 1 , wherein the probe further includes a photoplethysmography (PPG) sensor. 20. The system of claim 1 , further comprising an optical fiber having a plurality of sensors, wherein the data analysis unit is further configured to: determine a plurality of strain values at respective ones of the plurality of sensors; and estimate a location of the optical fiber in the esophagus based on the strain values. 21 . The system of claim 20, wherein the sensors comprise fiber Bragg gratings or Rayleigh scattering-type strain sensors. 22. The system of claim 1 , further comprising electromagnetic sensors configured to be used to triangulate the position of the tendrils in space. 23. The system of claim 1 , wherein a first tendril of the one or more tendrils includes a retraction mechanism configured to move the esophagus away from a cardiac ablation site. 24. The system of claim 23, wherein the retraction mechanism includes an air cavity that, when pressurized in the esophagus, moves the esophagus. 25. The system of claim 23, wherein the retraction mechanism is configured to induce a vacuum in the space between the esophageal wall and the probe to move the tissue away from the heart. 26. The system of claim 1 , wherein a first tendril of the one or more tendrils is pre-strained, and includes an inner core configured to be selectively inserted into or withdrawn from the first tendril by the user to move the esophagus. 27. The system of claim 26, further including an actuator wire adjacent to the core to induce a bending moment on the core. 28. The system of claim 1 , further comprising: a signal processor configured to analyze received signals; and an analog multiplexer to communicate the second signal to the signal processor. 29. The system of claim 1 , wherein the indication is one or both of a visual warning and an audio warning. 30. A method for detecting changes in esophageal tissue properties due to cardiac ablation, the method comprising: deploying a probe in a patient’s esophagus, wherein the probe includes a sheath and a plurality of tendrils extendible by a user from an end of the sheath, wherein the plurality of tendrils are configured to contact respective esophageal tissue at different locations in the esophagus when extended from the end of the sheath, wherein each tendril of the plurality of tendrils includes a plurality of nodes spaced apart along the tendril, and wherein each node of the plurality of nodes is configured to introduce signals into esophageal tissue and/or receive signals affected by esophageal tissue; introducing a first signal into esophageal tissue using a first node; determining, during a cardiac ablation procedure and based on a second signal received using a second node, a value of a characteristic indicative of actual changes to esophageal tissue properties of the patient’s esophagus in proximity to a cardiac ablation site and between the second node and a third node; determining, during the cardiac ablation procedure and based on the value, whether actual esophageal tissue change has occurred; and providing, during the cardiac ablation procedure, an indication of actual esophageal tissue change when determining that actual esophageal tissue change has occurred. 31 . The method of claim 30, wherein the third node is the first node. 32. The method of claim 30, further comprising, after deploying the probe in the patient’s esophagus, extending the plurality of tendrils from the end of the sheath. 33. The method of claim 30, further comprising determining that actual esophageal tissue change has occurred when the value of the characteristic, or a change in the value of the characteristic, passes a threshold. 34. The method of claim 30, further comprising: deploying an optical fiber in the esophagus, wherein the optical fiber includes a plurality of temperature sensors; determining, during the cardiac ablation procedure, a temperature value of esophageal tissue of the esophagus using a first sensor of the plurality of temperature sensors; and providing an indication of potential esophageal tissue change due to the cardiac ablation procedure based upon the temperature value. 35. The method of claim 34, wherein the probe includes the optical fiber, and wherein the method further comprises, after deploying the probe in the patient’s esophagus, extending the optical fiber and the plurality of tendrils from the end of the sheath. 36. The method of claim 30, wherein determining the value of the characteristic comprises determining a shear-wave velocity of a sound wave propagating between two locations in the esophagus corresponding to the second and third nodes. 37. The method of claim 30, wherein determining the value of the characteristic is based upon one or more differences between the second signal and a third signal received using the third node. 38. The method of claim 37, wherein the first signal is a current-controlled signal introduced between the first node and a fourth node, the difference is a differential voltage, and the characteristic is at least one of a bioimpedance, a resistance, a reactance or a conductance. 39. The method of claim 37, wherein the first signal is a sound wave, the difference is a propagation time, and the characteristic is a shear-wave velocity. 40. The method of claim 30, further comprising activating a retraction mechanism of a first tendril of the plurality of tendrils to move the patient’s esophagus away from the cardiac ablation site. 41 . The method of claim 40, wherein activating the retraction mechanism includes pressurizing an air cavity of the first tendril. 42. The method of claim 40, wherein activating the retraction mechanism includes inducing a vacuum in the space between the esophageal wall and the probe. 43. The method of claim 40, wherein the first tendril is pre-strained, and wherein activing the retraction mechanism includes withdrawing an inner core from the first tendril. 44. The method of claim 30, wherein the indication includes one or both of a visual warning and an audio warning. 45. A non-transitory, machine-readable storage medium comprising instructions that, when executed, cause a system for detecting actual esophageal tissue changes due to cardiac ablation to, after a probe is deployed in a patient’s esophagus, wherein the probe includes a sheath and a plurality of tendrils extendible by a user from an end of the sheath, wherein the plurality of tendrils are configured to contact respective esophageal tissue at different locations in the esophagus when extended from the end of the sheath, wherein each tendril of the plurality of tendrils includes a plurality of nodes spaced apart along the tendril, and wherein each node of the plurality of nodes is configured to introduce signals into esophageal tissue and/or receive signals affected by esophageal tissue: introduce a first signal into esophageal tissue using a first node; determine, during a cardiac ablation procedure and based on a second signal received using a second node, a value of a characteristic indicative of actual change to esophageal tissue of the patient’s esophagus in proximity to a cardiac ablation site and between the second node and a third node; determine, during the cardiac ablation procedure and based on the value, whether actual esophageal tissue change has occurred; and provide, during the cardiac ablation procedure, an indication of actual esophageal tissue change when determining that actual esophageal tissue change has occurred. 46. The non-transitory, machine-readable storage medium of claim 45, wherein the third node is the first node. 47. The non-transitory, machine-readable storage medium of claim 45, wherein the instructions, when executed, cause the system to determine that actual esophageal tissue change has occurred when the value of the characteristic, or a change in the value of the characteristic, passes a threshold. 48. The non-transitory, machine-readable storage medium of claim 45, wherein the instructions, when executed, cause the system to: determine, during the cardiac ablation procedure, a temperature value of esophageal tissue of the esophagus using at least one of a fiber Bragg grating, a Rayleigh scatter sensor, or a thermocouple of an optical fiber; and provide an indication of potential esophageal tissue change due to the cardiac ablation procedure based upon the temperature value. 49. The non-transitory, machine-readable storage medium of claim 45, wherein the instructions, when executed, cause the system to determine a shear-wave velocity of a sound wave propagating between two locations in the esophagus corresponding to the second and third nodes. 50. The non-transitory, machine-readable storage medium of claim 45, wherein the instructions, when executed, cause the system to determine the value of the characteristic is based upon one or more differences between the second signal and a third signal received using the third node. 51 . The non-transitory, machine-readable storage medium of claim 45, wherein the characteristic is at least one of a bioimpedance, a resistance, a reactance or a conductance. 52. The non-transitory, machine-readable storage medium of claim 45, wherein the characteristic is a shear-wave velocity. | 2022 | University of Michigan | University | NA | https://umich.edu | USA | The patent focuses on a method to identify actual esophageal tissue changes during cardiac ablation using a probe in the esophagus with multiple extendible tendrils, each with multiple nodes. It addresses the safety concern that cardiac ablation can cause esophageal injury without operators detecting it directly. Multiple tendril-based contact points sense esophageal changes in real time. | הפטנט מתמקד בשיטה לזיהוי שינויי רקמה ושטיים אמיתיים במהלך אבלציה לבבית באמצעות גשוש בוושט עם מספר שלוחות הניתנות להרחבה, כל אחת עם מספר צמתים. הוא מתמודד עם דאגת הבטיחות שאבלציה לבבית יכולה לגרום לפגיעה ושטית מבלי שמפעילים יזהו אותה ישירות. נקודות מגע מרובות מבוססות שלוחות חשות שינויים ושטיים בזמן אמת. | 1 | 1 | 1 | [24 empty] | 1 |  |  |  |  | UNIV MICHIGAN | University of Michigan | UNIV MICHIGAN | University of Michigan | University of Michigan | 2023-04-17 | 2023 | 2023-10-26 | 2023 | 2022-04-18 | 2022 | Application | Utility Invention | inactive | WO2023US18785 | 322869205 | US20250261933A1 | WO
448 | WO2025254964A1 | SYSTEMS AND METHODS FOR THYROID NODULE ABLATION PLANNING AND TRACKING | Ablation systems and methods include displaying a representation of a 3D model of tissue, tracking movement of an ablation probe during an ablation pass, determining a result of the ablation pass based on the tracking and an energy setting(s), and displaying a portion of the tissue ablated as a result of the ablation pass and/or a portion of the tissue remaining after the ablation pass. Ablation planning systems and methods include generating a 3D segmentation of tissue, generating a 3D model of the tissue based on the 3D segmentation, displaying a representation of the 3D model, determining a planned ablation pass for moving an energized ablation probe through the tissue, determining a result of executing the planned ablation pass, and displaying one of a planned portion of the tissue to be ablated based on the determined result a planned portion of the tissue remaining after the determined result. | 1. WHAT IS CLAIMED IS: 1. A system for thyroid nodule ablation, comprising: an ablation probe configured to be energized to ablate tissue; a display; at least one processor; and at least one non-transitory computer-readable storage medium storing instructions that, when executed by the at least one processor, cause the at least one processor to: display, on the display, a representation of a 3D model of tissue of interest; track movement of the ablation probe during a first ablation pass wherein the ablation probe is energized according to at least one energy setting and moved through the tissue of interest; determine a result of the first ablation pass based on the movement tracking and the at least one energy setting, the result of the first ablation pass including an ablation volume and a location of the ablation volume within the tissue of interest; and display, on the display, at least one of a portion of the tissue of interest ablated as a result of the first ablation pass or a portion of the tissue of interest remaining after the first ablation pass. 2. The system according to claim 1 , wherein the at least one non-transitory computer-readable storage medium stores instructions that, when executed by the at least one processor, further cause the at least one processor to: track movement of the ablation probe during a second ablation pass wherein the ablation probe is energized according to at least one energy setting and moved through the tissue of interest; determine a result of the second ablation pass based on the movement tracking and the at least one energy setting, the result of the second ablation pass including an ablation volume and a location of the ablation volume within the tissue of interest; and update, as a result of the second ablation pass, the display of the at least one of the portion of the tissue of interest ablated or the portion of the tissue of interest remaining. 3. The system according to claim 2, wherein the at least one non -transitory computer-readable storage medium stores instructions that, when executed by the at least one processor, further cause the at least one processor to: determine, based upon the results of the first and second ablation passes, whether a threshold amount of the tissue of interest has been ablated; and display an indication of the result of the determination. 4. The system according to claim 1, wherein the display of the at least one of the portion of the tissue of interest ablated or the portion of the tissue of interest remaining includes visually distinguishing at least one portion of the displayed 3D model representation from at least one other portion of the displayed 3D model representation. 5. The system according to claim 1, wherein the display of the at least one of the portion of the tissue of interest ablated or the portion of the tissue of interest remaining includes displaying at least one of a percentage of a volume of the tissue of interest that has been ablated or a percentage of a volume of the tissue of interest that remains, respectively. 6. The system according to claim 1, wherein the at least one non-transitory computer-readable storage medium stores instructions that, when executed by the at least one processor, further cause the at least one processor to display, on the displayed 3D model representation, a representation of a position of the ablation probe relative to the tissue of interest. 7. The system according to claim 1, wherein the at least one non-transitory computer-readable storage medium stores instructions that, when executed by the at least one processor, further cause the at least one processor to display, on the displayed 3D model representation, a trajectory projection of the ablation probe relative to the tissue of interest. 8. The system according to claim 1, wherein the at least one non-transitory computer-readable storage medium stores instructions that, when executed by the at least one processor, further cause the at least one processor to display, on the displayed 3D model representation, a navigation route of the ablation probe through the tissue of interest. 9. The system according to claim 1, wherein the at least one non -transitory computer-readable storage medium stores instructions that, when executed by the at least one processor, further cause the at least one processor to display, on the displayed 3D model representation, an indication of a speed of the ablation probe during the first ablation pass relative to a planned speed. 10. The system according to claim 9, wherein: a representation of the ablation probe moving relative to the displayed 3D model representation during the first ablation pass indicates the speed of the ablation probe; and an avatar moving relative to the displayed 3D model representation during the first ablation pass indicates the planned speed. 11. A method of thyroid nodule ablation, comprising: displaying a representation of a 3D model of tissue of interest; tracking movement of the ablation probe during a first ablation pass wherein the ablation probe is energized according to at least one energy setting and moved through the tissue of interest; determining a result of the first ablation pass based on the movement tracking and the at least one energy setting, the result of the first ablation pass including an ablation volume and a location of the ablation volume within the tissue of interest; and displaying at least one of a portion of the tissue of interest ablated as a result of the first ablation pass or a portion of the tissue of interest remaining after the first ablation pass. 12. The method according to claim 11, further comprising: tracking movement of the ablation probe during a second ablation pass wherein the ablation probe is energized according to at least one energy setting and moved through the tissue of interest; determining a result of the second ablation pass based on the movement tracking and the at least one energy setting, the result of the second ablation pass including an ablation volume and a location of the ablation volume within the tissue of interest; and updating, as a result of the second ablation pass, the display of the at least one of the portion of the tissue of interest ablated or the portion of the tissue of interest remaining. 13. The method according to claim 12, further comprising: determining, based upon the results of the first and second ablation passes, whether a threshold amount of the tissue of interest has been ablated; and displaying an indication based on the determination. 14. The method according to claim 11, wherein displaying the at least one of the portion of the tissue of interest ablated or the portion of the tissue of interest remaining includes visually distinguishing at least one portion of the displayed 3D model representation from at least one other portion of the displayed 3D model representation. 15. The method according to claim 11, wherein displaying the at least one of the portion of the tissue of interest ablated or the portion of the tissue of interest remaining includes displaying at least one of a percentage of a volume of the tissue of interest that has been ablated or a percentage of a volume of the tissue of interest that remains, respectively. 16. The method according to claim 11, further comprising displaying, on the displayed 3D model representation, a representation of a position of the ablation probe relative to the tissue of interest. 17. The method according to claim 11, further comprising displaying, on the displayed 3D model representation, a trajectory projection of the ablation probe relative to the tissue of interest. 18. The method according to claim 11, further comprising displaying, on the displayed 3D model representation, a navigation route of the ablation probe through the tissue of interest. 19. The method according to claim 11, further comprising displaying, on the displayed 3D model representation, an indication of a speed of the ablation probe during the first ablation pass relative to a planned speed. 20. The method according to claim 19, wherein: a representation of the ablation probe moving relative to the displayed 3D model representation during the first ablation pass indicates the speed of the ablation probe; and an avatar moving relative to the displayed 3D model representation during the first ablation pass indicates the speed associated with the planned ablation pass. | 2024 | Medtronic | Medical and Healthcare | Large-Size Company | https://www.medtronic.com | USA | The patent focuses on thyroid nodule ablation systems that display a 3D model of tissue, track ablation probe movement, determine ablation pass results from tracking and energy settings, and show ablated/remaining tissue. It addresses the operator's difficulty in tracking ablation coverage of thyroid nodules across multiple passes. The visualization shows ablated and remaining portions for completion confirmation. | הפטנט מתמקד במערכות אבלציה לקשריות בלוטת התריס המציגות מודל תלת-ממדי של רקמה, עוקבות אחר תנועת גשוש האבלציה, קובעות תוצאות מעבר אבלציה ממעקב והגדרות אנרגיה, ומציגות רקמה שעברה אבלציה/נותרה. הוא מתמודד עם הקושי של המפעיל במעקב אחר כיסוי אבלציה של קשריות בלוטת התריס על פני מספר מעברים. ההמחזה מראה חלקים שעברו אבלציה ונותרו לאישור השלמה. | [7 empty] | 1 | 1 | [12 empty] | 1 | [9 empty] | 1 | MEDTRONIC XOMED LLC | Medtronic | MEDTRONIC XOMED LLC | Medtronic | Medtronic | 2025-05-30 | 2025 | 2025-12-11 | 2025 | 2024-06-04 | 2024 | Application | Utility Invention | active | WO2025US31681 | 325072864 | [NA] | WO
449 | EP4272678A1 | ABLATION ASSISTANCE USING NUMERIC MODELING | An example system for use in ablating target tissue includes memory configured to store anatomical and/or physiological information of a patient and processing circuitry communicatively coupled to the memory. The processing circuitry is configured to, based on the anatomical information and/or the physiological information, determine ablation parameters, the ablation parameters including a suggested positioning of at least energy delivery element of at least one catheter and/or an amount of energy to be delivered via the at least one energy delivery element to the target tissue during ablation. The processing circuitry is configured to output, for display, a representation of at least one of a suggested positioning of the at least one energy delivery element during the ablation, a representation of the target tissue, or a representation of the predicted tissue volume that will be ablated after delivery of ablation energy. | 1. A system for use in ablating target tissue of a patient, the system comprising: memory configured to store at least one of anatomical information of a patient or physiological information of the patient and a numerical model; and processing circuitry communicatively coupled to the memory, the processing circuitry being configured to: determine, based on the at least one of the anatomical information of the patient or the physiological information of the patient, ablation parameters using the numerical model, the ablation parameters comprising at least one of a suggested positioning of at least one energy delivery element of at least one catheter or an amount of energy to be delivered via the at least one energy delivery element to the target tissue of the patient during ablation; and output, for display, at least one of a representation of the suggested positioning of the at least one energy delivery element during the ablation, a representation of the target tissue, or a representation of predicted tissue that will be ablated after the delivery of ablation energy via the at least one energy delivery element. 2. The system of claim 1, wherein the numerical model comprises at least one of a machine learning model or an artificial intelligence model. 3. The system of claim 1 or 2, wherein the numerical model is trained on at least one of anatomical information, physiological information, patient information, energy delivered during ablation, or tissue that was ablated during ablation for a plurality of patients. 4. The system of one of the claims 1 to 3, wherein the processing circuitry is further configured to: after delivery of ablation energy via the at least one energy delivery element, determine a tissue that was ablated; and output for display a representation of the tissue that was ablated and/or wherein the processing circuitry is further configured to: after delivery of ablation energy via the at least one energy delivery element, determine a tissue that was ablated; and output for display a representation of surrounding tissue, the surrounding tissue comprising tissue adjacent to the tissue that was ablated and not comprising the tissue that was ablated. 5. The system of one of the claims 1 to 4, wherein the processing circuitry is further configured to: update the stored anatomical information based on the determination of the tissue that was ablated. 6. The system of one of the claims 1 to 5, further comprising the at least one catheter, wherein the at least one catheter comprises at least one of a pulsed field ablation catheter, a cryoablation catheter, a laser ablation catheter, an ultrasonic ablation catheter, a thermal ablation catheter, a carbon ion beam catheter, or a radio frequency ablation catheter, and/or wherein the processing circuitry is further configured to: control delivery of the at least one catheter into an anatomy of the patient via robotics towards the target tissue. 7. The system of one of the claims 1 to 6, further comprising: a user interface configured to display the at least one of the representation of the suggested positioning of the at least one energy delivery element during ablation, the representation of the target tissue, or the representation of the predicted tissue that will be ablated after the delivery of the ablation energy via the at least one energy delivery element, and/or further comprising: a user interface configured to receive a user input of a selection of the target tissue, and wherein the processing circuitry is configured to, in response to receiving the user input of the selection of the target tissue, output for display at least one of the representation of the target tissue or the representation of the predicted tissue that will be ablated after the delivery of the ablation energy via the at least one energy delivery element. 8. The system of one of the claims 1 to 7, wherein the processing circuitry is further configured to: receive the at least one of the anatomical information or the physiological information from one or more of a computed tomography device, a magnetic resonance imaging device, an ultrasound device, a pacing device, an electrophysiology mapping device, or a non-invasive mapping device; and determine a distance from the at least one energy delivery element to the target tissue, wherein the energy to be delivered to the target tissue during ablation is further based on the distance from the at least one energy delivery element to the target tissue. 9. The system of one of the claims 1 to 8, wherein the processing circuitry is further configured to issue an alert when the at least one energy delivery element is within a treatment range of the target tissue. 10. The system of one of the claims 1 to 9, wherein the processing circuitry is further configured control the at least one energy delivery element to automatically begin delivering energy when the at least one energy delivery element is within the treatment range of the target tissue. 11. The system of one of the claims 1 to 10, wherein the processing circuitry is further configured to control the at least one energy delivery element during delivery of ablation energy to adjust the ablation parameters, and/or wherein the processing circuitry is further configured to: control the at least one energy delivery element to deliver non-therapeutic ablative energy; and test, based on the delivery of the non-therapeutic energy, assumptions in the numerical model. 12. A method comprising: determining, by processing circuitry and based on at least one of anatomical information of a patient or physiological information of the patient, ablation parameters using a numerical model, the ablation parameters comprising at least one of a suggested positioning of at least energy delivery element of at least one catheter or an amount of energy to be delivered via the at least one energy delivery element to target tissue of a patient during ablation; and outputting, for display by the processing circuitry, at least one of a representation of the suggested positioning of the at least one energy delivery element during the ablation, a representation of the target tissue or a representation of predicted tissue that will be ablated after the delivery of ablation energy via the at least one energy delivery element. 13. The method of claim 12, wherein the numerical model comprises at least one of a machine learning model or an artificial intelligence model, and/or wherein the numerical model is trained on at least one of anatomical information, physiological information, patient information, energy delivered during ablation, or tissue that was ablated during ablation for a plurality of patients. 14. The method of claim 12 or 13, further comprising: after the delivery of ablation energy via the at least one energy delivery element, determining, by the processing circuitry, a tissue that was ablated; and outputting, by the processing circuitry, for display a representation of the tissue that was ablated. 15. A non-transitory computer-readable storage medium storing instructions that, when executed, cause processing circuitry to: determine, based on the at least one of the anatomical information of the patient or the physiological information of the patient, ablation parameters using the numerical model, the ablation parameters comprising at least one of a suggested positioning of at least one energy delivery element of at least one catheter or an amount of energy to be delivered via the at least one energy delivery element to target tissue of the patient during ablation; and output, for display, a representation of at least one of the suggested positioning of the at least one energy delivery element during the ablation, a representation of the target tissue or a representation of predicted tissue that will be ablated after the delivery of ablation energy via the at least one energy delivery element. | 2022 | Medtronic | Medical and Healthcare | Large-Size Company | https://www.medtronic.com | USA | The patent focuses on a numeric-modeling system that uses anatomical and physiological information of a patient to determine ablation parameters. It addresses the challenge that fixed protocols don't account for patient anatomy/physiology, leading to variable ablation outcomes. The numeric modeling generates suggested parameters tailored to the individual patient. | הפטנט מתמקד במערכת מידול נומרי המשתמשת במידע אנטומי ופיזיולוגי של מטופל לקביעת פרמטרי אבלציה. הוא מתמודד עם האתגר שפרוטוקולים קבועים אינם מתחשבים באנטומיה/פיזיולוגיה של המטופל, מה שמוביל לתוצאות אבלציה משתנות. המידול הנומרי מייצר פרמטרים מוצעים המותאמים למטופל היחיד. | [8 empty] | 1 | [13 empty] | 1 | [7 empty] | 1 |  | MEDTRONIC INC | Medtronic | MEDTRONIC INC | Medtronic | Medtronic | 2023-05-02 | 2023 | 2023-11-08 | 2023 | 2022-05-03 | 2022 | Application | Utility Invention | active | EP20230171115 | 291867583 | US20230355301A1 | EP
450 | WO2024112808A1 | UTILIZATION OF THREE-DIMENSIONAL NAVIGATION TECHNOLOGY DURING LUNG DENERVATION PROCEDURES | A pulmonary treatment system for treatment of target tissue in an airway of a subject, including a catheter assembly comprising an ablation assembly or tip including an expandable member, at least one energy emitter, and a cooling member, wherein the ablation tip is configured to be positioned within the airway of the subject such that expansion of the expandable member enables the at least one energy emitter and cooling member to engage a wall of the airway, the cooling member configured to cool a portion of the surface of the wall of the airway to reduce damage to the airway disposed between the at least one energy emitter and the target tissue, and a control assembly including a display configured to depict a three-dimensional graphical rendering of the airway of the subject for preplanning prior to the procedure, with real-time tracking of the ablation tip during the treatment of the target tissue. | 1. CLAIMSWhat is claimed is: 1. A pulmonary treatment system for treatment of target tissue in an airway of a subject, the system comprising: a catheter assembly comprising an ablation assembly including an expandable member, at least one energy emitter, and a cooling member, wherein the ablation assembly is configured to be positioned within the airway of the subject such that expansion of the expandable member enables the at least one energy emitter and cooling member to engage a wall of the airway, the cooling member configured to cool a portion of the surface of the wall of the airway to reduce damage to the airway disposed between the at least one energy emitter and the target tissue, while the at least one energy emitter delivers energy to the target tissue to create one or more nerve attenuating lesions; and a control assembly a display, the control assembly being configured to display a three- dimensional graphical rendering of the airway of the subject for real-time tracking of the ablation assembly during the treatment of the target tissue. 2. The system of claim 1, further comprising an electromagnetic field generator configured to generate an electromagnetic field sufficient to surround a portion of the subject. 3. The system of claim 1, wherein the three-dimensional graphical rendering of the airway comprises a predefined map generated by a computed tomography scan (CT scan), and wherein the real-time tracking of the ablation assembly is generated by a comparison of the predefined map and real time imaging generated by a bronchoscope or other imaging assembly introduced into the airway during treatment. 4. The system of claim 1, further comprising a sensor positioned on the catheter assembly configured to sense one or more parameters of the airway and/or the ablation assembly. 5. The system of claim 4, wherein the sensor is positioned on the at least one energy emitter, the sensor being configured to determine a position of the at least one energy emitter within a three-dimensional geometric coordinate system. 6. The system of claim 4, wherein the sensor is positioned offset the at least one energy emitter at an offset distance, the sensor being configured to determine a position of the at least one energy emitter within a three-dimensional geometric coordinate system using the offset distance. 7. The system of claim 4, wherein the sensor is further configured to monitor at least one of a temperature, inflation pressure, coolant flow rate, tissue impedance, current, or power associated with the ablation assembly. 8. The system of claim 1, wherein the control assembly is configured to plot a determined position of the ablation assembly in real-time on the three-dimensional graphical rendering of the airway of the subject. 9. The system of claim 1, further comprising one or more reference datum selectively positionable on the subject and configured to follow movements of the subject while inhaling and exhaling, thereby enabling compensation of a determined position of the ablation assembly within the three-dimensional graphical rendering of the airway of the subject. 10. The system of claim 1 , wherein the control assembly is configured to enable one or more virtual waypoints to be positioned within the three-dimensional graphical rendering of the airway of the subject, thereby enabling preplanning of ablation assembly placement within the airway of the subject prior to treatment. 11. The system of claim 10, wherein the control assembly is further configured to limit electrical power to the at least one energy emitter until a sensed position of the ablation assembly corresponds with the one or more virtual waypoints. 12. The system of claim 10, wherein the control system is further configured to automatically provide an indication that a preplanned procedure has been completed at the one or more virtual waypoints during the treatment. 13. A method for pulmonary treatment of target tissue in an airway of a subject, the method comprising: positioning an ablation assembly of a catheter assembly within the airway of the subject, the ablation assembly including an expandable member, at least one energy emitter, and a cooling member, wherein the expandable member enables the at least one energy emitter and cooling member to engage a wall of the airway, the cooling member configured to cool a portion of the surface of the wall of the airway to reduce damage to the airway disposed between the at least one energy emitter and the target tissue, while the at least one energy emitter delivers energy to the target tissue to create one or more nerve attenuating lesions; displaying a three-dimensional graphical rendering of the airway of the subject for preplanning prior to the procedure; and tracking a position of the ablation assembly during the treatment of the target tissue. 14. The method of claim 13, further comprising generating an electromagnetic field sufficient to surround a portion of the subject. 15. The method of claim 13, further comprising sensing one or more parameters of interest in proximity to the ablation assembly. 16. The method of claim 13, further comprising determining a position of the at least one energy emitter within a three-dimensional geometric coordinate system established by an electromagnetic field generator. 17. The method of claim 16, further comprising controlling a level of power delivery to the at least one energy emitter based on the position of the at least one energy emitter. 18. The method of claim 17, wherein controlling the level of power delivery comprises one of lowering the level of power delivery or shutting off power delivery if the position of the ablation assembly is within a predetermined distance from a predetermined position of non-target tissue. 19. The method of claim 11, further comprising: positioning a sensor outside of the airway; determining in real-time a distance between the ablation assembly and the sensor outside of the airway; and controlling a level of power delivery to the at least one energy emitter based on the position of the at least one energy emitter. 20. The method of claim 19, wherein controlling the level of power delivery comprises one of lowering the level of power delivery or shutting off power delivery if the distance falls below a predetermined threshold distance. 21. The method of claim 19, wherein positioning a sensor outside of the airway comprises positioning an esophageal assembly having the sensor thereon within an esophagus of the subject, and wherein the distance is a distance measured between the sensor of the esophageal assembly and a sensor coupled to the catheter assembly. 22. The method of claim 11 , further comprising plotting a determined position of the ablation assembly in real-time on the three-dimensional graphical rendering of the airway of the subject. 23. The method of claim 11, further comprising using one or more reference datum to follow movements of the subject while inhaling and exhaling, thereby enabling compensation of a determined position of the ablation assembly within the three-dimensional graphical rendering of the airway of the subject. 24. The method of claim 11, further comprising positioning one or more virtual waypoints within the three-dimensional graphical rendering of the airway of the subject, thereby enabling preplanning of ablation assembly placement within the airway of the subject prior to treatment. 25. The method of claim 23, further comprising limiting electrical power to the at least one energy emitter until a sensed position of the ablation assembly corresponds with the one or more virtual waypoints. 26. The method of claim 23, further comprising providing an indication that a preplanned procedure has been completed at the one or more virtual waypoints during the treatment. 27. A pulmonary treatment system for treatment of target tissue in an airway of a subject, the system comprising: a catheter assembly comprising an ablation assembly including an expandable member, at least one energy emitter, and a cooling member, wherein the ablation assembly is configured to be positioned within the airway of the subject such that expansion of the expandable member enables the at least one energy emitter and cooling member to engage a wall of the airway, the cooling member configured to cool a portion of the surface of the wall of the airway to reduce damage to the airway disposed between the at least one energy emitter and the target tissue, while the at least one energy emitter delivers energy to the target tissue to create one or more nerve attenuating lesions; an electromagnetic field generator configured to generate an electromagnetic field sufficient to surround a portion of the subject; a sensor positioned on the catheter assembly, the sensor being configured to determine a position of the at least one energy emitter within a three-dimensional geometric coordinate system established by the electromagnetic field generator; and a control assembly including a display configured to depict a three-dimensional graphical rendering of the airway of the subject for preplanning prior to the procedure, and to plot a determined position of the ablation assembly in real-time on the three-dimensional graphical rendering of the airway of the subject during the treatment of the target tissue, wherein the control assembly is further configured to enable one or more virtual waypoints to be positioned within the three-dimensional graphical rendering of the airway of the subject, thereby enabling preplanning of ablation assembly placement within the airway of the subject prior to treatment, wherein the control assembly is further configured to limit electrical power to the at least one energy emitter until a sensed position of the ablation assembly corresponds with the one or more virtual waypoints, and wherein the control system is further configured to automatically provide an indication that preplanned procedure has been completed at the one or more virtual waypoints during the treatment. | 2022 | NUVAIRA | Medical and Healthcare | Small-size company | http://www.nuvaira.com/ | USA | The patent focuses on a pulmonary treatment system for lung denervation using a catheter with an expandable member, energy emitter, and cooling member, with 3D navigation guidance. It addresses the limitations of 2D-guided lung denervation procedures where target placement and tissue protection are imprecise. 3D navigation enables accurate ablation in airway nerves. | הפטנט מתמקד במערכת טיפול ריאתי לדה-אינרבציה ריאתית באמצעות קטטר עם אלמנט מתרחב, פולט אנרגיה ורכיב קירור, עם הנחיית ניווט תלת-ממדית. הוא מתמודד עם המגבלות של פרוצדורות דה-אינרבציה ריאתית מונחות דו-ממדית שבהן מיקום מטרה והגנת רקמה אינם מדויקים. ניווט תלת-ממדי מאפשר אבלציה מדויקת בעצבי דרכי הנשימה. | 1 |  |  |  | 1 |  |  | 1 | [18 empty] | 1 | [5 empty] | NUVAIRA INC | NUVAIRA INC | NUVAIRA INC | NUVAIRA INC | NUVAIRA INC | 2023-11-21 | 2023 | 2024-05-30 | 2024 | 2022-11-21 | 2022 | Application | Utility Invention | inactive | WO2023US80777 | 304838952 | CN120569168A¦EP4622574A1¦US20240164842A1 | WO
451 | US20250366910A1 | GENERATOR AND CATHETER FOR TISSUE ABLATION | An ablation treatment catheter may include a first conductive wire formed of a first material. The first conductive wire is configured to be connected to a power generator and to carry a current to deliver ablation energy to a biological tissue. The catheter may include a second wire coupled to the first conductive wire. The second wire formed of a second material. The second wire forming a junction with the first conductive wire. The catheter may include a first electrode comprising an ablation element and a temperature sensor. The ablation element includes part of the first conductive wire and is configured to deliver the ablation energy to the biological tissue. The temperature sensor includes the junction of the first conductive wire and the second wire such that a thermocouple is formed at the junction to measure temperature at the first electrode. | 1. (canceled) 2. A system comprising: a catheter for ablating tissue, the catheter comprising an electrode for delivering ablation energy, the electrode comprising a junction between a first wire and a second wire, the first wire formed of a first material and the second wire formed of a second material different from the first material, the first and second wires capable of withstanding a voltage of 200 volts or above for pulsed-field ablation, wherein the junction is configured to deliver the ablation energy and forms a thermocouple between the first wire and the second wire for temperature sensing at the junction; and a power generator coupled to the catheter for powering the catheter, the power generator comprising a power driver capable of delivering voltage to the catheter in at least two voltage ranges, the at least two ranges comprising a first voltage range for radiofrequency ablation and a second voltage range above 200 volts for pulsed-field ablation. 3. The system of claim 2, wherein the catheter comprises a plurality of electrodes, at least a first electrode capable of delivering an ablation energy independent of a second electrode. 4. The system of claim 3, wherein the plurality of electrodes comprises a plurality of conductive wires and share a common return wire. 5. The system of claim 3, wherein the plurality of electrodes are located at predetermined positions of the catheter and the plurality of electrodes are configured to deliver at different ablation energy levels to form a position differential of ablation among the electrodes. 6. The system of claim 2, further comprising a controller configured to regulate a current based on a thermocouple voltage detected at the thermocouple. 7. The system of claim 2, wherein the first wire, the second wire, and the power driver form circuitry for a current carrying the ablation energy to be delivered. 8. The system of claim 2, wherein the power driver is configured to generate alternating current, and the power driver further comprises a frequency filter circuit configured to filter the alternating current to a target frequency, and whether the electrode of the catheter is configured to deliver the ablation energy via the alternating current. 9. The system of claim 8, wherein the target frequency is higher than 200 kHz. 10. The system of claim 8, wherein the alternating current is a monochromatic frequency current. 11. The system of claim 2, wherein the first material has a first impedance and the second material has a second impedance that is higher than the first impedance. 12. The system of claim 2, wherein the power generator is configured to alternate between the first voltage range and the second voltage range based on a preprogrammed ablation protocol. 13. The system of claim 2, wherein the power generator is configured to operate in a closed-loop system that adjusts the delivered voltage based on temperature feedback received from the thermocouple. 14. The system of claim 2, wherein the power generator is further configured to modulate a frequency of alternating current for carrying the ablation energy for selective tissue depth ablation. 15. The system of claim 2, wherein the power generator further comprises a programmable controller for customizing a sequence of voltage ranges delivered during ablation. 16. A catheter for ablating tissue, the catheter comprising: an electrode for delivering ablation energy, the electrode comprising a junction between a first wire and a second wire, the first wire formed of a first material and the second wire formed of a second material different from the first material, the first and second wires capable of withstanding a voltage of 200 volts or above for pulsed-field ablation, wherein the junction is configured to deliver the ablation energy and forms a thermocouple between the first wire and the second wire for temperature sensing at the junction, wherein the catheter is configured to be powered by a power generator that comprises a power driver capable of delivering voltage to the catheter in at least two voltage ranges, the at least two ranges comprising a first voltage range for radiofrequency ablation and a second voltage range above 200 volts for pulsed-field ablation. 17. The catheter of claim 16, further comprising a plurality of electrodes, at least a first electrode capable of delivering an ablation energy independent of a second electrode. 18. The catheter of claim 17, wherein the plurality of electrodes are located at predetermined positions of the catheter and the plurality of electrodes are configured to deliver at different ablation energy levels to form a position differential of ablation among the electrodes. 19. The catheter of claim 16, wherein the thermocouple of the catheter is configured to serve as a temperature sensor in communication with the power generator to sense the temperature at the junction by voltage. 20. The catheter of claim 16, wherein the first wire and the second wire are configured to be connected to the power driver of the power generator to form circuitry for a current carrying the ablation energy to be delivered. 21. A power generator for powering a catheter that is configured to ablates tissue, the power generator comprising: a power driver capable of delivering voltage to the catheter in at least two voltage ranges, the at least two ranges comprising a first voltage range for radiofrequency ablation and a second voltage range above 200 volts for pulsed-field ablation, and wherein the power generator is configured to power the catheter that comprises an electrode for delivering ablation energy, the electrode comprising a junction between a first wire and a second wire, the first wire formed of a first material and the second wire formed of a second material different from the first material, the first and second wires capable of withstanding a voltage of 200 volts or above for pulsed-field ablation, wherein the junction is configured to deliver the ablation energy and forms a thermocouple between the first wire and the second wire for temperature sensing at the junction. | 2024 | PHYSCADE | AI-driven healthcare company | Small-size company | https://www.accessdata.fda.gov/cdrh_docs/pdf25/K250749.pdf | USA | The patent focuses on a tissue ablation catheter and generator with a first conductive wire connected to a power generator and a second wire of different material coupled to it forming a junction. It addresses the design challenge of connecting different conductor materials in catheters without performance loss. The dissimilar-metal junction design optimizes ablation energy delivery. | הפטנט מתמקד בקטטר אבלציית רקמה ומחולל עם חוט מוליך ראשון המחובר למחולל הספק וחוט שני מחומר שונה המחובר אליו ויוצר צומת. הוא מתמודד עם אתגר העיצוב של חיבור חומרי מוליכים שונים בקטטרים ללא אובדן ביצועים. עיצוב צומת המתכת השונה מבצע אופטימיזציה של מסירת אנרגיית האבלציה. |  |  |  |  | 1 | [5 empty] | 1 | 1 | [9 empty] | 1 | [8 empty] | 1 |  | PHYSCADE INC | PHYSCADE INC | PHYSCADE INC | PHYSCADE INC | PHYSCADE INC | 2024-11-08 | 2024 | 2025-12-04 | 2025 | 2024-06-03 | 2024 | Application | Utility Invention | active | US202418942230 | 324793396 | US12150702B1¦US20250366899A1¦WO2025255029A1 | US
452 | US20250143781A1 | ENERGY DELIVERY SYSTEMS WITH LESION INDEX | Provided herein are systems, devices, and methods for treating tissue of a patient. The system includes a first energy delivery device having a first energy delivery element to be positioned proximate target tissue of the patient, a second energy delivery element, an energy delivery console to provide energy between the first energy delivery element and the second energy delivery element, and a user interface having a display for providing information to a user. The energy provided by the energy delivery console creates one or more electric fields within an energy delivery volume proximate the first energy delivery element, and at least one electric field within the energy delivery volume is sufficient to irreversibly electroporate target tissue within the energy delivery volume. | 1.-70. (canceled) 71. An energy delivery system, comprising: one or more energy delivery devices having an energy delivery portion comprising at least one energy element configured to supply energy to target tissue; and a processor configured to execute an algorithm to calculate a volume-distributed score (VDS) for a volume surrounding the energy-delivery portion of the one or more devices, wherein calculated VDS represents an estimation of therapy effectiveness for a future delivery of energy and/or of therapy effectiveness of an energy delivery already performed. 72. The system of claim 71, wherein the target tissue comprises tissue selected from the group consisting of: cardiac tissue; nerve tissue; vessel wall tissue; organ tissue; brain tissue; lung tissue; kidney tissue; liver tissue; stomach tissue; muscle tissue; and combinations thereof. 73. The system of claim 71, wherein the processor is further configured to render the VDS in a visual form on a display of a user interface. 74. The system of claim 73, wherein the visual form includes the VDS in an elongate form. 75. The system of claim 73, wherein the processor is further configured to limit the display of the VDS to include only VDS values that exceed a threshold. 76. The system of claim 75, wherein the processor is configured to set or change the threshold in response to a user input via the user interface and to visually render the variability of scoring differences on the display in different areas of the target tissue. 77. The system of claim 73, wherein the processor is further configured to display the VDS in a visual form in relation to the target tissue. 78. The system of claim 73, wherein the processor is further configured to display the VDS as at least one projection onto an anatomic model. 79. The system of claim 73, wherein the volume of space represented by the VDS can extend far enough to encompass regions that are highly probable, marginally probably, and/or minimally probable of receiving a therapeutically effective delivery of energy for irreversible electroporation. 80. The system of claim 73, wherein the processor is configured to render a spatial distribution of VDS values in the display using color, opacity, intensity, and/or another variable graphical parameter to indicate differentiation. 81. The system of claim 71, wherein the therapeutically effective delivery of energy causes an alteration of the target tissue, wherein such an alteration can comprise: an electrical alteration of tissue where cells of the tissue are stunned or killed; a biophysical alteration of tissue where the tissue is thickened, tissue is thinned, and/or edema results; a biochemical alteration of tissue; a material alteration of tissue where tissue is stiffened, softened, made porous, and/or made brittle, and/or combinations of these. 82. The system of claim 71, wherein the processor is further configured to calculate the VDS based one or more metric, comprising at least one of electrical field strength; area of contact; angle of contact; force of contact; temperature; impedance; one or more tissue properties; blood properties; and/or combinations of these. 83. The system of claim 71, wherein the processor is further configured to calculate the VDS by taking into account one or more of: a configuration of the at least one energy delivery element; a configuration of an energy waveform and/or other energy delivery parameter; and/or blood geometry; target tissue geometry; and/or properties relative to the one or more delivery devices. 84. The system of claim 71, wherein the algorithm is configured to use a field model that is based on: analytical calculations and/or equations; discrete and/or numerical methods using randomized point sources with modeled boundary conditions; and/or a look-up table and/or estimation and/or optimization algorithm with data fitting based on a priori data including at least one or a simulation, a database, and/or a training dataset. 85. The system of claim 71, wherein the algorithm includes an optimization algorithm that includes statistical models, pattern matching, machine learning, artificial intelligence, and/or other data-driven techniques based on historical data to calculate the VDS. 86. The system of claim 71, wherein the algorithm includes a vulnerability algorithm that is executable to processes a set of actual VDS values of a region to qualitatively label and/or quantitatively display areas in which delivered energy is below a predetermined effectiveness threshold. 87. The system of claim 86, wherein the vulnerability algorithm is executable to process a set of VDS values to identify data features that are deterministically and/or probabilistically ineffective. 88. The system of claim 86, wherein the vulnerability algorithm is executable to search for spatial gaps, edges, channels, and/or spatial heterogeneity in the VDS values. 89. The system of claim 86, wherein the vulnerability algorithm is executable to utilize image processing and/or pattern matching techniques to assess vulnerability. 90. The system of claim 71, wherein the at least one energy delivery element includes at least one energy delivery element disposed on an ablation catheter. | 2021 | ENCHANNEL MEDICAL | Medical and Healthcare | Mid-Size Company | https://enchannel.com/ | China | The patent focuses on energy delivery systems with a lesion index, using a first energy delivery element near target tissue, a second energy delivery element, an energy console, and a user interface. It addresses the operator's need for an intuitive index that summarizes lesion progress during energy delivery. The lesion index gives clinicians a clear, quantitative read of treatment progress. | הפטנט מתמקד במערכות מסירת אנרגיה עם מדד נגע, המשתמשות באלמנט מסירת אנרגיה ראשון ליד רקמת המטרה, אלמנט מסירת אנרגיה שני, קונסולת אנרגיה וממשק משתמש. הוא מתמודד עם הצורך של המפעיל במדד אינטואיטיבי המסכם התקדמות נגע במהלך מסירת אנרגיה. מדד הנגע נותן לרופאים קריאה ברורה וכמותית של התקדמות הטיפול. | 1 |  |  |  |  | 1 | [5 empty] | 1 | [9 empty] | 1 | 1 | [7 empty] | 1 |  | ACUTUS MEDICAL INC | ACUTUS MEDICAL INC | ENCHANNEL MEDICAL, LTD. | ENCHANNEL MEDICAL LTD | ENCHANNEL MEDICAL LTD | 2022-07-27 | 2022 | 2025-05-08 | 2025 | 2021-07-27 | 2021 | Application | Utility Invention | active | US202218291334 | 321107213 | AU2022318897A1¦CN117915848A¦EP4376744A1¦EP4376744A4¦JP2024528886A¦WO2023009586A1 | US
453 | WO2026030782A1 | A MICROWAVE ABLATION CATHETER | The present invention provides a microwave ablation catheter designed for renal nerve ablation, incorporating an advanced cooling system to optimize performance. The catheter features a body with distal and proximal portions, housing a microwave antenna within the distal portion. This antenna, in electrical communication with a microwave generator, delivers microwave energy for tissue ablation. There is a cooling system comprising an external coolant pump that directs cooling fluid from a first internal lumen to a second internal lumen. The fluid contacts the microwave antenna before entering the second lumen. A first thermocouple within the second lumen monitors temperature and communicates with the coolant pump to dynamically adjust the flow rate of the cooling fluid based on sensed temperature thresholds. | 1. THE CLAIMS DEFINING THE INVENTION ARE AS FOLLOWS: 1. A catheter adapted for delivering microwave energy in ablating tissue, the catheter comprising: a catheter body having a distal portion and a proximal portion; a microwave antenna positioned between a first internal lumen and a second internal lumen, and located within the distal portion of the catheter body, wherein the microwave antenna is in electrical communication with a microwave generator to deliver microwave energy to the antenna; an antenna cooling system comprising; an external coolant pump for directing cooling fluid from the first internal lumen to the second internal lumen, wherein the cooling fluid from the first internal lumen contacts the microwave antenna prior to entering the second internal lumen; and a first thermocouple positioned in the second internal lumen and in communication with the external cooling pump, wherein the external cooling pump is configurated to dynamically adjust the flow rate of the cooling fluid to the microwave antenna when predetermined temperature thresholds are sensed by the thermocouple. 2. The catheter of Claim 1, wherein the flow rate of the cooling fluid is between 30 ml/min- lOOml/min. 3. The catheter of Claim 1, wherein the first internal lumen is an internal in-flow lumen and the second internal lumen is an internal out-flow lumen. 4. The catheter of Claim 1, wherein the catheter includes at least one electrode located on the outer periphery of the catheter, wherein the electrode is configured to perform at least one function from the group of: sensing, pacing, mapping, and ablation. 5. The catheter of Claim 1, wherein the catheter further comprises a second thermocouple positioned at the electrode located on the outer periphery of the catheter. 6. The catheter of Claim 1, wherein the first thermocouple is configured to monitor the cooling efficiency of the microwave antenna. 7. The catheter of Claim 5, wherein the second thermocouple is configured to measure the temperature of the blood or tissue. 8. The catheter of Claim 1 , wherein the catheter further comprises a guide wire lumen extending from the proximal end to the distal end, wherein the guide wire lumen allows for a guide wire to extend therethrough. 9. The catheter of Claim 1, wherein the micro wave antenna is compatible with standard coaxial cables having a diameter of about 1.13 mm with an impedance value of about 50 Ohms. 10. The catheter of Claim 1, wherein the antenna cooling system is in fluid communication to a proximal end of the first internal lumen. 11. The catheter of Claim 1, wherein the catheter comprises a lubricious coating to the outer surface of the catheter. 12. The catheter of Claim 1, wherein the microwave ablation catheter operates at a frequency ranging between 300 MHz to 300 GHz, preferably between 2 GHz to 3 GHz, 13. The catheter of Claim 12, wherein the micro wave ablation catheter operates at a frequency of 2.45 GHz 14. The catheter of Claim 1, wherein the micro wave antenna may be selected from the group consisting of monopole, dipole, slotted, or quarter-wavelength antennas 15. The catheter of any one of Claims 1 to 14, wherein the catheter body comprises; a first elongate tubular member having a proximal end portion and distal end portion, and houses a second elongate tubular member and contains plurality of lumens; wherein the second tubular member includes a proximal end portion and a distal end portion extending from the distal end of the first tubular member; a catheter shaft extending from the proximal end to the distal end of the catheter; and a knob-like control mechanism attached to the proximal end of the first tubular member, wherein the knob-like control mechanism is configured to control the movement, deflection, or positioning of the catheter within the body of the subject. 16. The catheter of Claim 15, wherein there includes at least one lumen defined between the proximal end and the distal end of the first elongate member. 17. The catheter of Claim 15, wherein there includes at least one lumen defined between the proximal end and the distal end of the second elongate tubular member. 18. The catheter of Claim 15, wherein catheter comprises a catheter sheath that surrounds and protects the catheter shaft. 19. The catheter of Claim 15, wherein the catheter further comprises: a handle, wherein the handle includes aa proximal end and a distal end and the catheter device extending from the distal end of the handle; an operating mechanism for effecting relative displacement between the first elongate tubular member and the second elongate tubular member of the catheter device. 20. The catheter of Claim 19, wherein the operating mechanism is a knob that is axially slidable on the handle of the catheter body. 21. The catheter of Claim 19, wherein the catheter has a length of about 900 mm. 22. The catheter of Claim 19, wherein the catheter may or may not be deflectable. 23. The system of Claim 19, wherein a guide wire lumen is not required if the catheter is deflectable. 24. The system of Claim 19, wherein the guide wire lumen is required if the catheter is not deflectable. 25. The system of Claim 19, wherein the distal end of the first tubular member and the distal end of the second tubular member are joined through a process of plastic welding. 26. The system of Claim 19, wherein the material of the first tubular member and the second tubular member is Polyether block amide (PEBAX). 27. The system of Claim 19, wherein the tip of the catheter may bend bidirectionally at an angle ranging between 0 to ±90 degrees for an effective ablation procedure. 28. The system of Claim 19, wherein the catheter is configured with an integrated pressure sensor designed for direct measurement of blood pressure. | 2024 | CATHRX | Medical and Healthcare | Small-size company | https://cathrx.com | Australia | The patent focuses on a microwave ablation catheter for renal nerve ablation with an advanced cooling system and a microwave antenna in the distal portion. It addresses the limitation of conventional renal denervation tools where heating risks vessel damage during microwave delivery. The cooling-system-equipped catheter optimizes microwave performance for renal nerve ablation. | הפטנט מתמקד בקטטר אבלציית מיקרוגלים לאבלציית עצב כלייתי עם מערכת קירור מתקדמת ואנטנת מיקרוגלים בחלק הדיסטלי. הוא מתמודד עם המגבלה של כלי דה-אינרבציה כלייתית קונבנציונליים שבהם חימום מסכן בנזק לכלי דם במהלך מסירת מיקרוגלים. הקטטר המצויד במערכת קירור מבצע אופטימיזציה של ביצועי המיקרוגלים לאבלציית עצב כלייתי. |  |  |  |  | 1 | 1 | [6 empty] | 1 | [12 empty] | 1 | [6 empty] | CATHRX LTD | CATHRX LTD | CATHRX LTD | CATHRX LTD | CATHRX LTD | 2025-07-02 | 2025 | 2026-02-12 | 2026 | 2024-08-06 | 2024 | Application | Utility Invention | active | WO2025AU50710 | 326553220 | [NA] | WO
454 | WO2026058249A1 | CRYO-ASSISTED IRREVERSIBLE ELECTROPORATION (IRE) ABLATION SYSTEM AND METHOD | The cryo-assisted Irreversible Electroporation (IRE) ablation system includes an IRE ablation generator, a coolant supply, cryo-assisted IRE ablation stylet(s), and a cryo-assisted IRE ablation controller for controlling the IRE ablation generator and the cryo-assisted IRE ablation stylet(s). The ablation stylets include a stylet needle having longitudinally directed strip electrode(s) and a freezing module for cooling a longitudinal directed ice-forming surface for freezing a generally hemi-ellipsoidal volume of target tissue. The controller repeatedly switches the ablation stylets between an IRE ablation energy mode and a non-cryoablation freezing mode whereupon, in an absence of ice formation, each cryo-assisted IRE ablation stylet generates a pre-ice 3D electric field pattern in target tissue and, subsequent to a generally hemi-ellipsoidal ice formation, the ice formation modifies the pre-ice 3D electric field pattern to a post-ice 3D electric field pattern different from the pre-ice 3D electric field pattern. | 1. Claims: 1. A cryo-assisted Irreversible Electroporation (IRE) ablation system comprising: a) an IRE ablation generator for providing at least one burst of IRE ablation energy wherein each at least one burst of IRE ablation energy includes at least one IRE ablation pulse; b) a coolant supply for providing coolant having a coolant temperature between from -5 °C to about 5 °C; c) at least one cryo-assisted IRE ablation stylet including: i) an elongated flexible stylet catheter for connection with the cryo- assisted IRE ablation controller and the coolant supply, and ii) an elongated rigid stylet needle mounted on the elongated flexible stylet catheter for introduction into target tissue in a tumorous body organ or body tissue, the elongated rigid stylet needle having a proximal stylet needle end and a distal stylet needle end, the distal stylet needle end having at least one longitudinally directed strip electrode for generating a 3D electric field pattern in target tissue, and a freezing module including at least one thermoelectric device for cooling a longitudinal directed ice-forming surface circumferentially spaced apart from the at least one longitudinally directed strip electrode for freezing a generally hemi-ellipsoidal volume of target tissue in the immediate vicinity of the longitudinally directed ice-forming surface; and d) a cryo-assisted IRE ablation controller for controlling the IRE ablation generator and each at least one cryo-assisted IRE ablation stylet for repeatedly switching a cryo-assisted IRE ablation stylet between two operation modes: an IRE ablation energy mode for energizing its at least one longitudinally directed strip electrode for delivering IRE ablation energy to target tissue and not operating its freezing module, and a non-cryoablation freezing mode for operating its freezing module for freezing target tissue and not energizing its at least one longitudinally directed strip electrode, whereupon, in an absence of ice formation in target tissue in the immediate vicinity of the cryo-assisted IRE ablation stylet, the cryo-assisted IRE ablation stylet generates a pre-ice 3D electric field pattern therein and, subsequent to generation of a generally hemi-ellipsoidal volume of frozen target tissue in the immediate vicinity of the cryo-assisted IRE ablation stylet, the generally hemi-ellipsoidal volume of frozen target tissue modifies the preice 3D electric field pattern to a post-ice 3D electric field pattern different from the pre-ice 3D electric field pattern. 2. The system according to claim 1 wherein the cryo-assisted IRE ablation controller further controls a short circuit protection module for selectively short circuiting a cryo-assisted IRE ablation stylet’s freezing module from before the IRE ablation energy mode until after the IRE ablation energy mode. 3. The system according to claim 2 wherein a cryo-assisted IRE ablation stylet includes a short circuit protection module under the control of the cryo- assisted IRE ablation controller. 4. The system according to any one of claims 1 to 3 wherein a cryo- assisted IRE ablation stylet’s proximal stylet needle end includes a visual marker for assisting an operator’s rotational orientation of the cryo-assisted IRE ablation stylet in target tissue. 5. The system according to any one of claims 1 to 4 wherein a cryo- assisted IRE ablation stylet’s distal stylet needle end includes an ultra-sound marker for assisting an operator’s rotational orientation of the cryo-assisted IRE ablation stylet in target tissue. 6. The system according to any one of claims 1 to 5 wherein a cryoassisted IRE ablation stylet’s distal stylet needle end includes a temperature sensor for sensing the target tissue’s temperature adjacent the longitudinal directed ice-forming surface. 7. The system according to any one of claims 1 to 6 wherein a cryoassisted IRE ablation stylet’s each at least one longitudinally directed strip electrode subtends an about 90° to about 180° arc angle in a transverse crosssection of the cryo-assisted IRE ablation stylet. 8. The system according to any one of claims 1 to 7 wherein a cryoassisted IRE ablation stylet’s ice-forming surface subtends an about 90° to about 150° arc angle in a transverse cross-section of the cryo-assisted IRE ablation stylet. 9. The system according to any one of claims 1 to 8 for providing coolant to a cryo-assisted IRE ablation stylet at a flow rate of between about 20 ml/minute to 50 ml/minute to maintain its longitudinally directed strip electrode above freezing temperature. 10. A cryo-assisted Irreversible Electroporation (IRE) ablation stylet for use in the cryo-assisted Irreversible Electroporation (IRE) ablation system according to any one of claims 1 to 9. 11. A cryo-assisted IRE ablation controller for use in the cryo-assisted Irreversible Electroporation (IRE) ablation system according to any one of claims 1 to 9. 12. A cryo-assisted Irreversible Electroporation (IRE) ablation method comprising the steps of: a) providing an IRE ablation generator for providing IRE ablation energy at least one burst of IRE ablation energy wherein each at least one burst of IRE ablation energy includes at least one IRE ablation pulse; b) providing a coolant supply for providing coolant having a coolant temperature between from -5°C to about 5°C; c) providing at least one cryo-assisted IRE ablation stylet including: i) an elongated flexible stylet catheter for connection with the cryoassisted IRE ablation controller and the closed coolant circuit; and ii) an elongated rigid stylet needle mounted on the elongated flexible stylet catheter for introduction into target tissue in a tumorous body organ or body tissue, the elongated rigid stylet needle having a proximal stylet needle end and a distal stylet needle end, the distal stylet needle end having at least one longitudinally directed strip electrode for generating a 3D electric field pattern in target tissue, and a freezing module including at least one thermoelectric device for cooling a longitudinal directed ice-forming surface circumferentially spaced apart from the at least one longitudinally directed strip electrode for freezing a generally hemi-ellipsoidal volume of target tissue in the immediate vicinity of the longitudinally directed ice-forming surface; and d) providing a cryo-assisted IRE ablation controller for controlling the IRE ablation generator and each at least one cryo-assisted IRE ablation stylet, the cryo-assisted IRE ablation controller repeatedly switching a cryo- assisted IRE ablation stylet between two operation modes: an IRE ablation energy mode for energizing its at least one longitudinally directed strip electrode for delivering IRE ablation energy to target tissue and not operating its freezing module, and a non-cryoablation freezing mode for operating its freezing module for freezing target tissue and not energizing its at least one longitudinally directed strip electrode, whereupon, in an absence of ice formation in target tissue in the immediate vicinity of the cryo-assisted IRE ablation stylet, the cryo-assisted IRE ablation stylet generates a pre-ice 3D electric field pattern therein and, subsequent to generation of a generally hemi-ellipsoidal volume of frozen target tissue in the immediate vicinity of the cryo-assisted IRE ablation stylet, the generally hemi-ellipsoidal volume of frozen target tissue modifies the preice 3D electric field pattern to a post-ice 3D electric field pattern different from the pre-ice 3D electric field pattern. 13. The method according to claim 12 and further comprising the step of selectively short circuiting a cryo-assisted IRE ablation stylet’s freezing module from before the IRE ablation energy mode until after the IRE ablation energy mode. 14. The method according to either claim 12 or 13 and further comprising the step of providing coolant to a cryo-assisted IRE ablation stylet at a flow rate of between about 20 ml/minute to 50 ml/minute to maintain its longitudinally directed strip electrode above freezing temperature. | 2024 | BERGER ABRAHAM¦HAZAN ORI EMANUEL | NA | NA | NA | NA | The patent focuses on a cryo-assisted IRE ablation system that combines an IRE generator, coolant supply, cryo-assisted IRE ablation stylets, and a controller. It addresses the limitation of pure IRE that may not eliminate all target tissue, where added cryo-cooling boosts effectiveness. Combining IRE pulses with freezing through one stylet broadens treatment effect. | הפטנט מתמקד במערכת אבלציית IRE בעזרת קריו המשלבת מחולל IRE, אספקת נוזל קירור, סטיילטים לאבלציית IRE בעזרת קריו ובקר. הוא מתמודד עם המגבלה של IRE טהור שעלול לא לסלק את כל רקמת המטרה, שבו קירור קריו נוסף מגביר את היעילות. שילוב פולסי IRE עם הקפאה דרך סטיילט אחד מרחיב את אפקט הטיפול. |  |  |  |  | 1 | [6 empty] | 1 |  |  | 1 | [13 empty] | 1 |  |  |  | BERGER ABRAHAM¦HAZAN ORI EMANUEL | BERGER ABRAHAM¦HAZAN ORI EMANUEL | BERGER ABRAHAM¦HAZAN ORI EMANUEL | BERGER ABRAHAM¦HAZAN ORI EMANUEL | BERGER ABRAHAM¦HAZAN ORI EMANUEL | 2025-09-10 | 2025 | 2026-03-19 | 2026 | 2024-09-12 | 2024 | Application | Utility Invention | active | WO2025IL50781 | 326933134 | [NA] | WO
455 | WO2024131363A1 | BIOLOGICAL CAVITY ABLATION SYSTEM, BIOLOGICAL CAVITY ABLATION METHOD, AND STORAGE MEDIUM | A biological cavity ablation system, a biological cavity ablation method, and a storage medium. The biological cavity ablation system comprises an energy module (10), a delivery module (20), and a control module (30). The energy module (10) is used for generating ablation energy directed at a target area (50) of a biological cavity. The delivery module (20) is used for transmitting the ablation energy to the target area (50) and for detecting in real-time the temperature of the target area (50). The control module (30) is used for detecting in real time the parameter values corresponding to the ablation energy generated by the energy module (10). The control module (30), on the basis of the parameter values and the temperature values of the target area (50), regulates in real-time the ablation energy of the energy module (10). This configuration, by means of the detected parameter values of the ablation energy and the temperature of the target area (50) during an ablation treatment process, enables feedback regulation of the ablation energy, providing a function of biological monitoring and feedback. This configuration can increase the safety coefficient of the ablation energy output process, thereby avoiding excessively-high actual temperatures of the target area (50), ensuring that ablation treatment is carried out in a safe and gentle manner, and reducing the chances of complications arising in patients postoperatively. | 1. A biological cavity ablation system, characterized by comprising: An energy module, which is used to generate ablation energy to a target area in the biological cavity; A delivery module, which is used to transmit the ablation energy to the target area and detect the temperature of the target area in real time; A control module is used to detect in real time the parameter value corresponding to the ablation energy generated by the energy module, the parameter value includes a current value and/or a voltage value, the delivery module and the energy module are both connected to the control module, and the control module adjusts the ablation energy of the energy module in real time according to the parameter value and the temperature value of the target area fed back by the delivery module. 2. The biological cavity ablation system according to claim 1, characterized in that the energy module comprises: A signal generating unit, which is used to generate an ablation electrical signal; A signal filtering unit, used for performing low-pass filtering on the ablation electrical signal at a preset frequency; The power amplification unit is used to amplify the ablation electrical signal processed by the signal filtering unit by a preset multiple to obtain the ablation energy. 3. The biocavity ablation system according to claim 1 is characterized in that the delivery module includes a catheter and an ablation head, the proximal end of the catheter is connected to the energy module, the ablation head is arranged at the distal end of the catheter, the ablation head is used to transmit the ablation energy to the target area of the biocavity, and the ablation head is used to detect the temperature of the target area. 4. The biocavity ablation system according to claim 3 is characterized in that the catheter has a curved tube segment, the curved tube segment is connected to the ablation head, and the curved tube segment can be offset relative to the axis of the catheter. 5. The biological cavity ablation system according to claim 3, characterized in that the ablation head comprises: A substrate connected to the conduit, the substrate having a reference surface; A plurality of ablation electrodes, which are installed on the substrate at intervals and are all located on the reference plane, and are electrically connected to the energy module; A temperature detection unit is mounted on the base and located on the reference plane; the temperature detection unit is communicatively connected with the control module. 6. The biological cavity ablation system according to claim 5 is characterized in that the ablation head includes at least two groups of electrode rows arranged side by side at intervals, each group of the electrode rows includes a plurality of the ablation electrodes arranged in a colinear manner, and the temperature detection unit is located between two adjacent groups of the electrode rows. 7. The biocavity ablation system according to claim 1 is characterized in that the control module controls the energy module to stop generating the ablation energy according to preset conditions, and generates an alarm signal according to the preset conditions; the preset conditions are configured as: the current value of the ablation energy is greater than a current threshold, and/or the temperature of the target area is greater than a temperature threshold. 8. The biocavity ablation system according to claim 1 is characterized in that the biocavity ablation system also includes a human-computer interaction module, and the human-computer interaction module includes a touch component and/or a visualization component; the touch component is connected to the control module, and the touch component is triggered to drive the control module to adjust the working state of the energy module; the visualization component includes an endoscope and a display, the endoscope is used to extend into the biocavity to collect endoscopic images of the target area, and the display is used to visualize the endoscopic image. 9. The biological cavity ablation system according to claim 1, characterized in that the control module is configured to execute: Obtain the ablation energy at sampling time tmtom+nThe current value corresponding to each sampling moment in; Calculate the sampling time t respectivelym+1tom+nThe current value corresponding to each sampling time in and the sampling time tmThe corresponding current value difference is obtained, thereby obtaining (n-1) different deviation values; According to (n-1) different deviation values, a current increment is obtained, and based on the current increment and the sampling time tm+nThe corresponding current value is obtained at sampling time tm+n+1The corresponding current value; According to the sampling time tm+n+1The corresponding current value outputs a PWM signal with a corresponding duty cycle, and controls the energy module to output the ablation energy according to the PWM signal; Wherein, n is greater than or equal to 3, m is greater than or equal to 0, and both m and n are integers. 10. The biocavity ablation system according to claim 9, characterized in that when n is equal to 3, the current increment is equal to the sampling time tm+2The corresponding deviation value and sampling time tm+1The difference between the corresponding deviation values; When n is greater than 3, the current increment is equal to the average value of (n-1) different deviation values. 11. The biocavity ablation system according to claim 9 is characterized in that the biocavity ablation system includes a signal routing circuit module connected to both the energy module and the delivery module, the signal routing circuit module divides the ablation energy into at least two signals, one of which goes to the control module so that the control module detects the parameter value of the ablation energy, and the other goes to the delivery module to transmit the ablation energy to the target area. 12. A biological cavity ablation method, characterized by comprising: delivering ablative energy to a target area in the biological cavity; Real-time detection of parameter values corresponding to the ablation energy, and real-time detection of the temperature value of the target area, wherein the parameter values include voltage values and/or current values; The ablation energy is regulated in real time according to the parameter value and the temperature value of the target area. 13. The biological cavity ablation method according to claim 12, characterized in that the step of regulating the ablation energy according to the parameter value comprises: Obtain the ablation energy at sampling time tmtom+nThe current value corresponding to each sampling moment in; Calculate the sampling time t respectivelym+1tom+nThe current value corresponding to each sampling time in and the sampling time tmThe corresponding current value difference is obtained, thereby obtaining (n-1) different deviation values; According to (n-1) different deviation values, a current increment is obtained, and based on the current increment and the sampling time tm+nThe corresponding current value is obtained at sampling time tm+n+1The corresponding current value; According to the sampling time tm+n+1The corresponding current value outputs a PWM signal with a corresponding duty cycle, and controls the energy module to output the ablation energy according to the PWM signal; Wherein, n is greater than or equal to 3, m is greater than or equal to 0, and both m and n are integers. 14. A storage medium having a readable and writable program stored thereon, characterized in that when the program is executed, the biological cavity ablation method as described in claim 12 or 13 is implemented. | 2022 | SHANGHAI MICROPORT ENTPATH MEDTECH | Medical and Healthcare | NA | https://www.microport.com | China | The patent focuses on a biological cavity ablation system with energy, delivery, and control modules that generates and delivers ablation energy to a target area inside a cavity. It addresses the challenge of precisely delivering ablation energy in confined biological cavities where access and energy distribution are limited. The control module orchestrates delivery for accurate cavity-targeted ablation. | הפטנט מתמקד במערכת אבלציה לחלל ביולוגי עם מודולי אנרגיה, מסירה ובקרה המייצרים ומספקים אנרגיית אבלציה לאזור מטרה בתוך חלל. הוא מתמודד עם האתגר במסירה מדויקת של אנרגיית אבלציה בחללים ביולוגיים סגורים שבהם הגישה והתפלגות האנרגיה מוגבלות. מודול הבקרה מתזמר את המסירה לאבלציה מדויקת ממוקדת חלל. | 1 |  |  |  | 1 | [16 empty] | 1 | [8 empty] | 1 |  | SHANGHAI MICROPORT ENTPATH MEDTECH CO LTD | SHANGHAI MICROPORT ENTPATH MEDTECH CO LTD | SHANGHAI MICROPORT ENTPATH MEDTECH CO LTD | SHANGHAI MICROPORT ENTPATH MEDTECH CO LTD | SHANGHAI MICROPORT ENTPATH MEDTECH CO LTD | 2023-11-07 | 2023 | 2024-06-27 | 2024 | 2022-12-23 | 2022 | Application | Utility Invention | inactive | WO2023CN130114 | 313308736 | CN118236148A¦CN118236148B | WO
456 | WO2022147472A1 | INTEGRATED PROBE-ABLATION DEVICE AND SYSTEMS FOR SENSORY NERVE ABLATION IN A JOINT | Integrated probe-ablation provides a combined stimulation and ablation minimally invasive mechanism for delivering both bipolar electrical stimulation and laser ablation of afferent nerve fibers of a desired sensory nerve of a patient under only local anesthesia without having to withdraw the mechanism. A rigid cannula provides minimally invasive access for an integrated probe-ablation device to contact a target nerve tissue proximate a joint of the patient with a distal end having a working end surface that defines a perimeter to contact the target nerve tissue and position a captured portion between the perimeter and a bone structure of the joint. Electrodes proximate the working end surface contact and are controlled provide bipolar electrical stimulation to test by response of the patient whether the target nerve tissue is the desired sensory nerve, and an optical fiber emit laser then is controlled to emit light to ablate the captured portion. | 1. CLAIMS1. An integrated probe-ablation system for electrical stimulation and laser ablation of afferent nerve fibers of a desired sensory nerve of a patient under only local anesthesia, the system comprising: a rigid cannula configured to provide minimally invasive access to a target nerve tissue in a target area proximate a joint of the patient; an integrated probe-ablation device configured to be inserted into the cannula once a distal end of the cannula is positioned proximate the target nerve tissue, the integrated probeablation device having a proximal portion and a distal end and including: a handle at the proximal portion; a working end surface at the distal end, the working end surface defining a generally concave shaped cavity having a perimeter configured to contact the target nerve tissue and position a portion of the target nerve tissue between the perimeter and a bone structure of the joint; and a rigid shaft between the handle and the working end surface; at least a pair of electrical conductors extending through the shaft from the handle to the working end surface with each electrical conductor having an electrode proximate the working end surface that is configured to contact and provide bipolar electrical stimulation to the portion of the target nerve tissue positioned between the perimeter and the bone structure of the joint; at least one optical fiber extending through the shaft from the handle to the working end surface and having a terminus end proximate the working end surface that22 is configured to emit laser light to ablate the portion of the target nerve tissue positioned in the generally concave shaped cavity defined by the working end surface; a control system configured to provide both electrical energy and laser energy to the integrated probe-ablation device without having to withdraw the integrated probe-ablation device from the cannula, the control system including: an electrical stimulation generator configured to be operably coupled to at least the pair of electrical conductors proximate the handle to provide electrical stimulation energy to the electrodes to test whether the target nerve tissue is the desired sensory nerve; and a laser generator configured to be operably coupled to at least the one optical fiber proximate the handle to provide laser light to the terminus end. 2. The integrated prove-ablation system of claim 1 , wherein the joint is the vertebral facetjoint and the bone structure is a concavity formed between a superior articulating process and an adjacent transverse process of the vertebral facet-joint. 3. The integrated prove-ablation system of claim 1, wherein the joint is one of an amphiarthroses joint and a diathroses joint. 4. The integrated prove-ablation system of claim 1, wherein the target nerve tissue is in a spinal column of the patient and the joint is one of a spinal joint, a sacroiliac joint or an intradiscal joint. 5. The integrated prove-ablation system of claim 1, wherein the target nerve tissue is in a joint capsule of the patient and the joint is a diathroses joint. 6. The integrated probe-ablation system of claim 1 , wherein the integrated probe-ablation device further comprises at least one fluid channel extending through the shaft from the handle to an output port proximate the working end surface, and the control system further comprises a fluid delivery system operably connected to the handle to selectively deliver fluid from the output port to the portion of the target nerve tissue. 7. The integrated probe-ablation system of claim 1, wherein a portion of the target nerve tissue extends at least partially into the generally concave shaped cavity when the target nerve tissue is positioned between the perimeter and the bone structure. 8. The integrated probe-ablation system of claim 1 , wherein the integrated probe-ablation device further comprises a temperature sensor at the distal end to monitor a temperature of the target nerve tissue. 9. The integrated probe-ablation system of claim 1, wherein the perimeter of the working end surface includes a rim around the working end surface having a set of protrusions used to facilitate penetration of the working end surface into the target nerve tissue to reduce movement of the working end surface relative to the target nerve tissue. 10. The integrated probe-ablation system of claim 9, wherein the set of protrusions of the rim around the working end surface are formed in a pattern selected from one of a serrated pattern or a scalloped pattern. 11. An integrated probe-ablation device for electrical stimulation and laser ablation of afferent nerve fibers of a desired sensory nerve of a patient under only local anesthesia, the device being configured to be inserted into a rigid cannula configured to provide minimally invasive access to a target nerve tissue in a target area proximate one of an amphiarthroses joint and a diathroses joint of the patient and configured to be operably coupled to an electrical stimulation generator that provides electrical stimulation energy to selectively test whether a target nerve tissue is the desired sensory nerve and to a laser generator that provides laser light to selectively ablate the target nerve tissue, the integrated probe-ablation device comprising: a rigid shaft configured to be selectively insertable into the rigid cannula and having a proximal portion and a distal end with a working end surface at the distal end configured to contact the target nerve tissue and position a portion of the target nerve tissue between the working end surface and a bone structure of the joint; at least a pair of electrical conductors extending through the shaft from the proximal portion to the working end surface with each electrical conductor having an electrode proximate the working end surface that is configured to contact and provide bipolar electrical stimulation to the portion of the target nerve tissue positioned between the working end surface and the bone structure of the joint; and at least one optical fiber extending through the shaft from the proximal portion to the working end surface and having a terminus end proximate the working end surface that is25 configured to emit laser light to ablate the portion of the target nerve tissue positioned between the working end surface and the bone structure of the joint, the integrated prove-ablation device being configured to provide both electrical energy and laser to the target nerve tissue without having to withdraw the integrated probe-ablation device from the cannula. 12. The integrated prove-ablation device of claim 11, wherein the joint is the vertebral facet-joint and the bone structure is a concavity formed between a superior articulating process and an adjacent transverse process of the vertebral facet-joint. 13. The integrated prove-ablation device of claim 11, wherein the target nerve tissue is in a spinal column of the patient and the joint is one of a spinal joint, a sacroiliac joint or an intradiscal joint. 14. The integrated prove-ablation device of claim 11, wherein the target nerve tissue is in a joint capsule of the patient and the joint is a diathroses joint. 15. The integrated prove-ablation device of claim 11, wherein the working end surface defines a generally concave shaped cavity having a perimeter configured to contact the target nerve tissue and position a portion of the target nerve tissue between the perimeter and a bone structure of the joint. 2616. The integrated probe-ablation system of claim 13 , wherein the perimeter of the working end surface includes a rim around the working end surface having a set of protrusions used to facilitate penetration of the working end surface into the target nerve tissue to reduce movement of the working end surface relative to the target nerve tissue. 17. The integrated probe-ablation system of claim 16, wherein the set of protrusions of the rim around the working end surface are formed in a pattern selected from one of a serrated pattern or a scalloped pattern. 18. The integrated probe-ablation system of claim 11, wherein the integrated probeablation device further comprises a temperature sensor at the distal end to monitor a temperature of the target nerve tissue. 19. A method for integrated electrical stimulation and laser ablation of afferent nerve fibers of a desired sensory nerve of a patient under only local anesthesia, the method comprising: inserting a rigid cannula through a skin of the patient to provide minimally invasive access with a distal end of the cannula positioned proximate a target nerve tissue in a target area proximate a joint of the patient; inserting an integrated probe-ablation device having a rigid shaft between a proximal portion and a distal end into the cannula until a working end surface at the distal end contacts the target nerve tissue adjacent a bone structure of the joint and positions a portion of the target nerve tissue between the working end surface and the bone structure of the joint, wherein the integrated probe-ablation devices includes: 27 at least a pair of electrical conductors extending through the shaft from the proximal portion to the working end surface with each electrical conductor having an electrode proximate the working end surface that is configured to contact and provide bipolar electrical stimulation to the portion of the target nerve tissue positioned between the perimeter and the bone structure of the joint; and at least one optical fiber extending through the shaft from the handle to the working end surface and having a terminus end proximate the working end surface that is configured to emit laser light to ablate the portion of the target nerve tissue positioned in the generally concave shaped cavity defined by the working end surface; selectively providing electrical energy to at least the pair of electrical conductors to provide electrical stimulation energy to the electrodes to determine whether the target nerve tissue is the desired sensory nerve based on observation and experience of the patient in response; and within a short period of time after the target nerve tissue is determined to be the desired sensory nerve, and without withdrawing the integrated probe-ablation device from the cannula, selectively providing laser energy to the to at least the one optical fiber to provide laser light energy to the terminus end to ablate the target nerve tissue with laser energy. 20. The method of claim 19, wherein the short period of time is less than two seconds. 21. The method of claim 19, wherein the working end surface at the distal end is maintained in contact with the target nerve tissue adjacent the bone structure of the joint and with the28 portion of the target nerve tissue positioned between the working end surface and the bone structure of the joint during the short period of time. 22. The method of claim 19, wherein selectively providing electrical energy to the at least the pair of electrical conductors provides electrical stimulation energy of between 0-3 volts at a frequency selected from the set consisting of about 50 Hz for sensory simulation and about 2 Hz for motor stimulation to the electrodes for a period of a few seconds, to allow for the operator to observe the presence or absence of the desired response. 23. The method of claim 19, wherein selectively providing laser energy to the to at least the one optical fiber provides laser light energy to generate photocoagulation of the target nerve tissue to form a lesion having a width of at least 5 millimeters and a depth of at least 5 millimeters. 24. An integrated probe-ablation system for combined electrical stimulation and laser ablation of afferent nerve fibers of a desired sensory nerve of a patient under only local anesthesia, the system comprising: a rigid cannula that provides minimally invasive access channel having a diameter of between 1-7 mm to a target nerve tissue in a target area proximate a joint of the patient; an integrated probe-ablation device having a rigid shaft between a proximal portion and a distal end, the distal end having a working end surface that defines a perimeter configured to contact the target nerve tissue and position a captured portion29 of the target nerve tissue between the perimeter and a bone structure of the joint, the device including: at least a pair of electrical conductors that extend through the shaft to corresponding electrodes proximate the working end surface; and at least one optical fiber that extends through the shaft to a terminus end proximate the working end surface, wherein the electrodes are configured to contact and provide bipolar electrical stimulation to the captured portion of the target nerve tissue to perform a test by a response of the patient of whether the target nerve tissue is the desired sensory nerve, and wherein the at least one fiber optic is configured to emit laser light to ablate the captured portion of the target nerve tissue during a short period of time if the response confirms that the target nerve tissue is the desired sensory nerve, and wherein the working end surface at the distal end is configured to be maintained in contact with the target nerve tissue adjacent the bone structure of the joint and with the captured portion of the target nerve tissue positioned between the working end surface and the bone structure of the joint during both test and the short period of time, such that the integrated probe-ablation device can deliver both bipolar electrical stimulation and laser ablation of afferent nerve fibers of the desired sensory nerve of the patient without having to withdraw the integrated probe-ablation device from the cannula; and a control system operably coupled to the integrated probe-ablation device to provide both electrical energy and laser energy to the integrated probe-ablation device. 3025. The system of claim 24, wherein the working end surface defines a generally concave shape within the perimeter that facilitates positioning of the captured portion of the target nerve tissue between the working end surface and the bone structure of the joint such that the captured portion extend at least partially into the generally concave shape. 26. The system of claim 25, wherein the perimeter of the working end surface includes a rim around the working end surface having a set of protrusions used to facilitate penetration of the working end surface into the target nerve tissue to reduce movement of the working end surface relative to the target nerve tissue. 27. The integrated probe-ablation system of claim 26, wherein the set of protrusions of the rim around the working end surface are formed in a pattern selected from one of a serrated pattern or a scalloped pattern. 28. The system of claim 24, wherein the working end surface defines a generally flat shape within the perimeter. 29. The system of claim 24, wherein the at least one fiber optic is a plurality of fiber optics. 30. The system of claim 29, wherein the plurality of fiber optics are arranged in a pattern on the working end surface selected from one of a linear pattern, a circular pattern or a generallyX-shaped pattern. 3131. The system of claim 24, wherein the terminus end of at least one fiber optic is flush with the working surface. 32. The system of claim 24, wherein the terminus end of at least one fiber optic is recessed a distance back from the working surface to provide an effective focal length for delivery of laser energy to the captured portion of the target nerve tissue. 33. The system of claim 24, wherein the terminus end of at least one fiber optic is recessed a distance back from a plane defined by the perimeter of the working surface to provide an effective focal length for delivery of laser energy to the captured portion of the target nerve tissue. 34. The system of claim 24, further comprising one or more optical elements that are operably positioned relative to the terminus end of at least one fiber optic to alter emission of the laser energy. 35. The system of claim 34, wherein the one or more optical elements are selected from the set consisting of a diffusive element or a deflective element. 36. The system of claim 24, wherein the control system is configured to selectively provide electrical energy to the at least the pair of electrical conductors as electrical stimulation energy of 0-3 volts at a frequency selected from the set consisting of about 50 Hz for sensory32 simulation and about 2 Hz for motor stimulation to the electrodes for a period of several seconds. 37. The system of claim 24, wherein the control system is configured to selectively provide laser energy to the to at least the one optical fiber provides laser light to generate photocoagulation of the target nerve tissue to form a lesion having a width of at least 5 millimeters and a depth of up to 5 millimeters. 38. The system of claim 24, wherein the diameter of the cannula is less than 6 mm and a diameter of the perimeter of the working end surface is at least 4 mm. 39. The system of claim 24, wherein the integrated probe-ablation device further comprises a temperature sensor at the distal end to monitor a temperature of the target nerve tissue and wherein the control system selectively provides laser energy in response to the temperature sensor. 33 | 2020 | DMG GROUP | NA | NA | NA | NA | The patent focuses on an integrated probe-ablation device that delivers both bipolar electrical stimulation and laser ablation of afferent nerve fibers under local anesthesia, without withdrawing the device. It addresses the workflow burden of separate stimulation and ablation tools when treating sensory nerves in joints. The integrated design enables stimulation and ablation in one minimally invasive insertion. | הפטנט מתמקד במכשיר משולב גשוש-אבלציה המספק הן גירוי חשמלי דו-קוטבי והן אבלציית לייזר של סיבי עצבים אפרנטיים תחת הרדמה מקומית, ללא משיכת המכשיר. הוא מתמודד עם נטל זרימת העבודה של כלי גירוי ואבלציה נפרדים בעת טיפול בעצבים תחושתיים במפרקים. העיצוב המשולב מאפשר גירוי ואבלציה בהחדרה זעיר-פולשנית אחת. | 1 |  |  |  | 1 | [8 empty] | 1 | [15 empty] | 1 |  |  | DMG GROUP LLC | DMG GROUP LLC | DMG GROUP LLC | DMG GROUP LLC | DMG GROUP LLC | 2021-12-30 | 2021 | 2022-07-07 | 2022 | 2020-12-31 | 2020 | Application | Utility Invention | inactive | WO2021US73195 | 283590972 | [NA] | WO
457 | WO2024129656A2 | SYSTEMS AND METHODS FOR PLANNING AND/OR NAVIGATING TO TREATMENT ZONES IN A MEDICAL PROCEDURE | Systems and methods are described for planning, and/or navigating during, a medical procedure. In one aspect, an expected trajectory of treatment tool that is extendable from a flexible elongate device is determined based on a pose of the flexible elongate device. A projected ablation zone is determined, in relation to a model representing an internal patient anatomy and a target lesion, based on the expected trajectory and an insertion distance of the treatment tool. A display device is caused to display a graphical user interface depicting the expected tool trajectory and the projected ablation zone. In another aspect, a position of a flexible elongate device is determined in relation to a model representing a target lesion. Trajectories along which the treatment tool can approach the target lesion are determined based on a position of the target lesion and the position of the flexible elongate device. | 1. WHAT IS CLAIMED: 1. A method for planning, or navigating during, a medical procedure, the method comprising: obtaining, by one or more processors, a model representing an internal anatomy of a patient and a target lesion; determining, by the one or more processors and based on a pose of a flexible elongate device, an expected tool trajectory of a treatment tool extendable from the flexible elongate device; determining, by the one or more processors and based on the expected tool trajectory and an insertion distance of the treatment tool, a projected treatment zone in relation to the model; and causing, by the one or more processors, a display device to display a graphical user interface depicting (i) the expected tool trajectory, and (ii) the projected treatment zone, in relation to the model. 2. The method of claim 1, wherein the insertion distance comprises a virtual insertion distance indicated by a user. 3. The method of claim 2, wherein the virtual insertion distance is indicated by the user via a virtual insertion distance control on the graphical user interface. 4. The method of claim 3, wherein the virtual insertion distance control comprises a virtual slider. 5. The method of claim 3 or 4, wherein the graphical user interface depicts the expected tool trajectory as a path that extends beyond the virtual insertion distance, and wherein the method further comprises: causing, by the one or more processors, the projected treatment zone to move along the path on the graphical user interface in response to user input via the virtual insertion distance control. 6. The method of any one of claims 2-5, further comprising: limiting, by the one or more processors, the virtual insertion distance indicated by the user to a known maximum insertion distance of the treatment tool. 7. The method of claim 1, wherein the flexible elongate device comprises an actual flexible elongate device, the method further comprising: determining the pose of the flexible elongate device based on shape sensor data generated using a shape sensor of the actual flexible elongate device. 8. The method of claim 7, wherein the shape sensor comprises a fiber-optic shape sensor. 9. The method of claim 7, wherein the shape sensor comprises a plurality of electromagnetic sensors. 10. The method of claim 7, wherein the insertion distance comprises an actual insertion distance. 11. The method of claim 10, further comprising: determining, by the one or more processors, the actual insertion distance using an insertion sensor of the flexible elongate device. 12. The method of claim 7, further comprising: causing, by the one or more processors and in response to detecting that a treatment is performed at a treatment site, the graphical user interface to depict a completed treatment zone at the treatment site in relation to the model. 13. The method of claim 12, further comprising: determining, by the one or more processors, one or more of a size or shape of the completed treatment zone based on one or more of a power or duration associated with the treatment. 14. The method of claim 12, further comprising: determining, by the one or more processors, one or more of a size or shape of the completed treatment zone based on one or more of (i) proximity of the treatment site to one or more critical structures of the patient or (ii) one or more tissue characteristics of the patient at the treatment site. 15. The method of claim 12, further comprising: determining, by the one or more processors, one or more of a size or shape of the completed treatment zone based on one or more of an actual size or shape associated with the treatment at the treatment site. 16. The method of claim 1, wherein the flexible elongate device comprises a virtual flexible elongate device, and wherein the method further comprises: causing, by the one or more processors, the graphical user interface to depict a planned pose of the virtual flexible elongate device. 17. The method of claim 1, further comprising: in response to user input, causing, by the one or more processors, the graphical user interface to change the projected treatment zone to a marked treatment zone that has a fixed position irrespective of further adjustments to the insertion distance. 18. The method of claim 17, wherein the marked treatment zone has a different visual appearance than the projected treatment zone. 19. The method of claim 18, wherein the marked treatment zone has one or more of a different color, different shading, or different pattern than the projected treatment zone. 20. The method of claim 17, wherein the user input is made by a user via a virtual marking control on the graphical user interface. 21. The method of claim 20, wherein the insertion distance comprises a virtual insertion distance indicated by the user via a virtual insertion distance control on the graphical user interface, and wherein the method further comprises: in response to user manipulation of the virtual insertion distance control occurring after the projected treatment zone is changed to the marked treatment zone, causing, by the one or more processors, the graphical user interface to display a new treatment zone in relation to the model. 22. The method of any one of claims 1-21, further comprising: causing, by the one or more processors, the graphical user interface to depict at least a portion of the flexible elongate device. 23. The method of any one of claims 1-211, wherein the projected treatment zone comprises an ellipsoid surrounding a position located at the insertion distance along the expected tool trajectory. 24. The method of any one of claims 1-21, further comprising: determining, by the one or more processors, a major axis of the target lesion; and causing, by the one or more processors, the graphical user interface to depict the major axis of the target lesion. 25. The method of any one of claims 1-21, further comprising: determining, by the one or more processors, a viewing angle for viewing the medical procedure using an intra-operative imaging device. 26. The method of claim 25, further comprising: causing, by the one or more processors, the graphical user interface to depict the viewing angle, wherein the graphical user interface includes a virtual control enabling a user to adjust a viewing angle of the model. 27. The method of any one of claims 1-21, further comprising: causing, by the one or more processors, the graphical user interface to adjust one or more of a size or shape of the projected treatment zone based on proximity of the projected treatment zone to one or more structures within the patient. 28. The method of any one of claims 1-21, further comprising: causing, by the one or more processors, the graphical user interface to adjust one or more of a color, shade, or pattern of the projected treatment zone based on overlap between the projected treatment zone and one or more structures within the patient. 29. The method of any one of claims 1-21, further comprising: causing, by the one or more processors, the graphical user interface to adjust one or more of a size or shape of the projected treatment zone based on user input. 30. The method of claim 21, further comprising: in response to the user input, setting, by the one or more processors, one or more of a power parameter or a duration parameter of medical equipment to a value that corresponds to one or more of the adjusted size or shape of the projected treatment zone. 31. The method of claim 29, further comprising: limiting, by the one or more processors, a maximum size of the projected treatment zone based on a known limitation of the medical equipment. 32. The method of any one of claims 1-21, further comprising: determining, by the one or more processors, a suggested probe trajectory for a next treatment site; and causing, by the one or more processors, the graphical user interface to depict the suggested probe trajectory. 33. The method of any one of claims 1-21, wherein the flexible elongate device comprises a catheter containing the treatment tool. 34. The method of any one of claims 1-21, wherein the medical procedure comprises an ablation procedure, the treatment tool comprises an ablation probe, and the projected treatment zone comprises a projected ablation zone. 35. The method of claim 34, wherein the ablation probe comprises an electroporation probe. 36. The method of claim 34, wherein the ablation probe comprises one or more of a radiofrequency ablation probe, a microwave ablation probe, a cryoablation probe, or a thermal ablation probe. 37. The method of claim 33, wherein the treatment tool comprises an injection tool. 38. The method of any one of claims 1-21, wherein the model comprises a lung airway model. 39. The method of any one of claims 1-21, wherein obtaining the model includes obtaining an initial model based on pre-operative imaging data. 40. The method of claim 39, wherein obtaining the model further includes: updating the initial model based on intra-operative imaging data. 41. The method of any one of claims 1-21, wherein obtaining the model includes obtaining a model based on intra-operative imaging data. 42. A system comprising: one or more processors; a display device; and one or more non-transitory, computer-readable media storing instructions that, when executed by the one or more processors, cause the one or more processors to: obtain a model representing an internal anatomy of a patient and a target lesion; determine, based on a pose of a flexible elongate device, an expected tool trajectory of a treatment tool extendable from the flexible elongate device; determine, based on the expected tool trajectory and an insertion distance of the treatment tool, a projected treatment zone in relation to the model; and cause a display device to display a graphical user interface depicting (i) the expected tool trajectory, and (ii) the projected treatment zone, in relation to the model. 43. The system of claim 42, configured to perform the method of any one of claims 2- 41. 44. One or more non-transitory, computer-readable media storing instructions that, when executed by one or more processors, cause the one or more processors to: obtain a model representing an internal anatomy of a patient and a target lesion; determine, based on a pose of a flexible elongate device, an expected tool trajectory of a treatment tool extendable from the flexible elongate device; determine, based on the expected tool trajectory and an insertion distance of the treatment tool, a projected treatment zone in relation to the model; and cause a display device to display a graphical user interface depicting (i) the expected tool trajectory, and (ii) the projected treatment zone, in relation to the model. 45. The one or more non-transitory, computer-readable media of claim 44, wherein the instructions, when executed by the one or more processors, cause the one or more processors to perform the method of any one of claims 2-41. 46. A method for planning, or navigating during, a medical procedure, the method comprising: obtaining, by one or more processors, a model representing a target lesion within a patient; determining, by the one or more processors, a position of a flexible elongate device in relation to the model; and determining, by the one or more processors and based on a position of the target lesion and the position of the flexible elongate device, a plurality of trajectories along which a treatment tool extendable from the flexible elongate device can approach the target lesion to collectively provide treatment coverage across at least one dimension of the target lesion. 47. The method of claim 46, further comprising: causing, by the one or more processors, a display device to indicate the plurality of trajectories in relation to the model. 48. The method of claim 46, wherein the plurality of trajectories comprises angles that the flexible elongate device can be manipulated to form. 49. The method of claim 46, wherein the plurality of trajectories comprises axes with which the flexible elongate device can be manipulated to align. 50. The method of claim 46, wherein the model includes a first plurality of points representing the target lesion, and wherein determining the plurality of trajectories includes: calculating a second plurality of points by projecting the first plurality of points onto a plane, the plane being at least partially defined by the position of the flexible elongate device; identifying a major axis of a two-dimensional shape defined by the second plurality of points; and determining the plurality of trajectories based on the major axis. 51. The method of claim 50, wherein identifying the major axis includes using principal component analysis to determine the major axis. 52. The method of claim 50, wherein determining the plurality of trajectories includes determining a suggested initial trajectory of the flexible elongate device for the medical procedure. 53. The method of claim 52, wherein determining the suggested initial trajectory includes determining the suggested initial trajectory based on a projected point, of the second plurality of points, that is furthest from a point at a distal portion of the flexible elongate device. 54. The method of claim 53, wherein determining the suggested initial trajectory includes determining the suggested initial trajectory based on (i) a point, of the first plurality of points, that corresponds to the projected point, and (ii) the point at the distal end of the flexible elongate device. 55. The method of claim 52, wherein determining the plurality of trajectories includes determining a suggested one or more subsequent trajectories of the flexible elongate device for the medical procedure. 56. The method of claim 55, wherein determining the suggested one or more subsequent trajectories includes determining the suggested one or more subsequent trajectories based on a predetermined overlap of projected treatment zones and a distance between (i) a distal end of the flexible elongate device and (ii) a furthest point of the target lesion from the distal end of the flexible elongate device. 57. The method of claim 56, wherein determining the plurality of trajectories includes: determining the distance between a distal end of the flexible elongate device and the furthest point of the target lesion by (i) calculating a third plurality of points by projecting at least some of the first plurality of points onto a vector, the vector being at least partially defined by the position of the flexible elongate device, and (ii) selecting a furthest point, of the third plurality of points, from the distal end of the flexible elongate device. 58. The method of claim 57, wherein the vector comprises a pointing direction of the flexible elongate device. 59. The method of claim 57, wherein the vector is defined between (i) a point at the distal end of the flexible elongate device and (ii) a centroid of the target lesion as represented by the first plurality of points. 60. The method of claim 58, wherein the plane is normal to the vector. 61. The method of claim 57, further comprising: determining, by the one or more processors, a viewing angle for the medical procedure based on the major axis and the vector. 62. The method of claim 61, wherein the viewing angle provides an imaging axis that is orthogonal to a plane defined by the major axis and the vector. 63. The method of claim 61, further comprising: causing, by the one or more processors, a display device to indicate the viewing angle. 64. The method of claim 63, further comprising: causing, by the one or more processors and in response to a user activating a control, an intra-operative imaging device to reorient in accordance with the viewing angle. 65. The method of claim 61, further comprising: causing, by the one or more processors and in response to determining the viewing angle, an intra-operative imaging device to reorient in accordance with the viewing angle. 66. The method of any one of claims 46-65, wherein the flexible elongate device comprises an actual flexible elongate device, and wherein determining the position of the flexible elongate device includes determining the position based on shape sensor data generated using a shape sensor of the actual flexible elongate device. 67. The method of any one of claims 46-65, wherein the flexible elongate device comprises a virtual flexible elongate device. 68. The method of any one of claims 46-65, wherein the flexible elongate device comprises a catheter containing the treatment tool. 69. The method of any one of claims 46-65, wherein the medical procedure comprises an ablation procedure and the treatment tool comprises an ablation probe. 70. The method of claim 69, wherein the ablation probe comprises an electroporation probe. 71. The method of claim 69, wherein the ablation probe comprises one or more of a radiofrequency ablation probe, a microwave ablation probe, a cryoablation probe, or a thermal ablation probe. 72. The method of any one of claims 46-65, wherein the model comprises a lung airway model. 73. The method of any one of claims 46-65, wherein obtaining the model includes obtaining an initial model based on pre-operative imaging data. 74. The method of claim 73, wherein obtaining the model further includes: updating the initial model based on intra-operative imaging data. 75. The method of any one of claims 46-65, wherein obtaining the model includes obtaining a model based on intra-operative imaging data. 76. A system comprising: one or more processors; and one or more non-transitory, computer-readable media storing instructions that, when executed by the one or more processors, cause the one or more processors to: obtain a model representing a target lesion within a patient; determine a position of a flexible elongate device in relation to the model; and determine, based on a position of the target lesion and the position of the flexible elongate device, a plurality of trajectories along which a treatment tool extendable from the flexible elongate device can approach the target lesion to collectively provide treatment coverage across at least one dimension of the target lesion. 77. The system of claim 76, configured to perform the method of any one of claims 47-75. 78. One or more non-transitory, computer-readable media storing instructions that, when executed by one or more processors, cause the one or more processors to: obtain a model representing a target lesion within a patient; determine a position of a flexible elongate device in relation to the model; and determine, based on a position of the target lesion and the position of the flexible elongate device, a plurality of trajectories along which a treatment tool extendable from the flexible elongate device can approach the target lesion to collectively provide treatment coverage across at least one dimension of the target lesion. 79. The one or more non-transitory, computer-readable media of claim 78, wherein the instructions, when executed by one or more processors, cause the one or more processors to perform the method of any one of claims 47-75. | 2022 | Intuitive Surgical | Medical and Healthcare | Large-Size Company | https://isrg.intuitive.com/ | USA | The patent focuses on systems that determine the expected trajectory of a treatment tool from a pose of a flexible elongate device, plus a projected ablation zone relative to a model of patient anatomy and target lesion. It addresses the operator's need to plan and visualize ablation zones before delivery to avoid missing the lesion. The projection helps plan and navigate to treatment zones during the procedure. | הפטנט מתמקד במערכות הקובעות את המסלול הצפוי של כלי טיפול מתוך תנוחת מכשיר ארוך גמיש, בתוספת אזור אבלציה צפוי ביחס למודל של אנטומיית מטופל ונגע מטרה. הוא מתמודד עם הצורך של המפעיל לתכנן ולהמחיש אזורי אבלציה לפני המסירה כדי להימנע מהחמצת הנגע. ההקרנה עוזרת לתכנן ולנווט לאזורי טיפול במהלך הפרוצדורה. |  | 1 | [5 empty] | 1 | [15 empty] | 1 |  |  | 1 | [5 empty] | INTUITIVE SURGICAL OPERATIONS | Intuitive Surgical | INTUITIVE SURGICAL OPERATIONS | Intuitive Surgical | Intuitive Surgical | 2023-12-12 | 2023 | 2024-06-20 | 2024 | 2022-12-14 | 2022 | Application | Utility Invention | inactive | WO2023US83510 | 313114498 | CN120358999A¦EP4633509A2¦WO2024129656A3 | WO
458 | WO2024198530A1 | CRYOABLATION PROBE FOR BRONCHIAL VAGAL BLOCKADE AND SURGICAL INSTRUMENT | A cryoablation probe for bronchial vagal blockade and a surgical instrument. The cryoablation probe comprises a handle portion (1), an extension portion (2) and a probe portion (3), wherein the probe portion (3) is connected to the handle portion (1) by means of the extension portion (2); and the probe portion (3) comprises a linear section (31) and a curved section (32), the curved section (32) being connected to the extension portion (2) by means of the linear section (31). A working part of the cryoablation probe is subjected to curved deformation, such that in clinical use, the curved section (32) can better fit to a bronchial wall so as to perform cryoablation on a bronchial vagus nerve at a fitting position. Thus, ablation surgery can be efficiently completed; in addition, nerves on a tracheal wall can be selectively ablated and blocked by adjusting the direction of the curved section (32). Compared with cryoballoons in the prior art, the present invention can not only improve the surgical efficiency, but can also avoid parts needing to be avoided so as to avoid side effects of surgery, thereby improving the surgical effect. | 1. A cryoablation probe for bronchovagal nerve blocking, comprising a handle portion, an extension portion and a probe portion, wherein the probe portion is connected to the handle portion via the extension portion, and is characterized in that: The probe portion includes a straight segment and a curved segment, and the curved segment is connected to the extension portion through the straight segment; Also includes a controller, the controller is configured to adjust the temperature of the probe portion to a second temperature within a preset time when the temperature of the probe portion reaches the first temperature; The second temperature is higher than the first temperature; The axial angle between the plane where the center curve of the curved segment is located and the straight line segment is not less than 45 degrees. 2. The cryoablation probe for bronchovagal blockade according to claim 1 is characterized in that: the controller controls the cooling amount of the fluid introduced into the gas so that the temperature of the probe portion reaches the first temperature from the initial temperature; The relationship between the first temperature, the initial temperature and the fluid cooling capacity can be expressed as: ΔT1=λ*Q1*Δt1, where ΔT1 represents the temperature difference from the first temperature to the initial temperature, λ represents the thermal conductivity of bronchial tissue, Q1 represents the amount of fluid cooling transmitted per unit time, and Δt1 represents the time consumed by the probe from the initial temperature to the first temperature. 3. The cryoablation probe for bronchovagal nerve blocking according to claim 1 is characterized in that: the controller controls the fluid heat of the gas to make the probe temperature reach the second temperature within a preset time; The relationship between the first temperature, the second temperature, the preset time and the fluid heat can be expressed as: ΔT2=λ*Q2*Δt2, wherein ΔT2 represents the temperature difference from the first temperature to the second temperature, λ represents the thermal conductivity of bronchial tissue, Q2 represents the fluid heat transmitted per unit time, and Δt2 represents the preset time. 4. The cryoablation probe for bronchovagal nerve blocking according to claim 1 is characterized in that the probe portion has a clamping space for clamping a bronchoscope. 5. The cryoablation probe for bronchovagal nerve blocking according to claim 4 is characterized in that: the controller is based on the ablation image of the ablation target area collected by the bronchoscope, and the controller is configured to determine the temperature of the probe part according to the ablation image. 6. The cryoablation probe for bronchovagal blockade according to claim 5 is characterized in that: determining the probe temperature according to the ablation image comprises: converting the ablation image into a grayscale image; Calculating the average grayscale value of the grayscale image; The probe portion temperature is determined according to a matching relationship between the average gray value and the tissue temperature. 7. The cryoablation probe for bronchovagal nerve blocking according to claim 1 is characterized in that: the probe portion comprises an outer shell and an inner tube, the outer shell is connected to the extension portion, the inner tube is inserted into the extension portion and the outer shell, and a first channel is provided between the outer shell, the extension portion and the inner tube; A heat exchanger is provided in the handle portion, a second channel is provided between the heat exchanger and the handle portion, the second channel is connected to the first channel, one end of the inner hole of the heat exchanger is connected to the inside of the inner tube, and the other end is connected to the gas delivery pipe. 8. The cryoablation probe for bronchovagal nerve blocking according to claim 5 is characterized in that: the curved segment includes a working segment and a transition segment, the working segment is connected to the straight segment through the transition segment, and the center line of at least part of the working segment is an arc curve. 9. The cryoablation probe for bronchovagal nerve blocking according to claim 1 is characterized in that the maximum outer contour dimension of the curved segment is between 6 and 8 mm. 10. A surgical device, characterized in that it comprises a bronchoscope and a cryoablation probe for bronchovagal nerve blocking as described in any one of claims 1 to 9, wherein the probe portion is fixedly connected to the bronchoscope via a straight segment, and the lens of the bronchoscope faces the curved segment. | 2023 | CHINA JAPAN FRIENDSHIP HOSPITAL CHINA JAPAN INST OF CLINICAL MEDICAL SCIENCES | University | NA | https://www.zryhyy.com.cn/ | China | The patent focuses on a cryoablation probe for bronchial vagal blockade with a handle, extension portion, and probe portion containing a linear and curved section. It addresses the difficulty of cryoablating the vagal nerves around bronchial branches with a straight probe. The curved section conforms the probe to the airway anatomy for targeted vagal blockade. | הפטנט מתמקד בגשוש קריו-אבלציה לחסימת ואגוס סימפוני עם ידית, חלק הארכה וחלק גשוש המכיל מקטע ליניארי ומעוקל. הוא מתמודד עם הקושי בקריו-אבלציה של עצבי הוואגוס סביב הסתעפויות הסימפונות עם גשוש ישר. המקטע המעוקל מתאים את הגשוש לאנטומיית דרכי הנשימה לחסימת ואגוס ממוקדת. |  |  |  |  | 1 |  |  |  | 1 | [5 empty] | 1 | [5 empty] | 1 | [5 empty] | 1 | [5 empty] | CHINA JAPAN FRIENDSHIP HOSPITAL CHINA JAPAN INST OF CLINICAL MEDICAL SCIENCES | CHINA JAPAN FRIENDSHIP HOSPITAL CHINA JAPAN INST OF CLINICAL MEDICAL SCIENCES | CHINA JAPAN FRIENDSHIP HOSPITAL CHINA JAPAN INST OF CLINICAL MEDICAL SCIENCES | CHINA JAPAN FRIENDSHIP HOSPITAL CHINA JAPAN INST OF CLINICAL MEDICAL SCIENCES | CHINA JAPAN FRIENDSHIP HOSPITAL CHINA JAPAN INST OF CLINICAL MEDICAL SCIENCES | 2023-12-14 | 2023 | 2024-10-03 | 2024 | 2023-03-24 | 2023 | Application | Utility Invention | inactive | WO2023CN138711 | 288435001 | CN115998408A¦CN115998408B | WO
459 | WO2025027549A1 | DEVICE FOR TREATING CANCER | A device and method of treating cancerous tissue from a blood vessel, such as a blood vessel proximate a pancreas having pancreatic cancer. The device includes a disruption apparatus and a fixation element fixing a position of the disruption apparatus within the blood vessel. The disruption apparatus is configured to disrupt cancer cells of the pancreatic cancer from within the blood vessel. The method includes delivering the device through the blood vessel to a position adjacent to the cancerous tissue, and disrupting the cancerous tissue by delivering heat or energy to the cancer. Other embodiments are described and claimed. | 1. CLAIMSWhat is claimed is: 1. A catheter configured for insertion into one or more blood vessels proximate a pancreas to treat pancreatic cancer, the catheter comprising: a disruption apparatus being configured to disrupt cancer cells of the pancreatic cancer from within the blood vessel; and a fixation element fixing a position of the disruption apparatus within the blood vessel. 2. A catheter configured for insertion into one or more blood vessels proximate a pancreas having pancreatic cancer, the catheter comprising: a disruption apparatus configured to disrupt cancer cells of the pancreatic cancer. 3. The catheter of claims 1 or 2, wherein the disruption apparatus includes an ablation unit configured to emit ablation energy to ablate the cancer cells. 4. The catheter of claim 3, wherein the ablation unit is an ablation transducer. 5. The catheter of claim 4, wherein the ablation transducer comprises lead zirconate titanate. 6. The catheter of claim 4, wherein the ablation transducer comprises a capacitive micromachined ultrasonic transducer. 7. The catheter of claim 4, wherein the ablation transducer comprises a piezoelectric micromachined ultrasonic transducer. 8. The catheter of claim 3, wherein the ablation unit includes one or more ablation electrodes. 9. The catheter of any one of claims 4-7, wherein the ablation transducer is configured to emit ultrasound energy simultaneously in more than 270 degrees arc to the blood vessel. 10. The catheter of claim 9, wherein the ablation transducer is configured to emit the ultrasound energy simultaneously in a 360 degrees arc to the blood vessel. 11. The catheter of any one of claims 1-10, wherein the disruption apparatus is configured to disrupt cancer cells attached to, wrapped around, and/or spread along a length of the blood vessel. 12. The catheter of claim 3, wherein the ablation energy includes unfocused acoustic energy. 13. The catheter of claim 3, wherein the disruption apparatus further comprises a cavitation unit configured to emit cavitation energy to disrupt the cancer cells. 14. The catheter of claim 13, wherein the cavitation unit is a cavitation transducer. 15. The catheter of any one of claims 13-14, wherein the cavitation unit includes one or more cavitation electrodes. 16. The catheter of claim 15, wherein the one or more cavitation electrodes include a distal cavitation electrode distal to the ablation unit and a proximal cavitation electrode proximal to the ablation unit. 17. The catheter of claim 16, wherein the distal cavitation electrode and the proximal cavitation electrode are separated from one another by at least 5 mm. 18. The catheter of any one of claims 13-17, wherein the ablation unit includes a telescoping transducer slidable over the one or more cavitation electrodes. 19. The catheter of any one of claims 13-18, wherein the cavitation energy includes focused acoustic energy. 20. The catheter of claim 19, wherein the focused acoustic energy has a frequency in a range of 50 kHz to 200 kHz. 21. The catheter of any one of claims 13-20, wherein the ablation unit and the cavitation unit have respective active regions configured to output energy and a separation distance between the active regions is less than 5 mm. 22. The catheter of any one of claims 13-21, further comprising a flexible feature located between the ablation unit and the cavitation unit. 23. The catheter of claim 22, wherein the ablation unit and the cavitation unit each are attached to a backing member, and wherein the flexible feature comprises a portion of the backing member with an opening that spirals around that portion of the backing member. 24. The catheter of claim 23, wherein the flexible feature comprises a flexible boot between and connected to the ablation unit and the cavitation unit. 25. The catheter of any one of claims 1-24, further comprising an imaging unit configured to transmit an imaging signal to detect the cancer cells. 26. The catheter of claim 25, wherein the imaging unit includes an imaging transducer. 27. The catheter of claim 26, wherein the imaging transducer comprises a piezoelectric material. 28. The catheter of any one of claims 25-27, wherein the imaging signal has a frequency of 20 MHz. 29. The catheter any one of claims 25-28, wherein the imaging unit is a portion of an ablation unit of the disruption apparatus. 30. The catheter of any one of claims 1-29, wherein the fixation element comprises: a balloon having an interior; and a catheter shaft having a fluid channel, wherein the interior of the balloon is in fluid communication with the fluid channel. 31. The catheter of claim 30, wherein the balloon is compliant and includes a balloon body having a balloon wall radially surrounding the disruption apparatus. 32. The catheter of claim 31, wherein the compliant balloon is formed from polyether-based thermoplastic polyurethane having a Shore D durometer in a range of 50 to 60. 33. The catheter of claim 32, wherein the polyether-based thermoplastic polyurethane is Pellethane having a Shore D durometer of 55. 34. The catheter of any one of claims 30-33, wherein the balloon includes one or more perforations to deliver a therapeutic agent to the cancer cells. 35. The catheter of any one of claims 1-34, wherein the catheter further comprises a scraping unit configured to remove the cancer cells. 36. The catheter of any one of claims 1-35, wherein the catheter further comprises a collection unit configured to collect the cancer cells. 37. A device configured to treat cancerous tissue from a blood vessel comprising: a catheter shaft having a distal end, a proximal end, and an interior fluid passage; and a disruption apparatus located in the interior, wherein the disruption apparatus is configured to disrupt cancer cells of the cancerous tissue, and wherein the disruption apparatus includes a plurality of transducers connected by a flexible feature configured to facilitate navigating the device through the blood vessel. 38. The device of claim 37, wherein the plurality of transducers include an ultrasound transducer. 39. The device of claim 37 or 38, wherein the disruption apparatus further comprises a cryoprobe. 40. The device of any one of claims 37-39, wherein the disruption apparatus is configured to operate in a focused mode. 41. The device of any one of claims 37-39, wherein the disruption apparatus is configured to operate in an unfocused mode. 42. The device of any one of claims 37-41, further comprising an imaging unit located in the interior. 43. The device of claim 42, wherein the imaging unit includes an imaging transducer. 44. The device of claim 43, wherein the imaging transducer includes one or more elements. 45. The device of claim 44, wherein the one or more elements include a plurality of elements arranged in an array. 46. The device of claim 45, wherein the array is a single cylindrical array. 47. The device of claim 45, wherein the array is a multi-row cylindrical array. 48. The device of claim any one of claims 37-47, further comprising a balloon having an interior in fluid communication with the interior fluid passage, wherein the balloon is configured to receive a cooling fluid circulating through the interior via the interior fluid passage of the device when the disruption apparatus is activated. 49. The device of any one of claim 37-48, further comprising a scraping unit configured to scrape the disrupted cancer cells from the blood vessel. 50. The device of any one of claim 37-49, further comprising a collection unit positioned proximate to the disruption apparatus. 51. A method of treating a patient with cancer, comprising: delivering a device through a blood vessel of the patient to a position to cancerous tissue proximate a cancerous organ; and disrupting the cancerous tissue. 52. The method of claim 51, wherein disrupting the cancerous tissue includes ablating cancer cells. 53. The method of claim 52, wherein disrupting the cancerous tissue includes charring the cancer cells. 54. The method of any one of claims 51-53, wherein ablating the cancer cells includes activating an ablation unit of a catheter to emit unfocused ultrasound energy. 55. The method of any one of claims 51-53, wherein ablating the cancer cells includes activating an ablation unit of a catheter to emit focused ultrasound energy. 56. The method of any one of claims 51-53, wherein ablating the cancer cells includes activating an ablation unit of a catheter to emit radio frequency energy. 57. The method of any one of claims 51-53, wherein ablating the cancer cells includes activating an ablation unit of a catheter to emit cold fluid. 58. The method of any one of claims 51-56, wherein ablating the cancer cells includes raising a temperature of the cancerous tissue to a temperature between 56°C to 100°C. 59. The method of any one of claims 51-58, wherein disrupting the cancerous tissue comprises collapsing the cancerous tissue using cavitation. 60. The method of claim 59, wherein collapsing the cancerous tissue comprises collapsing membranous organelles within the cancerous tissue. 61. The method of claim 59, wherein the cavitation breaks cancer cells of the cancerous tissue into smaller pieces and one or more of leaves a tumor antigen intact or exposes an immunogenic moiety within the tumor antigen to stimulate an antitumor immune system of the patient. 62. The method of claim 61, further comprising capturing the smaller pieces that enter the bloodstream. 63. The method of any one of claims 51-62, wherein the blood vessel comprises one or more of a portal vein, a superior mesenteric vein, a superior mesenteric artery, a celiac artery, a celiac trunk, a splenic artery, or a hepatic artery. 64. The method of any one of claims 51-63, further comprising excising the cancerous tissue from the blood vessel. 65. The method of any one of claims 51-64, further comprising delivering chemotherapy to the blood vessel. 66. The method of any one of claims 51-65, further comprising scraping disrupted cancerous tissue from the blood vessel with a scraping unit. 67. The method of claim 66, wherein the scraping unit is a scalpel. 68. The method of claim 67, wherein the scalpel is a harmonic scalpel. 69. The method of any one of claims 66-68, wherein the scraping unit is located on the device. 70. The method of any one of claims 51-69 further comprising discarding disrupted cancerous tissue with a collection unit by moving the device through the blood vessel in a longitudinal direction. 71. The method of claim 70, wherein the collection unit is positioned proximate to an ablation unit used to disrupt the cancerous tissue. 72. The method of any one of claims 51-71, further comprising imaging the blood vessel during the treatment. 73. The method of claim 72, wherein imaging the blood vessel is done by an imaging unit located at a distal end of the device. 74. The method of any one of claims 51-73, further comprising delivering an anti-cancerous drug to the cancerous tissue. 75. The method of claim 74, wherein the anti-cancerous drug includes an activin inhibitor. 76. The method of claim 75, wherein the anti-cancerous drug includes pegvorhyaluronidase (PEGPH20). 77. The method of any one of claims 51-76, wherein disrupting the cancerous tissue includes triggering a specific cellular anti-cancer immune response. 78. A method of treating a patient having pancreatic cancer, comprising: intravascularly positioning a catheter within a blood vessel of the patient and adjacent to a cancer of the patient; and ablating cancer cells of the cancer by delivering energy to the cancer via the catheter. 79. The method of claim 78, wherein ablating the cancer cells includes charring the cancer cells to ease removal of the cancer cells from the blood vessel. 80. The method of claim 78, wherein the catheter comprises a balloon having an interior, the balloon in proximity to a distal end of the catheter, and an ablation unit inside the balloon; wherein said ablation comprises actuating the ablation unit to emit energy; the method further comprising circulating cooling fluid through the interior of the balloon to inflate the balloon into contact with the walls of the blood vessel, and circulating cooling fluid through the interior of the balloon during the ablating to protect the walls of the blood vessel from heat generated during the ablating. 81. The method of claim 78, wherein the cancer cells are located in the walls of the blood vessel; wherein the catheter comprises a balloon having an interior, the balloon in proximity to a distal end of the catheter, and an ablation unit inside the balloon; wherein the ablating comprises actuating the ablation unit to emit energy; the method further comprising circulating cooling fluid through the interior of the balloon to inflate the balloon an amount that does not bring the balloon into contact with the walls of the blood vessel, and circulating cooling fluid through the interior of the balloon during the ablating. 82. The method of claim 78, wherein the cancer cells are located in the walls of the blood vessel; wherein the catheter comprises an ablation unit in proximity to a distal end thereof; wherein the ablating comprises placing the ablation unit in the blood vessel in contact with blood flowing therethrough, and actuating the ablation unit to emit energy. 83. The method of claim 78, wherein the cancer cells are located in the walls of the blood vessel; wherein the catheter comprises a balloon having an interior, the balloon in proximity to a distal end of the catheter, and an ablation unit proximal to and outside of the balloon; the method further comprising circulating through the interior of the balloon to inflate the balloon into contact with the walls of the blood vessel to occlude the blood vessel before the ablating; and wherein the ablating comprises actuating the ablation unit to emit energy. 84. A method for treating cancer, comprising: delivering a balloon delivery catheter to treatment sites in a plurality of blood vessels adjacent to a pancreas; inflating the balloon delivery catheter at each of the treatment sites; delivering energy at each of the treatment sites, wherein an amount of the delivered energy is effective to disrupt cancer cells adjacent to the treatment sites; deflating the balloon delivery catheter at each of the treatment sites; and withdrawing the balloon delivery catheter from the body lumen. 85. The method of claim 84, wherein the delivered heat or energy is effective to ablate the cancer cells. 86. The method of claim 84, wherein the delivered heat or energy is effective to collapse the cancer cells. 87. The method of claim 84, wherein the method comprises infusing a formulation from the inflated balloon delivery catheter to the vessel wall adjacent to nerves or nerve endings at each of the treatment sites. 88. The method of claim 84, wherein the energy is acoustic energy. 89. The method of claim 84, wherein the energy is RF energy. 90. The method of claim 84, wherein the energy is microwave energy. 91. The method of claim 84, wherein the energy is heat energy. 92. A treatment system comprising a catheter of any of claims 1-36, preferably according to claim 33, and a reservoir containing a therapeutic agent for use in a method of treating cancer in a patient, wherein the catheter delivers the therapeutic agent to the cancerous tissue, such as one or more selected from anticancer agents, therapeutic antibodies, immunotherapeutic agents, nucleic acid molecules, radioisotopes, thymidylate synthase inhibitors, vinca alkaloid agents, cisplatin or other platinum agents and other therapeutic agents described herein. 93. The system for use according to claim 92, wherein microbubbles are used to enhance local delivery of therapeutic agents. 94. The system for use according to claim 92 or 93, wherein the method of treating cancer in a patient includes the method as defined in any of claims 50-82 or as defined in any combination of claims 50-82. 95. The system for use according to any of claims 92-94, wherein the cancer is pancreatic cancer. 96. The system for use according to any of claims 92-94, wherein the therapeutic agent is contained in the reservoir in dissolved form in a treatment solution, wherein treatment solution is delivered to the cancerous tissue. 97. A catheter system comprising: the catheter of any one of claims 1 to 36, wherein the catheter has a guidewire lumen; and a guidewire configured to be accommodated in the guidewire lumen of the catheter. 98. A method for treating cancer, comprising: delivering a balloon delivery catheter to treatment sites in a plurality of blood vessels adjacent to a pancreas, wherein at least one of the treatment sites is a cancer treatment site and at least one of the treatment sites is a pain treatment site; inflating the balloon delivery catheter at each of the treatment sites; delivering energy at each of the cancer treatment sites, wherein an amount of the delivered energy is effective to disrupt cancer cells adjacent to each cancer treatment site, and delivering energy at each of the pain treatment sites, wherein an amount of the delivered energy is effective to ablate at least one nerve adjacent to each pain treatment site; deflating the balloon delivery catheter at each of the treatment sites; and withdrawing the balloon delivery catheter from the body lumen. | 2023 | Otsuka | Medical and Healthcare | Large-Size Company | https://www.otsuka.com | Japan | The patent focuses on a device that treats cancer cells from within a blood vessel, using a disruption apparatus and a fixation element to anchor it in vessels near the pancreas. It addresses the difficulty of treating pancreatic cancer with conventional approaches given the organ's location and surgical challenges. The intravascular disruption approach enables a different route to treat pancreatic cancer cells. | הפטנט מתמקד במכשיר המטפל בתאי סרטן מתוך כלי דם, באמצעות מכשיר שיבוש ואלמנט קיבוע לעיגונו בכלי דם ליד הלבלב. הוא מתמודד עם הקושי בטיפול בסרטן הלבלב בגישות קונבנציונליות בהינתן מיקום האיבר ואתגרים ניתוחיים. גישת השיבוש התוך-וסקולרית מאפשרת מסלול שונה לטיפול בתאי סרטן הלבלב. |  |  | 1 | [25 empty] | 1 |  |  |  | OTSUKA MEDICAL DEVICES CO LTD | Otsuka | OTSUKA MEDICAL DEVICES CO LTD | Otsuka | Otsuka | 2024-07-31 | 2024 | 2025-02-06 | 2025 | 2023-07-31 | 2023 | Application | Utility Invention | inactive | WO2024IB57431 | 318989631 | [NA] | WO
460 | WO2024163399A1 | USER INTERFACE FOR THREE DIMENSIONAL IMAGING AND TREATMENT | A user interface presents a 3D view of the tissue and treatment plan. The 3D view comprises a plurality of transverse images arranged along a one or more longitudinal images. The user adjusts the treatment profile with input to the user interface, and an updated treatment profile is shown on the other views. The user to one or more of zoom, pan, or rotate the 3D view with the treatment profile overlaid on the 3D view, and the treatment profile moves with the 3D view to maintain registration with the 3D view. In some embodiments, a 3D treatment plan is generated in accordance with a plurality of angles between a treatment probe and the one or more tissue structures. An AI algorithm can be used to identify tissue structures and plan the treatment angles and energy delivery in accordance with the tissue structures, which can provide a more customized treatment. | 1. CLAIMSWHAT IS CLAIMED IS: 1. A method of planning a treatment, the method comprising: receiving a plurality of transverse images and one or more longitudinal images of a tissue; generating an arrangement of the plurality of transverse images along the one or more longitudinal images in a three-dimensional (3D) view, the three dimensional view comprising the plurality of transverse images at a plurality corresponding locations along the one or more longitudinal images; overlaying a representation of a three dimensional (3D) treatment profile on the 3D view of the plurality of transverse images and the one or more longitudinal images; and providing to a display of a user interface, the 3D view with the representation overlaid on the one or more longitudinal images and one or more of the plurality of transverse images. 2. The method of claim 1, wherein in the 3D view, the plurality of transverse images comprises a first portion on a first side of the one or more longitudinal images and a second portion on a second side of the one or more longitudinal images, the first portion in front of the one or more longitudinal images and the second portion behind one or more longitudinal images. 3. The method of claim 2, wherein the representation of the 3D treatment profile is overlaid in the 3D view on the first portion and the one or more longitudinal images, with the 3D representation extending along the first portion and the one or more longitudinal images from a common location where the first portion intersects the one or more longitudinal images. 4. The method of claim 2, wherein the one or more longitudinal images comprises an amount of transparency sufficient to view the second portion of one or more of the transverse images through the one or more longitudinal images. 5. The method of claim 4. wherein the representation of the 3D treatment profile is overlaid on the second portion on the second side and the amount of transparency is sufficient to view the representation of the 3D treatment profile on the second side through the one or more longitudinal images. 6. The method of claim 1, wherein the arrangement comprises a three dimensional (3D) arrangement and the user interface comprises an input for a user to one or more of zoom, pan or rotate the 3D arrangement in the 3D view. 7. The method of claim 6, wherein the representation of the 3D treatment profile moves with the 3D arrangement to maintain registration of the 3D treatment profile with the 3D arrangement in response to the user input to one or more of zoom, pan or rotate the 3D representation. 8. The method of claim 7, wherein the user input is configured to rotate the one or more longitudinal images and the plurality of transverse images from a first orientation to a second orientation, the first orientation showing a first portion of the plurality of transverse images in front of the one or more longitudinal images and a second portion of the plurality of transverse images behind the one or more longitudinal images, the second orientation showing the second portion of the plurality of transverse images in front of the one or more longitudinal images and the second portion of the transverse images behind the one or more longitudinal images. 9. The method of claim 7, wherein a representation of a treatment probe is show n extending along the one or more longitudinal images and the representation of the treatment probe moves with the 3D arrangement and the 3D treatment profile to maintain registration of the representation of the treatment probe with the 3D arrangement and the 3D treatment profile. 10. The method of claim 1 , wherein the representation of the 3D treatment profile comprises a three dimensional (3D) treatment volume overlaid on the arrangement. 11. The method of claim 10, wherein the 3D treatment volume extends along the one or more longitudinal images in a longitudinal direction and along one or more transverse images in a direction transverse to the one or more longitudinal images. 12. The method of claim 11, wherein an outer boundary of the 3D treatment volume is shown extending along the one or more longitudinal images transverse to the one or more longitudinal images and along the transverse image in the longitudinal direction. 13. The method of claim 1, wherein a representation of a treatment probe is shown in the 3D view extending in a longitudinal direction. 14. The method of claim 13, wherein the representation comprises an animation of the treatment probe. 15. The method of claim 14, wherein the animation of the treatment probe is configured to move to show an animation of the treatment probe delivering energy from an energy source in accordance with a 3D treatment plan. 16. The method of claim 13, wherein the representation comprises an image of the treatment probe from an imaging device. 17. The method of claim 1. wherein a representation of a 3D treatment profile in the 3D view comprises an intersection of the 3D treatment profile with the one or more longitudinal images and one or more of the plurality of transverse images. 18. The method of claim 1, w herein the plurality of transverse images corresponds to a plurality of predefined anatomical locations of an organ. 19. The method of claim 18, wherein the plurality of predefined anatomical locations comprises anatomical locations of first organ and a second organ in which the first organ and the second organ are visible in the one or more longitudinal images and one or more of the transverse images. 20. The method of claim 19, wherein the plurality of predefined anatomical locations has been identified with an artificial intelligence algorithm. 21. The method of claim 19, wherein the first organ comprises a prostate and the second organ comprises a bladder. 22. The method of claim 21, wherein the plurality of predefined anatomical locations comprises an intravesical prostatic protrusion (IPP). a bladder neck (BL), the a prostate (MID), a verumontanum (VERU), and a peripheral sphincter (P. SPH). 23. The method of claim 1, wherein the user interface is configured to receive a user input identity ing a selected image among the one or more longitudinal images and the plurality of transverse images shown in the 3D view and display the selected image with a 2D view'. 24. The method of claim 23, wherein the user input corresponds to a location of the selected image in the 3D view and a user input at the location with a pointing device or a touch screen display. 25. The method of claim 1. wherein the user interface comprises a plurality of user selectable inputs corresponding to an image to display, the plurality of user selectable inputs comprising a first input to display the 3D view, a second input to display the one or more longitudinal images and a third input for the user to select one or more of the transverse images. 26. The method of claim 1, wherein the user interface is configured to receive a user input selecting a transverse image among the plurality of transverse images and provide a two dimensional (2D) view of the transverse image with the representation overlaid on the 2D view of the transverse image in response to the user input. 27. The method of claim 1, wherein the user interface is configured to receive a user input selecting the one or more longitudinal images and provide a 2D view of the one or more longitudinal images with the representation overlaid on the 2D view of the one or more longitudinal images in response to the user input. 28. The method of claim 1, wherein the plurality of transverse images intersects the one or more longitudinal images at the plurality of corresponding locations. 29. The method of claim 28, wherein the plurality of transverse images comprises volumetric pixels (voxels) overlapping with the one or more longitudinal images at the plurality of corresponding locations. 30. The method of claim 1, wherein each of the plurality of transverse images appears substantially perpendicular to the one or more longitudinal images in the 3D view and optionally to within five degrees of perpendicular. 31. The method of claim 1, wherein the 3D view comprises a perspective view of the plurality7 of transverse images and the one or more longitudinal images. 32. The method of claim 1, wherein the plurality of transverse images and the one or more longitudinal images comprise images from a 3D tomographic image. 33. The method of claim 1 , wherein the plurality of transverse images and the one or more longitudinal images comprise ultrasound images and optionally wherein the ultrasound images comprise images from an ultrasound probe inserted into the patient. 34. The method of claim 1. wherein the one or more longitudinal images comprises a plurality of longitudinal images. 35. The method of claim 34, wherein the plurality of longitudinal images has been generated from a 3D volumetric image of the tissue. 36. The method of claim 35, wherein the plurality of longitudinal images has been generated in response to a plurality of angles of an energy source to treat the tissue. 37. The method of claim 34, wherein the plurality of longitudinal images comprises a first longitudinal image along a first portion of the treatment profile and a second longitudinal image along a second portion of the treatment profile. 38. The method of claim 37, wherein the arrangement shows a portion of the first longitudinal image along the first portion of the treatment profile and a portion of the second longitudinal image along the second portion of the treatment profile. 39. The method of claim 37, wherein the first longitudinal image comprises a first transparency along the first portion of the treatment profile and a second transparency along the second portion of the treatment profile greater than the first transparency to increase visibility of the second longitudinal image along the second portion. 40. The method of claim 39, wherein the second longitudinal image comprises a first transparency along the first portion of the treatment profile and a second transparency the second portion of the treatment profile, the first transparency greater than the second transparency to increase visibility of the first longitudinal image along the first portion of the treatment profile. 41. The method of claim 34, wherein the plurality of longitudinal images comprises a first longitudinal image at a first angle with respect to an axis of rotation and a second longitudinal image at a second angle with respect the axis of translation. 42. The method of claim 41, wherein the first angle corresponds to a first angle of rotation of an energy7 source about the axis and the second angle corresponds to a second angle of rotation of an energy- source and optionally wherein the first angle and the second angle correspond to angles of rotation of an energy- source inserted into the patient. 43. The method of claim 41, yvherein the first longitudinal image extends along a first portion of the treatment profile at the first angle and the second portion of the longitudinal image extends along a second portion of the treatment profile at the second angle. 44. The method of claim 34, wherein the plurality of longitudinal images comprises at least 3 longitudinal images, each at a different angle of rotation with respect to an elongate axis of a treatment probe. 45. A method of generating a treatment plan, the method comprising: receiving a plurality of transverse images of a tissue to be treated; for each of the plurality of transverse images, determining a position of a treatment probe and a boundary of the tissue, the boundary defining an area of the tissue: for each of the plurality of transverse images, determining a first treatment angle from the treatment probe to a first treatment location on a first side of the area in response to the boundary' and a second treatment angle from the treatment probe to a second treatment location on a second side of the area in response to the boundary; and generating the treatment plan in response to the first treatment angle and the second treatment angle of said each of the plurality of transverse images. 46. The method of claim 45, wherein the first treatment angle is oriented to provide a first tissue margin on the first side and the second treatment angle is oriented to provide a second tissue margin on the second side. 47. The method of claim 45, wherein the first treatment angle is selected to provide a first tissue margin thickness at the first treatment angle greater than a penetration depth of the energy' source at the first angle and the second treatment angle is selected to provide a second tissue margin thickness at the second angle greater than a depth of penetration of the energy source at the second angle. 48. The method of claim 46, wherein the first tissue margin corresponds to a first tissue margin angle between the first treatment angle and a first angle of a first margin of the boundary’ on the first side and the second tissue margin corresponds to a second tissue margin angle between the second treatment angle and a second angle of a second margin of the boundary on the second side. 49. The method of claim 48, wherein the first tissue margin angle is within a range from about 1 degree to about 15 degrees and optionally within a range from about 1 degree to about 10 degrees and further optionally within a range from about 2 degrees to about 10 degrees. 50. The method of claim 48, wherein the second tissue margin angle is within a range from about 1 degree to about 15 degrees and optionally within a range from about 1 degree to about 10 degrees and further optionally within a range from about 2 degrees to about 10 degrees. 51. The method of claim 46, wherein the treatment plan corresponds to a first depth of penetration of the energy source at the first treatment angle and the first tissue margin comprises a first margin thickness at the first treatment angle, the first margin thickness greater than the first depth. 52. The method of claim 50, wherein the first margin thickness is greater than the first penetration depth by an amount within a range from about 1 percent to about 15 percent and optionally within a range from about 1 percent to about 10 percent and further optionally within a range from about 2 percent to about 10 percent. 53. The method of claim 46, wherein the treatment plan corresponds to a second depth of penetration of the energy source at the second treatment angle and the second tissue margin comprises a second margin thickness corresponds the first treatment angle, the second margin thickness greater than the second penetration depth. 54. The method of claim 53, wherein the second margin thickness is greater than the second penetration depth by an amount within a range from about 1 percent to about 15 percent and optionally within a range from about 1 percent to about 10 percent and further optionally within a range from about 2 percent to about 10 percent. 55. The method of claim 45, wherein the first treatment angle is determined in response to a first boundary' angle between the treatment probe and the boundary of the tissue on the first side and the second treatment angle is determined in response to a second boundary angle between the treatment probe and the boundary of the tissue on the second side. 56. The method of claim 55, wherein the first treatment angle and the first boundary angle are arranged to provide a first tissue margin between the first treatment angle and the first boundary angle and wherein the second treatment angle and the second boundary' angle are arranged to provide a second tissue margin between the second angle and the second boundary' angle. 57. The method of claim 55, wherein the first boundary angle is determined from the treatment probe to a first margin of the boundary on the first side and the second boundary^ angle is determined from the treatment probe to a second margin of the boundary' on the second side. 58. The method of claim 57, wherein the first boundary angle corresponds to a first projection line from the probe to a first boundary location that is tangential to the boundary on the first side and the second margin of the tissue boundary corresponds to a second projection line from the probe that is tangential to the boundary on the second side. 59. The method of claim 57, wherein the first tissue margin extends between the first margin of the boundary and the first treatment location on the first side and the second tissue margin extends between the second margin of the boundary' and the second treatment location on the second side. 60. The method of claim 45, wherein, for each of the plurality of transverse images, the treatment plan is configured to treat tissue in accordance with a treatment profile corresponding to the first treatment angle and the second treatment angle. 61. The method of claim 60, wherein the treatment profile comprises a variable thickness between the first treatment angle and the second treatment angle. 62. The method of claim 60, wherein the treatment profile comprises a substantially uniform thickness between the first treatment angle and the second treatment angle and optionally wherein the substantially uniform thickness is uniform to within 20 percent. 63. The method of claim 60, wherein the treatment profile is separated from the probe with a gap between the probe and the treatment profile. 64. The method of claim 61, wherein the variable thickness comprises a first thickness corresponding to the first treatment angle, a second thickness corresponding to the second treatment angle and a central thickness at an angle between the first angle and the second angle, the central thickness of the treatment profile greater than the first thickness and the second thickness. 65. The method of claim 64, wherein the treatment plan is configured to treat the first thickness and the second thickness with a first scan of energy from the energy source between the first treatment angle and the second treatment angle and the central thickness with the first scan and a second scan of the energy from the energy source. 66. The method of claim 65, wherein the first scan comprises a first rotational scan of the energy source between the first treatment angle and the second treatment angle, and the second scan comprises a second rotational scan of the energy source, the second rotational scan comprising an angle of rotation less than the first rotational scan. 67. The method of claim 60, wherein the treatment plan is configured to remove tissue at a plurality of thicknesses at a plurality of corresponding angles between the first treatment angle and the second treatment angle. 68. The method of claim 67, wherein the treatment plan is configured to remove tissue to a first depth with a first pass of the energy source at a first angular subtense between the first treatment angle and the second treatment angle and to remove tissue to a second depth with a second pass of the energy source corresponding to a second angular subtense less than the first angular subsense, the second depth greater than the first depth. 69. The method of claim 68, wherein the second angular subtense corresponds to a third location on the first side of the area and a fourth location on the second side of the area. 70. The method of claim 69, wherein the third location is located at or below the first removal depth and the fourth location is located at or below the first removal depth. 71. The method of claim 67, wherein each of the plurality of thicknesses comprises a distance from a proximal location of the boundary to a distal location of the boundary for an angle of a plurality of corresponding treatment angles. 72. The method of claim 67, wherein the treatment plan is configured to adjust one or more of an intensity of energy from the energy source, a power of energy from the energy source, translational velocity of the energy source, a rotational velocity of the energy source, or a number of passes of the energy' source at the plurality of corresponding angles to treat the tissue in accordance with the plurality of thicknesses. 73. The method of claim 60, wherein the treatment profile comprises a plurality of thicknesses corresponding to a plurality of rotation angles of an energy source and each of the plurality of thicknesses is determined for a corresponding treatment angle of the plurality of rotation angles. 74. The method of claim 45, wherein the probe is separated from the boundary by a closest distance at a location of the boundary closest to the probe and by a farthest distance at a location of the boundary' farthest from the probe. 75. The method of claim 45, wherein the first location at the first treatment angle is separated from the probe by a first distance and the second location at the second treatment angle is separated from the probe by a second distance. 76. The method of claim 45, wherein the boundary comprises a boundary of an anatomical tissue structure separated from the probe by a distance and the treatment plan is configured to adjust delivery of the energy from the energy source in response to the distance and the first treatment angle and the second angle from said each of the plurality of transverse images. 77. The method of claim 76, wherein the boundary encloses the anatomical tissue structure and optionally comprises a perimeter around the anatomical tissue structure in said each of the plurality of transverse images. 78. The method of claim 45, wherein the boundary is determined by an artificial intelligence algorithm for said each of the plurality of transverse images. 79. The method of claim 45, wherein the location of the probe corresponds to a location of a directional energy source during treatment for said each of the plurality transverse images and the first angle corresponds to a first angle of the directional energy' source during treatment and the second angle corresponds to a second angle of the energy source during treatment. 80. The method of claim 45, wherein the first angle, the second angle of said each of the plurality of transverse images corresponds to a plurality of treatment angles of a three dimensional (3D) treatment plan. 81. The method of claim 80, wherein the first angle and the second and from said each the plurality of transverse images are combined to generate the 3D treatment plan. 82. The method of claim 80, wherein the plurality of transverse images is located along a treatment axis corresponding to axial translation of a directional energy source on the probe during treatment and said each of the plurality of transverse images is located at a position along the treatment axis. 83. The method of claim 82, wherein the treatment plan is configured to rotate the probe to direct energy' from an energy' source at the first angle to the first location and at the second angle to the second location for said each of the plurality of locations along the treatment axis. 84. The method of claim 83, wherein the treatment plan is configured to rotate the energy' source from the first angle to the second angle to scan the energy' source across the tissue between the first angle and the second angle. 85. The method of claim 45, wherein the treatment plan comprises machine readable instructions to move the probe to a plurality of axial locations along a treatment probe axis and to rotate the probe to a plurality of angles to deliver energy' from an energy source at the plurality of axial locations and the plurality of angles. 86. The method of claim 85, wherein the treatment plan comprises the first angle and the second angle for each of the plurality of locations and is configured to scan the energy source along the tissue between the first angle and the second angle at said each of the plurality of axial locations. 87. The method of claim 85, wherein the treatment plan is configured to adjust one or more of an intensity of energy from the energy source, a power of energy from the energy source, a translation of the probe along the treatment probe axis, a rotation of the probe about the treatment probe axis, a translational velocity of the treatment probe, a rotational velocity of the treatment probe, or a number of passes of the energy' source along the tissue, in order to treat the tissue to different depths in accordance with a treatment profile. 88. The method of claim 87, wherein the treatment plan is configured to fix an axial position of the treatment probe and rotate the treatment probe at the axial position to scan energy from the energy source along the tissue with a first pass between first angles and a second pass between second angles at the axial position. 89. The method of claim 88, wherein a first portion tissue between the first angles does not overlap with the second angles to treat tissue to a first depth with the first pass and wherein the second angles overlap with the first angles to treat tissue to a second depth with the first pass and the second pass. 90. The method of claim 87, wherein the treatment plan is configured to fix a rotational angle of the treatment probe about the elongate probe axis and translate the treatment probe at the rotational angle to scan energy from the energy source along the tissue with a first pass between first translational locations and a second pass between second translational location at the rotational angle. 91. The method of claim 85, wherein the treatment plan is configured to increase a translational velocity of the energy source on the treatment probe to decrease a treatment depth and decrease the translational velocity to increase the treatment depth. 92. The method of claim 85, wherein the treatment plan is configured to increase a rotational velocity' of the energy' source on the treatment probe to decrease a treatment depth and decrease the rotational velocity to increase the treatment depth. 93. The method of claim 45, wherein each of the plurality of transverse image comprises a second tissue and a boundary of the second tissue is determined and an output to a user interface generated in response to a location of the boundary of the second tissue. 94. The method of claim 93, wherein the location of the second tissue boundary is between the first angle and the second angle and the output is generated in response to the location of the second tissue boundary' between the first angle and the second angle. 95. The method of claim 93, wherein the boundary of the first tissue is located between the boundary of the second tissue and the output is generated in response to a distance between the probe and the location of the second tissue boundary. 96. The method of claim 93, wherein the treatment plan comprises an angle of rotation of an energy source beyond a margin of the boundary of the first tissue and the output is generated in response to the angle of rotation beyond the margin. 97. The method of claim 93, wherein the treatment plan corresponds to a depth of penetration of an energy source beyond a thickness of the tissue and the output is generated in response to the depth of penetration greater than the thickness. 98. The method of claim 97, wherein the thickness comprises a plurality of thicknesses at a plurality of rotation angles of an energy source and each of the plurality of thicknesses is determined for a corresponding angle of the plurality of rotation angles. 99. The method of claim 98, wherein said each of the plurality of thicknesses comprises a distance between a first location of the first tissue boundary and a second location of the first tissue boundary at the corresponding angle and the depth of penetration is determined for said each of the plurality of thicknesses. 100. The method of claim 93. wherein the output to the user interface comprises feedback to a user and optionally wherein the feedback comprises one or more of a notification or an alert. 101. The method of claim 93, wherein the output to the user interface is configured for a user to adjust a treatment profile shown on one or more of the plurality of transverse images in response to the feedback. 102. The method of claim 101, wherein the treatment profile substantially encloses the area, and the user interface is configured for the user to adjust the treatment profile with respect to the area by moving the treatment profile as shown on a display for each of the plurality of transverse images. 103. The method of claim 101 , wherein the user interface is configured for the user to adjust one or more of an angle of the treatment profile with respect to the probe or a depth of the treatment profile with respect to the probe. 104. The method of claim 93, wherein the boundary of the first tissue, the boundary of the second tissue, and a treatment profile are shown on each of the plurality of transverse images and the user interface is configured to adjust the treatment profile on the plurality' of transverse images. 105. The method of claim 93, wherein the boundary of the first tissue and the boundary of the second tissue is determined with an Al algorithm. 106. The method of claim 93, wherein the first tissue comprises a first anatomical tissue structure and the second tissue comprises a second anatomical tissue structure different from the first anatomical tissue structure. 107. The method of claim 106, wherein the first anatomical tissue structure comprises an anatomical tissue structure of a first organ and the second anatomical tissue structure comprises a second anatomical tissue structure of a second organ different from the first organ. 108. The method of claim 108, wherein the first anatomical tissue structure comprises one or more of an anatomical structure of the first tissue structure comprises one or more of a tissue wall, a vesicle, a lumen, a wall of a lumen, a bladder, a wall of a bladder, a neck of a bladder, a wall of a bladder neck, a ureteric orifice, an internal urethral orifice, an external urethral sphincter, a ureter, a wall of a ureter, a prostate, a lobe of a prostate, an intravesical prostatic protrusion, a capsule of a prostate, an internal sphincter, and external sphincter, an artery', a wall of an artery', a vein, a wall of a vein, or a lens of an eye, and the second anatomical tissue structure comprises one or more of wherein the anatomical structure of the second tissue structure comprises one or more of a tissue wall, a vesicle, a lumen, a wall of a lumen, a bladder, a wall of a bladder, a neck of a bladder, a wall of a bladder neck, a trigone tissue, a ureteric orifice, an internal urethral orifice, an external urethral sphincter, a ureter, a wall of a ureter, a prostate, a lobe of a prostate, an intravesical prostatic protrusion, a capsule of a prostate, a verumontanum of a prostate, an internal sphincter, and external sphincter, an artery, a wall of an artery, a vein, a wall of a vein, or a retina of an eye. 109. The method of claim 108, wherein the first anatomical tissue structure comprises the intravesical prostatic protrusion of the prostate and the second anatomical tissue structure comprises the trigone of the bladder. 110. The method or apparatus of any of the preceding claims wherein treatment comprises one or more of ablation, resection, or irradiation of the tissue. 111. The method of any of the preceding claims, wherein the method comprises treating the patient. 112. An apparatus, comprising: a processor configured to perform the method of any one of the preceding claims. 113. The method or apparatus of any of the preceding claims wherein the one or more longitudinal images comprises one or more of a sagittal image or a parasagittal image. 114. The method or apparatus of any of the preceding claims, wherein <TRUNCATED> | 2023 | PROCEPT BIOROBOTICS | Medical and Healthcare | Mid-Size Company | http://www.procept-biorobotics.com/ | USA | The patent focuses on a 3D imaging-and-treatment user interface that presents transverse images along longitudinal images, with adjustable treatment profiles and multi-view updates. It addresses the operator's difficulty in planning and adjusting 3D treatment profiles using 2D-only displays. Interactive 3D visualization with synchronized multi-view updates makes treatment planning intuitive. | הפטנט מתמקד בממשק משתמש להדמיה וטיפול תלת-ממדי המציג תמונות רוחב לאורך תמונות אורך, עם פרופילי טיפול הניתנים להתאמה ועדכוני תצוגה מרובה. הוא מתמודד עם הקושי של המפעיל בתכנון והתאמה של פרופילי טיפול תלת-ממדיים באמצעות תצוגות דו-ממדיות בלבד. המחזה תלת-ממדית אינטראקטיבית עם עדכוני תצוגה מרובה מסונכרנים הופכת תכנון טיפול לאינטואיטיבי. |  |  | 1 | [5 empty] | 1 | [13 empty] | 1 | 1 | [6 empty] | 1 |  | PROCEPT BIOROBOTICS CORP | PROCEPT BIOROBOTICS CORP | PROCEPT BIOROBOTICS CORP | PROCEPT BIOROBOTICS CORP | PROCEPT BIOROBOTICS CORP | 2024-01-30 | 2024 | 2024-08-08 | 2024 | 2023-02-01 | 2023 | Application | Utility Invention | inactive | WO2024US13449 | 313952730 | CN120916724A¦EP4658195A1¦JP2026504383A¦US20240252226A1¦US20240252255A1 | WO
461 | WO2024216286A2 | SYSTEMS AND METHODS FOR PULSED FIELD ABLATION | The present technology relates to pulsed field ablation (PFA) and associated systems and methods. Embodiments of the present disclosure include a PFA system that is delivered via a low-profile catheter. The PFA catheter is formed to provide a flexible, adaptive ablation tool under robotic control for targeted and precise delivery of energy to selected tissue regions. This enables enhanced tissue targeting and continuity of generated lesions. The ablation may be delivered by selected pairs of electrodes via, for example, electric pulses. | 1. CLAIMSI/'W e claim: 1. An ablation system for treating target tissues in a body, the ablation system comprising: an ablation catheter comprising an elongate shaft having a wall defining a lumen; a guidewire shaft shaped and sized to extend along the ablation catheter within the lumen; an expansible anchor structure coupled to a distal portion of the guidewire shaft, the anchor structure having a collapsed delivery configuration and an expanded anchoring configuration; an ablation member rotatably disposed at a distal end of the elongate shaft and extending axially and laterally therefrom, the ablation member comprising at least two energy delivering electrodes, the ablation member configured for rotation about an axis defined by the distal portion of the guidewire shaft; a pulse generator configured to generate and deliver electrical pulses to the ablation member, and a controller operably coupled to the pulse generator and to the ablation member, the controller adapted to selectively energize the at least two electrodes in a first mode, and to control rotation of the ablation member. 2. The ablation system of claim 1 wherein the at least two energy’ delivery electrodes are further configured, in a second mode, to measure electrical impedance of tissues adjacent thereto. 3. The ablation system of claim 2 wherein the controller is configured to repeatably alternate operation of the at least two electrodes from the first mode to the second mode. 4. The ablation system of claim 3 wherein the controller is configured to maintain a duration of tissue contact in the first mode considering the electrical impedance measured at the tissues while in the second mode. 5. The ablation system of claim 4 wherein the controller is configured to rotate the ablation member to a new tissue location at a step size considering at least one of (a) the duration of tissue contact, (b) the voltage level of delivered electrical pulses, (c) the electrical impedance measurement, and (d) stored positions that correspond to locations of prior tissue ablations. 6. The ablation system of claim 1 wherein the controller is configured to rotate the ablation member to a new tissue location at a predetermined step size considering at least one of (a) a previous commanded position of the ablation member and (b) the voltage level of delivered electrical pulses. 7. The ablation system of claim 1 wherein, upon rotation of the ablation member from a first tissue contact position to a second tissue contact position, the controller is further configured to position the ablation member such that at least one of the at least two energy delivering electrodes contacts a same portion of treated tissue in the first and second tissue contact positions. 8. The ablation system of claim 1 , further comprising a force sensor operably coupled with the ablation member and the controller, wherein the force sensor is configured to provide a signal indicating a tissue contact force of the ablation member. 9. The ablation system of claim 8 wherein the controller is further configured to adjust an angle and/or a position of the ablation member to alter the tissue contact force toward a target value thereof. 10. The ablation system of claim 1 wherein anchor structure is an expandable balloon. 11. The ablation system of claim 1 wherein anchor structure is a braid structure. 12. A method of ablating target tissues in a body, the method comprising: advancing a guidewire shaft comprising an anchoring structure toward an ostia of a pulmonary vein; deploying the anchoring structure within the ostia to anchor the guidewire; advancing an ablation catheter carrying an extended ablation arm having a pair of electrodes over the guidewire; contacting, with the pair of electrodes, target tissue at a first position; delivering energy- in a first pulse to the pair of electrodes; rotating the ablation arm to a next position adjacent to the prior position; and delivering energy- in a subsequent pulse to the pair of electrodes. 13. The method of claim 12 wherein the target tissue is adjacent an ostia, further comprising repeating rotation of the ablation arm and delivery of energy until a perimeter of tissue adjacent the ostia has received ablative energy-. 14. The method of claim 12, further comprising measuring at least one of a tissue temperature and a tissue impedance during and/or following the delivering energy. 15. The method of claim 14 wherein determining at least one of the next position and the energy of the subsequent pulse is made considering the measured tissue temperature and/or impedance. 16. The method of claim 14 wherein measuring the impedance comprises measuring impedance spectra. 17. The method of claim 14 wherein measuring the tissue temperature and/or the tissue impedance is performed by the pair of electrodes. 18. The method of claim 14 wherein measuring the tissue temperature and/or the tissue impedance is performed by a second pair of electrodes having electrodes adjacent the first pair of electrodes. 19. The method of claim 12, further comprising maintaining a selected tissue contact force during contacting tissue. 20. An ablation catheter for treating target tissues in a body, the ablation catheter comprising: an elongate shaft having a wall defining a lumen; a region of conductive surfaces disposed along a distal portion of the elongate shaft and configured to contact target tissue; a guidewire shaft disposed within the lumen and extending along the elongate shaft: and at least two electrodes movably disposed within the lumen and configured for electrical coupling to the region of conductive surfaces, wherein each electrode is independently movable along the guidewire shaft to interface with a selected portion of the region of conductive surfaces to delivery energy therethrough. 21. The ablation catheter of claim 20 wherein each electrode of the at least two electrodes comprises: an electrode lumen sized for passing the guidewire shaft therethrough, and an electrode wire extending proximally therefrom and configured for coupling to a pulse generator source. 22. The ablation catheter of claim 21, further comprising a robotic controller operably coupled with the electrode wires, the robotic controller configured to move and to energize electrodes of the at least two electrodes. 23. The ablation catheter of claim 21 wherein the electrode wire of is further configured to move its corresponding electrode of the at least two electrodes along the guidewire shaft. 24. The ablation catheter of claim 20 wherein movement of at least two electrodes is constrained to prevent direct tissue contact therewith. 25. The ablation catheter of claim 20 wherein the region of conductive surfaces is free of any fixed electrical connections. 26. The ablation catheter of claim 25 wherein the region of conductive surfaces comprises a plurality of contact pads. 27. The ablation catheter of claim 26 wherein each contact pad of the plurality of contact pads is spaced at a regular interval along the distal portion of the elongate shaft. 28. The ablation catheter of claim 20 wherein the region of conductive surfaces comprises porosities that penetrate from an interior to an exterior of the wall. 29. The ablation catheter of claim 28 wherein the porosities are configured to pass a conductive fluid from the interior to the exterior of the wall. 30. The ablation catheter of claim 29 wherein the conductive fluid has a higher conductivity than blood or the target tissue. 31. The ablation catheter of claim 20 wherein the distal portion of the elongate shaft is configurable between a generally straightened configuration and a curved configuration. 32. The ablation catheter of claim 20, further comprising an expansible anchor structure coupled to a distal portion of the guidewire shaft, the anchor structure having a collapsed delivery configuration and an expanded anchoring configuration for anchoring to tissue. 33. The ablation catheter of claim 32 wherein the expansible anchor structure comprises an anchor electrode. 34. A catheter system for delivering ablative energy to target tissue of a patient, the catheter comprising: an elongate guide component; a shaft positioned around and extending along the elongate guide component; a connector positioned around the elongate guide component and spaced apart from the shaft; a plurality of electrode strips coupled between the shaft and the connector, wherein the electrode strips are arranged radially around the elongate guide component, and wherein the electrode strips are deformable, wherein the catheter system is movable between — a first configuration in which the plurality of electrode strips extend linearly along the elongate guide component and are radially compressed, while the shaft and the connector are spaced apart by a first distance, and a second configuration in which the plurality electrode strips are axially compressed and radially expanded, while the shaft and the connector are spaced apart by a second distance that is smaller than the first distance. 35. The catheter system of claim 34 wherein the catheter system further comprises an expansible structure that carries the plurality’ of electrode strips. 36. The catheter system of claim 35 wherein the plurality of electrode strips comprise an electrode array, and further wherein expansion of the expansible structure further causes expansion of at least one electrode in the electrode array. | 2023 | SHIFAMED HOLDINGS | Medical and Healthcare | Small-size company | http://shifamed.com/ | USA | The patent focuses on a pulsed-field ablation system delivered via a low-profile catheter formed for flexible, adaptive ablation under robotic control. It addresses the limitations of bulky, manually controlled PFA tools that lack the precision needed for selected tissue regions. Robotic control plus a low-profile catheter enable targeted, precise PFA delivery. | הפטנט מתמקד במערכת אבלציה בשדה פולסי המסופקת באמצעות קטטר בעל פרופיל נמוך שנוצר לאבלציה גמישה ומסתגלת תחת בקרה רובוטית. הוא מתמודד עם המגבלות של כלי PFA מסורבלים בבקרה ידנית החסרים את הדיוק הנדרש לאזורי רקמה נבחרים. בקרה רובוטית בתוספת קטטר בעל פרופיל נמוך מאפשרים מסירת PFA ממוקדת ומדויקת. | 1 |  |  |  | 1 | 1 | [5 empty] | 1 | [12 empty] | 1 | [7 empty] | SHIFAMED HOLDINGS LLC | SHIFAMED HOLDINGS LLC | SHIFAMED HOLDINGS LLC | SHIFAMED HOLDINGS LLC | SHIFAMED HOLDINGS LLC | 2024-04-15 | 2024 | 2024-10-17 | 2024 | 2023-04-14 | 2023 | Application | Utility Invention | inactive | WO2024US24678 | 315636490 | WO2024216286A3 | WO
462 | WO2025257797A1 | MICROWAVE CATHETER SYSTEM HAVING SHORTING DIODE FOR ABLATION AND TEMPERATURE MEASUREMENTS | Systems and methods for measuring cable pollution to accurately measure parameters during ablation such as temperature of the target tissue for safe and efficacious ablation of target tissue are provided. For example, the systems described herein comprise a switch, e.g., a Dicke switch, a radiometer, and a reference termination positioned at a proximal side of a coaxial cable connecting a generator to a catheter comprising a microwave antenna, and a diode circuit positioned at a distal side of the coaxial cable, e.g., within a handle of the catheter. The diode circuit may be biased to a closed state to create a short circuit at the distal end of the cable to thereby isolate the contribution of the cable to the radiometric signals traveling through the cable, e.g., cable pollution, which may be used to compute the target tissue temperature. | 1. WHAT IS CLAIMED: 1. A system for measuring tissue temperature, the system comprising: a catheter having a proximal region and a distal region, the catheter configured to be operatively coupled to a generator via a cable; an antenna disposed at the distal region of the catheter, the antenna configured to measure a radiometer temperature of a target tissue; a diode circuit disposed at the proximal region of the catheter, the diode circuit comprising a diode configured to transition between an open state where a transmission path exists between a radiometer and the antenna, and a closed state where a short circuit is created at a distal end of the cable to block the transmission path between the radiometer and the antenna; a reference termination disposed proximal to the cable, the reference termination comprising a temperature sensor configured to measure a reference temperature; a controller having instructions that, when executed, cause the controller to: transition the diode to the closed state and receive, via the radiometer disposed proximal to the cable, a first radiometric signal from the cable when the diode is in the closed state, the first radiometric signal indicative of cable temperature; transition the diode to the open state and receive, via the radiometer, a second radiometric signal from the cable when the diode is in the open state, the second radiometric signal indicative of the radiometer temperature and the cable temperature; calculate a first radiometric measurement based on a comparison of the first radiometric signal and the reference temperature; calculate a second radiometric measurement based on a comparison of the second radiometric signal and the reference temperature; and calculate a target tissue temperature based on a comparison of the first and second radiometric measurements. 2. The system of claim 1, wherein the antenna comprises a switching antenna configured to switch between emitting energy to ablate the target tissue and measuring the radiometer temperature generated as a result of the energy emission. 3. The system of claim 2, wherein the controller is configured to selectively switch between an ablation state where the switching antenna emits energy for a first time period and a measurement state where the radiometer receives radiometric signals from the cable and the reference temperature from the reference termination for a second time period. 4. The system of claim 3, wherein the first time period is at least 80% percent of a sum of the first and second time periods. 5. The system of claim 3, wherein, in the measurement state, the radiometer receives the first radiometric signal from the cable and the reference temperature from the reference termination in an alternating manner for a third time period, and receives the second radiometric signal from the cable and the reference temperature from the reference termination in an alternating manner for a fourth time period, and wherein the second time period is the sum of the third and fourth time periods. 6. The system of claim 5, wherein the third and fourth time periods are each 100 milliseconds. 7. The system of claim 1, wherein the controller is configured to selectively transition a switch operatively coupled to the radiometer between a first state where the radiometer receives radiometric signals from the cable, and a second state where the radiometer receives the reference temperature from the reference termination in an alternating manner. 8. The system of claim 1, wherein the diode, in the closed state, is configured to cause radiometric signals traveling distally through the cable towards the antenna to be reflected proximally through the cable towards the radiometer. 9. The system of claim 1, wherein the diode, in the closed state, is configured to connect an inner conductor of the cable to an outer conductor of the cable. 10. The system of claim 1, wherein the controller is configured to transition the diode from the open state to the closed state by injecting a bias through the cable. 11. The system of claim 10, wherein the diode circuit comprises a capacitor in cascade with the diode, the capacitor disposed distal to the diode and configured to block the bias from entering the antenna. 12. The system of claim 1, wherein the temperature sensor comprises at least one of a thermistor configured to measure the reference temperature or an integrated circuit temperature sensor. 13. The system of claim 1, wherein the first radiometric signal comprises radiometric noise generated based on a temperature of the radiometer and thermal noise generated by the cable, and accounts for cable loss. 14. The system of claim 1, wherein the controller is configured to perform a cable calibration to account for cable loss of the cable, and wherein the controller is configured to calculate the target tissue temperature during an ablation procedure while accounting for the cable loss of the cable. 15. The system of claim 14, wherein the cable calibration is based on a calculation of the first and second radiometric measurements prior to the ablation procedure and a known tissue temperature. 16. The system of claim 1, wherein the catheter comprises a handle at the proximal region, and wherein the diode circuit is disposed in the handle. 17. The system of claim 16, wherein the catheter comprises a cooling channel extending through the handle, the cooling channel configured to receive a cooling fluid to cool the diode circuit. 18. The system of claim 16, further comprising the cable, wherein the cable is configured to couple the handle to the generator, and wherein the generator comprises the radiometer, the reference termination, and the controller. 19. The system of claim 1, further comprising a carrier configured to mount a substrate of the diode circuit, the carrier configured to electrically couple the diode circuit to the antenna and the cable. 20. A method for measuring tissue temperature, the method comprising: providing an antenna configured to be positioned in or adjacent to a target tissue, the antenna disposed at a distal region of a catheter and operatively coupled to a generator via a cable; transitioning a diode of a diode circuit disposed at a proximal region of the catheter to a closed state to create a short circuit at a distal end of the cable to block a transmission path between the antenna and a radiometer disposed proximal to the cable; receiving, via the radiometer, a first radiometric signal from the cable when the diode is in the closed state, the first radiometric signal indicative of cable temperature; transitioning the diode to an open state to unblock the transmission path between the radiometer and the antenna; receiving, via the radiometer, a second radiometric signal from the cable when the diode is in the open state, the second radiometric signal indicative of a radiometer temperature measured by the antenna and the cable temperature; calculating a first radiometric measurement based on a comparison of the first radiometric signal and a reference temperature measured by a temperature sensor of a reference termination disposed proximal to the cable; calculating a second radiometric measurement based on a comparison of the second radiometric signal and the reference temperature; and calculating a target tissue temperature based on a comparison of the first and second radiometric measurements. 21. The method of claim 20, wherein transitioning the diode from the open state to the closed state comprises injecting a bias through the cable. 22. The method of claim 20, further comprising selectively switching between causing the antenna to emit energy to ablate the target tissue and causing the antenna to measure the radiometer temperature generated as a result of the energy emission. 23. The method of claim 20, further comprising selectively switching between an ablation state where the antenna emits energy for a first time period and a measurement state where the radiometer receives radiometric signals from the cable and the reference temperature from the reference termination for a second time period. 24. The method of claim 23, wherein the first time period is at least 80% percent of a sum of the first and second time periods. 25. The method of claim 20, further comprising selectively transitioning a switch operatively coupled to the radiometer between a first state where the radiometer receives radiometric signals from the cable, and a second state where the radiometer receives the reference temperature from the reference termination in an alternating manner. 26. The method of claim 20, further comprising performing a cable calibration to account for cable loss of the cable, and wherein calculating the target tissue temperature during an ablation procedure comprises calculating the target tissue temperature while accounting for the cable loss of the cable. | 2024 | HEPTA MEDICAL | Medical and Healthcare | Start Up | https://www.heptamed.com/ | France | The patent focuses on a microwave catheter system with a Dicke switch, radiometer, reference termination, and a shorting diode for accurate temperature measurement during ablation. It addresses cable-pollution effects in radiometric thermometry that bias temperature readings during microwave ablation. The shorting-diode and switching architecture enable accurate temperatures for safer, more efficacious ablation. | הפטנט מתמקד במערכת קטטר מיקרוגלים עם מתג Dicke, רדיומטר, סיום ייחוס ודיודת קצר למדידת טמפרטורה מדויקת במהלך האבלציה. הוא מתמודד עם אפקטי זיהום כבל בתרמומטריה רדיומטרית המטים קריאות טמפרטורה במהלך אבלציית מיקרוגלים. ארכיטקטורת דיודת הקצר וההחלפה מאפשרת טמפרטורות מדויקות לאבלציה בטוחה ויעילה יותר. |  |  |  |  | 1 | [7 empty] | 1 | [6 empty] | 1 | [10 empty] | 1 |  | HEPTA MEDICAL SAS | HEPTA MEDICAL SAS | HEPTA MEDICAL SAS | HEPTA MEDICAL SAS | HEPTA MEDICAL SAS | 2025-06-12 | 2025 | 2025-12-18 | 2025 | 2024-06-13 | 2024 | Application | Utility Invention | active | WO2025IB56059 | 325146682 | [NA] | WO
463 | WO2024146781A1 | CROSS-PROBE ABLATION SYSTEM AND METHODS OF USING THE SAME | A system according to at least one embodiment of the present disclose includes a first probe including: a first electrode positioned on a distal end of the first probe; and a second electrode different from the first electrode and positioned proximal to the first electrode; a second probe including: a third electrode positioned on a distal end of the second probe; a fourth electrode different from the third electrode and positioned proximal to the third electrode; and one or more circuits that enable current to flow from the first electrode to the third electrode or vice versa and/or from the second electrode to the fourth electrode or vice versa. | 1. CLAIMS1. A system, comprising: a first probe comprising: a first electrode disposed on a distal end of the first probe; and a second electrode different from the first electrode and disposed proximal to the first electrode; a second probe comprising: a third electrode disposed on a distal end of the second probe; a fourth electrode different from the third electrode and disposed proximal to the third electrode; and one or more circuits that enable current to flow from the first electrode to the third electrode and/or from the second electrode to the fourth electrode. 2. The system of claim 1, further comprising: a coolant apparatus that pumps coolant into one or both of the first probe and the second probe. 3. The system of claim 2, wherein the coolant flows out of the first probe and into the second probe. 4. The system of claim 3, wherein the coolant flows out of the second probe and back to the coolant apparatus. 5. The system of any of claims 1-4, further comprising: a generator that provides the current to the one or more circuits for delivery to electrodes of the first probe and the second probe, wherein optionally the first probe further comprises: a thermocouple disposed distal to the first electrode. 6. The system of any of claims 1-5, further comprising: a display that renders information about at least one of the first probe and the second probe, wherein optionally the information comprises at least one of information about a temperature of the first electrode, information about a temperature of the second electrode, information about a power supplied to the first electrode, information about a power supplied to the second electrode, information about a measured impedance between the first electrode and the third electrode, and information about a measured impedance between the second electrode and the fourth electrode. 7. A system, comprising: a first probe comprising: a first electrode disposed on a distal end of the first probe; and a second electrode different from the first electrode and disposed proximal to the first electrode; and one or more circuits that enable current to flow from the first electrode to a third electrode of a second probe and/or from the second electrode to a fourth electrode of the second probe. 8. The system of claim 7, wherein the third electrode is disposed distal to the fourth electrode on the second probe. 9. The system of claim 8, wherein the one or more circuits are cyclable between a first mode and a second mode, wherein current flows between the first electrode and the second electrode and between the third electrode and the fourth electrode in the first mode, and wherein current flows between the first electrode and the third electrode and between the second electrode and the fourth electrode in the second mode. 10. The system of claim 9, further comprising: a display that displays information about at least one of the first probe and the second probe, wherein optionally the information comprises at least one of information about a temperature of the first electrode, information about a temperature of the second electrode, information about a power supplied to the first electrode, information about a power supplied to the second electrode, information about a measured impedance between the first electrode and the third electrode, and information about a measured impedance between the second electrode and the fourth electrode. 11. A system, comprising: a generator; a first probe comprising: a first electrode; and a second electrode different from the first electrode and disposed proximal to the first electrode; a second probe comprising: a third electrode; and a fourth electrode different from the third electrode and disposed proximal to the third electrode; and one or more circuits disposable between the generator and the first and second probes, the one or more circuits enabling current to flow from the first electrode to the third electrode and/or from the second electrode to the fourth electrode. 12. The system of claim 11, wherein the first probe comprises: a thermocouple that generates a temperature measurement of tissue adjacent to the first electrode. 13. The system of claim 12, further comprising: a processor; and a memory storing data thereon that, when executed by the processor, cause the processor to: receive the temperature measurement; and render, to a display, a visual depiction of the temperature measurement. 14. The system of claim 13, wherein the data further cause the processor to: modulate, based on the temperature measurement, a magnitude of the current generated by the generator that flows from the first electrode to the third electrode; and/or wherein the data further cause the processor to: decrease, when the temperature measurement is above a threshold value, a magnitude of the current generated by the generator that flows from the second electrode to the fourth electrode. 15. The system of any of claims 13-14, wherein the data further cause the processor to: render, to the display, information about at least one the first probe and the second probe, wherein optionally the information comprises at least one of information about the temperature measurement, information about a power supplied to the first electrode, information about a power supplied to the second electrode, information about a measured impedance between the first electrode and the third electrode, and information about a measured impedance between the second electrode and the fourth electrode. | 2023 | Medtronic | Medical and Healthcare | Large-Size Company | https://www.medtronic.com | USA | The patent focuses on a cross-probe ablation system using two probes, each with two electrodes (one distal, one proximal), to deliver energy across the probes. It addresses the limitation of single-probe ablation, where the geometry of the energy field is fixed by one device. Cross-probe configurations create flexible, customizable lesion shapes for varied targets. | הפטנט מתמקד במערכת אבלציה צולבת-גשושים המשתמשת בשני גששים, כל אחד עם שתי אלקטרודות (אחת דיסטלית, אחת פרוקסימלית), כדי לספק אנרגיה בין הגששים. הוא מתמודד עם המגבלה של אבלציה בגשוש יחיד, שבה הגיאומטריה של שדה האנרגיה קבועה על ידי מכשיר אחד. תצורות צולבות-גשושים יוצרות צורות נגע גמישות והניתנות להתאמה למטרות מגוונות. | 1 |  |  |  | 1 | [5 empty] | 1 | [10 empty] | 1 | [8 empty] | 1 |  | MEDTRONIC HOLDING CO SARL | Medtronic | MEDTRONIC HOLDING CO SARL | Medtronic | Medtronic | 2023-12-19 | 2023 | 2024-07-11 | 2024 | 2023-01-05 | 2023 | Application | Utility Invention | inactive | WO2023EP86656 | 313742621 | EP4646160A1 | WO
464 | US20230076642A1 | INTERACTIVE MULTI-PROBE ABLATION GUIDANCE SYSTEM AND METHOD | The present invention is directed to a multi-probe ablation simulation and guidance system and method for use in tissue ablation procedures. In use, the relative locations of a plurality of ablation probes capable of providing ablation energy are determined, and the effect of energy provided by the probes based on the determined locations is predicted to identify a simulated ablation volume. This simulated ablation volume is compared with a target tissue volume. The relative locations of the probes can be adjusted based on the comparison between the simulated ablation volume and the target tissue volume, and the predicted effect rerun until the simulated ablation volume encompasses the target tissue volume to be ablated and necrotized. | 1. A tissue ablation system comprising: a multi-active probe controller; a plurality of ablation probes controlled by the controller and connected to a power source; an imaging device; a screen for displaying computer generated information and/or medical images; a computing system for executing programming code and algorithms; and a facility PACS network, with which the computing system interfaces. 2. The ablation system of claim 1, wherein the imaging device comprises a CT, MRI, PET, or ultrasound scanner device. 3. The ablation system of claim 1, wherein the screen is stand-alone component, or comprises a part of at least one of a Personal Computer (PC), the ablation controller, or the imaging device. 4. The ablation system of claim 1, wherein the computing system comprises a PC, an embedded system, a virtual machine, or a docker. 5. The ablation system of claim 1, wherein the computing system is directly or virtually connected to the screen. 6. The ablation system of claim 1, wherein the computing system interfaces with the controller, or to at least one ablation probe. 7. The ablation system of claim 1, wherein the computing system interfaces locally or remotely to the facility PACS network. 8. The ablation system of claim 1, wherein the computing system interfaces locally or remotely to the imaging device. 9. The ablation system of claim 1, further comprising a surgical tool tracking sub-system for tracking the intracorporeal position of the ablation probes. 10. The ablation system of claim 1, wherein the visual display is a 2D or 3D display. 11. A method of ablating a target tissue volume in a subject, comprising: identifying the position of at least two multi-active probes in a tissue volume of a patient; simulating at operational speed for the at least two multi-active probes the ablation process for said multi-active probes as positioned in the tissue volume with a computing device to predict the ablation tissue volume that would be necrotized if the ablation were performed; obtaining a visual display of the ablation tissue volume, said ablation tissue volume encompassing the tissue that would be necrotized if the ablation were performed with the at least two multi-active probes positioned in the tissue volume; adjusting the position of and/or energy provided by the at least two multi-active probes if the displayed ablation tissue volume does not encompass the target tissue to be ablated; repeating the simulating and obtaining a visual display steps until the ablation tissue volume encompasses the targeted tissue volume, where the simulation and display steps are conducted at an operational speed allowing for an interactive update of the displayed simulation relating to the adjustment of the probe positioning and/or energy provided; and performing the ablation. 11. The method of claim 11, wherein the ablation system is a radiofrequency ablation system, a microwave ablation system, a cryoablation system, or an irreversible electroporation ablation system. 12. The method of claim 11, wherein the positions of the at least two multi-active probes are identified with a surgical tool tracking system. 13. The method of claim 11, wherein the positions of the at least two multi-active probes are identified by acquiring an image of the tissue volume of the patient comprising the probes, and retrieving and processing the image with the computing device. 14. The method of claim 11, wherein the visual display is displayed by a screen or monitor. 15. The method of claim 11, further comprising: obtaining an image of the target tissue volume; and superimposing the image on the visual display of the ablation tissue volume. 16. The method of claim 11, further comprising conducting additional simulating and displaying steps simultaneously with the ablation step. 17. The method of claim 11, wherein the target tissue volume is a solid cancer or tumor. 18. The method of claim 11, wherein the target tissue volume is the atrium of the heart. 19. The method of claim 11, wherein the simulating step uses electrical and thermal characteristics of the tissue volume of the patient, the measured probe positions, the probe geometries and the electrical and thermal characteristics of the probes to predict the ablation tissue volume. 20. A method for predicting the ablation volume of an ablation procedure for ablating a target tissue volume in a subject, the method comprising: determining relative locations of a plurality of ablation probes capable of providing ablation energy; predicting the effect of energy collectively provided by the probes based on the determined locations to identify a simulated ablation tissue volume; comparing the simulated ablation tissue volume with the target tissue volume; adjusting the relative locations of the plurality of probes based on the comparison between the simulated ablation tissue volume and the target tissue volume, where the predicting step is conducted at an operational speed allowing for an interactive update of the predicted effect of energy collectively provided by the probes relative to the adjustments of the probe locations; and predicting an associated simulated ablation tissue volume relative to the adjusted locations until the simulated ablation tissue volume encompasses the target tissue volume. | 2020 | NE Scientific | Medical and Healthcare | Small-size company | http://www.newengland-scientific.com/ | USA | The patent focuses on a multi-probe ablation simulation and guidance system that determines the relative locations of multiple ablation probes and predicts the simulated ablation volume from their effects. It addresses the difficulty operators have predicting how multiple probes will combine to form a lesion. Simulation guides probe placement for the right combined ablation volume. | הפטנט מתמקד במערכת סימולציה והנחיה לאבלציה רב-גשושית הקובעת את המיקומים היחסיים של מספר גששי אבלציה ומנבאת את נפח האבלציה המדומה מהאפקטים שלהם. הוא מתמודד עם הקושי שיש למפעילים בניבוי כיצד גששים מרובים יתחברו ליצירת נגע. סימולציה מנחה את מיקום הגשוש לנפח האבלציה המשולב הנכון. | [7 empty] | 1 | 1 | [12 empty] | 1 | [6 empty] | 1 |  |  |  | NE SCIENT LLC | NE SCIENT LLC | NE SCIENTIFIC, LLC | NE SCIENTIFIC LLC | NE SCIENTIFIC LLC | 2022-10-20 | 2022 | 2023-03-09 | 2023 | 2020-04-22 | 2020 | Application | Utility Invention | active | US202217970097 | 287465152 | WO2021216750A1 | US
465 | WO2025146191A1 | FOCUSED ULTRASOUND ABLATION CATHETER AND DEVICE | Disclosed in the present invention are a focused ultrasound ablation catheter and a device. The focused ultrasound ablation catheter comprises a multi-chamber tube as a catheter body; a first ultrasound transducer, arranged at a distal end of the catheter body and electrically connected to a control part, the first ultrasound transducer being parallel to a central axis of the catheter body to be used for measuring a distance and a temperature of an inner wall of a blood vessel, thereby calculating an ablation target region and measuring a temperature of the target region for calculation of an output power of second ultrasound transducers and determination of an ablation end point; and at least two second ultrasound transducers, arranged at the distal end of the catheter body and electrically connected to the control part, the second ultrasound transducers and the central axis of the catheter body forming a set included angle, and the at least two second ultrasound transducers being symmetrical about the first ultrasound transducer for determining an ablation focus and performing focused ablation. By means of the catheter, energy transmitted by sound waves can be concentrated in a focused ablation region, and other regions are not ablated. | 1. An ultrasonic focused ablation catheter, characterized by comprising: The catheter body is a multi-lumen tube, comprising at least a guidewire lumen for inserting a guidewire and a guidewire lumen for inserting a guidewire; A first ultrasonic transducer is disposed at the distal end of the catheter body and is electrically connected to the control unit through a wire; wherein the first ultrasonic transducer is parallel to the central axis of the catheter body and is used to measure the distance to the inner wall of the blood vessel, thereby calculating the ablation target area; the first ultrasonic transducer is also used to measure the temperature of the target area, which is used for calculating the output power of the second ultrasonic transducer and determining the ablation endpoint; At least two second ultrasonic transducers are arranged at the distal end of the catheter body and are used to be electrically connected to the control unit through a wire; wherein the second ultrasonic transducer forms a set angle with the central axis of the catheter body, and at least two of the second ultrasonic transducers are symmetrical with respect to the first ultrasonic transducer, so as to determine the ablation focus and perform focused ablation. 2. The ultrasonic focused ablation catheter according to claim 1, characterized in that: The second ultrasonic transducer is rotatably arranged at the distal end of the catheter body to adjust the size of the set angle. 3. The ultrasonic focused ablation catheter according to claim 1, characterized in that: With the first ultrasonic transducer as the center, at least two of the second ultrasonic transducers can move relatively along the central axis of the catheter body to adjust the relative distance between the at least two ultrasonic transducers. 4. The ultrasonic focused ablation catheter according to claim 1, characterized in that: There are multiple first ultrasonic transducers, and the multiple first ultrasonic transducers are arranged in a circular ring around the circumferential direction of the catheter body; wherein two inclined second ultrasonic transducers are symmetrically distributed on both sides of each first ultrasonic transducer. 5. The ultrasonic focused ablation catheter according to claim 1, characterized in that: The second ultrasonic transducer is electrically connected to the controller to receive the current temperature of the ablation target area measured by the first ultrasonic transducer, and calculate the current output power; The second ultrasonic transducer releases ultrasonic waves of corresponding intensity according to the current output power to perform focused ablation on the ablation target area. 6. The ultrasonic focusing ablation catheter according to claim 1, characterized in that it also includes: A tip balloon, disposed on the end surface of the distal end of the catheter body and connected to the internal cavity of the catheter body, for blocking a blood vessel in a filled state; Wherein, the internal cavity includes a liquid inlet cavity for liquid inlet and a liquid outlet cavity for liquid outlet. 7. The ultrasonic focused ablation catheter according to claim 6, characterized in that: A flexible electrode is installed on the outside of the tip balloon, and a pressure sensor is installed on the inside of the tip balloon. Both the flexible electrode and the pressure sensor are electrically connected to the control unit for detecting impedance and water pressure. 8. An ultrasonic focused ablation device, characterized by comprising: A control unit, used for automatic control of ultrasonic ablation; An operating handle connected to the control unit; The ultrasonic focusing ablation catheter according to any one of claims 1 to 7, wherein the proximal end of the ultrasonic focusing ablation catheter is mounted on the operating handle and is electrically connected to the control unit; Wherein, the control unit continuously adjusts the output power of the second ultrasonic transducer according to the measured temperature of the first ultrasonic transducer until the ablation endpoint is reached; the control unit is also used to determine whether the ablation endpoint is reached according to the judgment logic. 9. The ultrasonic focused ablation device according to claim 8, characterized in that the judgment logic specifically includes: When the first ultrasonic transducer measures that the current temperature of the ablation target area reaches a set value and stabilizes for a preset time, and when the first ultrasonic transducer measures that the distance to the inner wall of the blood vessel reaches a theoretical value, it is determined that the ablation endpoint has been reached. 10. The ultrasonic focused ablation device according to claim 8, characterized in that the control unit at least comprises: The main control module is used to process signals and data and control automated ablation; A signal generator, connected to the main control module, for receiving and sending ultrasonic signals; A power amplifier connected to the signal generator to amplify the ultrasonic signal; A gating module connected to the power amplifier to select the number of signal channels; A phase voltage and current module, connected to the main control module, for detecting parameters and outputting a matching ultrasonic frequency; The impedance matching module is connected to the main control module and is used to detect the matching impedance and output an ultrasonic signal of matching power. | 2024 | SHANGHAI GOLDEN LEAF MED TEC | Medical and Healthcare | NA | https://www.brattea.com/en/aboutUsEn.html | China | The patent focuses on a focused ultrasound ablation catheter with a multi-chamber tube and a first ultrasound transducer at the distal end parallel to the catheter axis. It addresses the limitation of conventional ablation catheters that don't deliver focused ultrasound efficiently for targeted treatment. The focused-ultrasound design enables precise targeted ablation through a catheter. | הפטנט מתמקד בקטטר אבלציה באולטרסאונד ממוקד עם צינור רב-חללי ומתמר אולטרסאונד ראשון בקצה הדיסטלי במקביל לציר הקטטר. הוא מתמודד עם המגבלה של קטטרי אבלציה קונבנציונליים שאינם מספקים אולטרסאונד ממוקד ביעילות לטיפול ממוקד. עיצוב האולטרסאונד הממוקד מאפשר אבלציה ממוקדת מדויקת דרך קטטר. | 1 |  | 1 |  | 1 | [10 empty] | 1 | [7 empty] | 1 | [6 empty] | 1 |  | SHANGHAI GOLDEN LEAF MED TEC CO LTD | SHANGHAI GOLDEN LEAF MED TEC CO LTD | SHANGHAI GOLDEN LEAF MED TEC CO LTD | SHANGHAI GOLDEN LEAF MED TEC CO LTD | SHANGHAI GOLDEN LEAF MED TEC CO LTD | 2025-01-05 | 2025 | 2025-07-10 | 2025 | 2024-01-05 | 2024 | Application | Utility Invention | active | WO2025CN70630 | 322382181 | CN117982815A | WO
466 | EP4125612A1 | BRONCHIAL DENERVATION USING INTEGRATED A-MODE SIGNAL | Apparatus and methods for deactivating bronchial nerves extending along the secondary bronchial branches of a mammalian subject to treat asthma and related conditions. An ultrasonic transducer (11) is inserted into the bronchus as, for example, by advancing the distal end of a catheter (10) bearing the transducer into the secondary bronchial section to be treated. The ultrasonic transducer emits circumferential ultrasound so as to heat tissues throughout circular impact volume (13) as, for example, at least about 1 cm3 encompassing the bronchus to a temperature sufficient to inactivate nerve conduction but insufficient to cause rapid ablation or necrosis of the tissues. The treatment can be performed without locating or focusing on individual bronchial nerves. The apparatus and methods utilized for lung tumor ablation. | 1. Claims of equivalent WO 2021201963 A1 CLAIMS: 1. A method for treating adverse respiratory symptoms in a mammalian subject, comprising: inserting an ultrasound transducer into a bronchial tree of the mammalian subject; and operating an actuator or control unit, electrically connected to the ultrasound transducer, to energize the ultrasound transducer to emit a short pulse at a sub-therapeutic level; receiving via the ultrasound transducer an ultrasound echo from organic tissues of the mammalian subject in response to the short pulse, the ultrasound echo being a volume-integrated A- mode signal; operating the actuator or control unit to process the volume integrated A-mode signal, which represents an accumulated intensity of the circumferential ultrasound echoes; analyzing the volume integrated A-mode signal to determine locations of bronchial cartilage rings; and activating the ultrasound transducer to transmit ultrasound therapeutic waveform energy between adjacent ones of the bronchial cartilage rings. 2. The method of claim 1, further comprising moving the ultrasound transducer to position same adjacent a space or gap between the adjacent ones of the bronchial cartilage rings prior to the activating of the ultrasound transducer to transmit the ultrasound therapeutic waveform energy. 3. The method of claim 1 wherein the ultrasound transducer includes multiple separately activatable transducer elements, the activating of the ultrasound transducer includes energizing the separately activatable transducer elements to direct the ultrasound therapeutic waveform energy between the adjacent ones of the bronchial cartilage rings 4. The method of claim 1 , further comprising analyzing the volume integrated A-mode signal to detect presence of air pockets or trapped air between balloon and bronchus , the activating of the ultrasound transducer to transmit ultrasound therapeutic waveform energy being carried out only when analyzing of the volume integrated A-mode signal reveals an absence of air, thereby ensuring complete circumferential coupling. 5. The method of claim 1, further comprising operating the actuator or control unit to process the electrical signal to measure time delay of the ultrasound echo waveform relative to a time of emission of the short pulse to determine diameter of a bronchial section to be treated, wherein the activating of the ultrasound transducer to transmit ultrasound therapeutic waveform energy includes activating the ultrasound transducer at an acoustic power level in accordance with the determined size of the bronchial section. 6. The method of claim 1 wherein the activating of the ultrasound transducer includes transmitting an amount of ultrasound therapeutic waveform energy into an impact volume circumferentially surrounding a bronchial section, in an amount therapeutically effective to inactivate conduction of all the bronchial nerves in the impact volume. 7. The method of claim 1 , wherein for an impact volume of approximately 1 cm3, the ultrasound therapeutic waveform energy is transmitted at an acoustic power level of approximately 10 to approximately 50 watts for approximately 10 to approximately 20 seconds to provide an absorbed dose of approximately 100 to approximately 1000 joules throughout the impact volume. 8. The method of claim 1 , wherein the ultrasound therapeutic waveform energy is transmitted so as to maintain the temperature of a bronchial wall section surrounding the ultrasound transducer below 65°C while heating bronchial nerves in a circumferential impact volume about the bronchial wall section to above 42°C. 9. The method of claim 1 , wherein the inserting of the ultrasound transducer into the bronchial tree is performed through a working channel of a bronchoscope under visual guidance, or through a steerable sheath, or with a steerable ultrasound catheter through an oral intubating device, or under optical imaging guidance with an optical fiber inserted through the central lumen of the steerable ultrasound treatment catheter, or without a sheath or bronchoscope, directly through an oral intubation device with a steerable ultrasound catheter with a distance scale marking for monitoring degree of insertion after conducting a CT, MRI procedure to ascertain distance along a bronchial tree to the bronchial section. 10. The method of claim 1 , further comprising: inserting an air filled balloon into the esophagus of the patient; monitoring the volume integrated A-mode signal; advancing the treatment catheter distally in extreme cases beyond the first bifurcation until an esophageal signal caused by the air filled balloon inside the esophagus is diminished, thereby enabling prevention of peri esophageal nerve damage; and circulating a cooling fluid through the balloon, displacing the air, and thereby reducing periesophageal nerve damage. 11. The method of claim 1 , wherein the ultrasound transducer is mounted to a distal end of a catheter, the inserting of the ultrasound transducer into the bronchial tree includes inserting the catheter so that the ultrasound transducer is placed at a desired operating position determined at least in part based on a bending radius of a distal catheter portion monitored via strain gages. 12. The method of claim 1 , whereas the desired, catheter position is determined by monitoring the diameters of trachea and bifurcated bronchi. 13. Apparatus for inactivating bronchial nerve conduction in a mammalian subject, comprising: an ultrasound transducer adapted for insertion into a bronchial tree of the mammalian subject and for emitting ultrasound energy; and an actuator or control unit electrically connected to the ultrasound transducer, the actuator or control unit being configured to activate the ultrasound transducer to emit a short pulse at a sub- therapeutic level, the ultrasound transducer being operatively connected to the actuator or control unit for receiving an electrically encoded volume-integrated A- node signal or accumulated intensity of ultrasound echoes from organic tissues of the mammalian subject, the actuator or control unit being further configured to process the volume integrated A-mode signal. 14. The apparatus of claim 13 wherein the actuator or control unit is additionally configured to analyze the volume integrated A-mode signal to distinguish a relative minimum of the volume integrated A-mode signal with respect to degree of insertion of the ultrasound transducer in the bronchial tree, thereby enabling positioning of the ultrasound transducer to transmit therapeutically effective focused ultrasound energy into a treatment or impact volume between bronchial cartilage rings. 15. The apparatus of claim 14 wherein the actuator or control unit is configured to measure time delay of the volume integrated A-mode signal and therewith determine size of a bronchial section, the actuator or control unit being also configured to control the ultrasound transducer to vary the amount of the therapeutically effective focused ultrasound energy in accordance with the determined size of the bronchial section. 16. The apparatus of claim 13 wherein the ultrasound transducer includes a longitudinal array of separately activatable transducer elements and the actuator or control unit is configured to energize the separately activatable transducer elements in a sequence to direct the ultrasound therapeutic waveform energy between the adjacent ones of the bronchial cartilage rings. 17. The apparatus of claim 13 wherein the actuator or control unit is configured to analyze the volume integrated A-mode signal to facilitate complete circumferential coupling. 18. The apparatus of claim 13 wherein the ultrasound transducer is provided on a steerable ultrasound catheter with a distance scale marking for monitoring degree of insertion after conducting a CT, MRI procedure to ascertain distance along a bronchial tree to a bronchial section. 19. The apparatus of claim 13 wherein the ultrasound transducer is provided on a catheter with strain gages. 20. Apparatus for inactivating bronchial nerve conduction in a mammalian subject comprising: an ultrasound transducer adapted for insertion into a bronchial section of the mammalian subject and for emitting focused ultrasound energy; and an actuator or control unit electrically connected to the ultrasound transducer, the actuator or control unit being configured to control the ultrasound transducer to transmit focused ultrasound energy into an impact volume of at least approximately 1 cm3 at least partially surrounding the bronchial section, the focused ultrasound energy being in an amount effective to inactivate conduction of bronchial nerves throughout the impact volume. 21. The apparatus of claim 20, wherein the actuator or control unit is configured to control the ultrasound transducer to transmit focused ultrasound energy at an acoustic power level of approximately 10 to approximately 50 watts for approximately 10 to approximately 20 seconds to provide a dose of approximately 100 to approximately 1000 joules in the impact volume. 22. The apparatus of claim 20, wherein the actuator or control unit is configured to control the transducer so as to maintain temperature of a section of bronchial wall of the bronchial section below 65°C while achieving a temperature above 42°C throughout the impact volume. 23. The apparatus of claim 20, wherein the actuator is adapted to control the ultrasound transducer to transmit the focused ultrasound energy in a pulsed function to vary treatment times and temperatures to increase the treatment time while maintaining a predetermined dose, to homogenize temperature distribution throughout the impact volume. 24. The apparatus of claim 20, wherein the ultrasound transducer is a circular phased array and the actuator or control unit is adapted to generate quasi-simultaneous ultrasound imaging and treatment waveforms so that progress of an ablation procedure can be monitored in real time via 2D imaging. 25. The apparatus of claim 20, further comprising a catheter with a distal end and a proximal end, the ultrasound transducer being mounted to the catheter adjacent the distal end inside a compliant balloon, filled with circulating fluid, in order to cool the ultrasound transducer and to center and align the ultrasound transducer with the bronchial axis in order to facilitate homogenous or even ultrasound power distribution throughout the impact volume. 26. The apparatus of claim 25, wherein the catheter is constructed and arranged to hold the ultrasound transducer out of contact with a wall of the bronchial section to avoid wall damage. 27. The apparatus of claim 25, wherein the ultrasound transducer has an axis, the catheter is constructed and arranged to hold the axis of the ultrasound transducer generally parallel to an axis of the bronchial section, and the ultrasound transducer is adapted to transmit ultrasound energy in a 360° cylindrical pattern surrounding the axis of the ultrasound transducer to optimize power distribution throughout impact volume. 28. The apparatus of claim 25, wherein the compliant balloon includes a porous membrane engageable with the wall of the bronchial section with fluid for improved acoustic coupling while preserving longitudinal movability of the compliant balloon and the ultrasound transducer disposed therein: 29. The apparatus of claim 25, wherein the circulating fluid consists of a coupling and cooling fluid, further comprising means for circulating the coupling and cooling fluid, the actuator or control unit being configured for measuring fluid volume and pressure of the cooling and circulating fluid inside the compliant balloon, the actuator or control unit being further configured to determine a point of constant volume and rapid pressure increase and therewith determine a diameter of the bronchial section, thereby optimizing overall ultrasound power. 30. The apparatus of claim 20, further comprising a compliant balloon filled with a circulating fluid, said ultrasound transducer being disposed inside said compliant balloon, also comprising focusing componentry taken from a group consisting of electronic control, a fluid lens and a mechanical lens. 31. A method for inactivating bronchial nerve conduction in a mammalian subject comprising the steps of: inserting an ultrasound transducer into a bronchial section of the mammalian subject; and actuating the ultrasound transducer to transmit therapeutically effective focused ultrasound energy into an impact volume of at least approximately 1 cm3 at least partially surrounding the bronchial section so that the therapeutically effective focused ultrasound energy inactivates conduction of all the bronchial nerves in the impact volume. 32. The method of claim 31 , wherein the ultrasound energy is transmitted at an acoustic power level of approximately 10 to approximately 50 watts for approximately 10 to approximately 20 seconds to provide an absorbed dose of approximately 100 to approximately 1000 joules throughout the impact volume. 33. The method of claim 31 , wherein the actuating of the ultrasound transducer is performed so as to maintain temperature of the bronchial section below 65°C while heating bronchial nerves in the impact to above 42°C. 34. The method of claim 31 , wherein the activating of the ultrasound transducer to transmit ultrasound therapeutic waveform energy is performed without determining locations of any specific bronchial nerves. 35. The method of claim 31 wherein the activating of the transducer is performed so that a single application of ultrasound energy to a treatment section of each main bronchus is effective to inactivate conduction distal of the respective treatment section of all bronchial nerves surrounding the respective treatment section. 36. The method of claim 31 , further comprising the steps of: repositioning the ultrasound transducer in the bronchial tree after the step of actuating the transducer; and then repeating the step of actuating the ultrasound transducer. 37. The method of claim 31 , wherein the activating of the transducer includes emitting ultrasound energy in a pulsed function to increase the treatment time without increasing the dose, thereby evening temperature throughout the impact volume. 38 The method of claim 31 , wherein the ultrasound transducer is a phased array activatable to perform quasi-simultaneous 2D imaging and treatment, thereby enabling monitoring of an ablation process in real time.. 39. The method of claim 31 , wherein the inserting of the ultrasound catheter into the bronchial section is performed through a working channel of a bronchoscope under visual guidance. 40. The method of claim 31 , wherein the inserting of the ultrasound catheter is performed through a steerable sheath. 41. The method of claim 31 , wherein the inserting of the ultrasound catheter is performed without a sheath or bronchoscope with a steerable ultrasound catheter through an oral intubating device. 42. The method of claim 31 , wherein the inserting of the ultrasound catheter is performed under optical imaging guidance through an optical fiber inserted through the central lumen of the steerable ultrasound treatment catheter. 43. The method of claim 31, wherein the inserting of the ultrasound transducer is performed without a sheath or bronchoscope, directly through an oral intubation device with a steerable ultrasound catheter with a distance scale marking for monitoring degree of insertion, further comprising conducting a GT, RI procedure prior to the inserting of the ultrasound catheter to ascertain distance along a bronchial tree to the bronchial section. 44. The method of claim 43, whereas the desired catheter position is determined based on the bending radius of the distal catheter portion monitored via strain gages. 45. The method of claim 43, whereas the desired catheter position is determined by monitoring the diameters of trachea vs. bifurcated bronchi. 46. The method of claim 31 , wherein the inserting of the ultrasound transducer is performed over a guide wire which has been placed through a working channel of a bronchoscope. 47. The method of claim 31 , further comprising the steps of: applying non-therapeutic ultrasound energy at a power level less than a power level of the therapeutically effective ultrasonic energy; receiving reflected non-therapeutic ultrasonic energy and generating echo signals representing the reflected energy; and optimizing the longitudinal position of the treatment volume so it is laying between cartilage rings. 48. The method of claim 31 , further comprising determining a diameter of the bronchial section based on echo signals before the actuating of the ultrasound transducer to apply the therapeutically effective ultrasonic energy. 49. The method of claim 48, further comprising adjusting the power applied to the ultrasound transducer during activating thereof to emit therapeutically effective focused ultrasound energy in an amount based at least in part on the determined diameter of the bronchial section. 50. A probe for use in bronchial nerve inactivation, the probe comprising: an ultrasound transducer adapted for transmitting focused ultrasound energy; and a catheter with a distal end and a proximal end, the transducer being mounted to the catheter adjacent the distal end, the catheter and transducer being constructed and arranged to allow positioning of the distal end and transducer within a bronchial section and to allow a compliant balloon to contact the bronchial wall. 51 : The probe of claim 50, wherein the catheter is constructed and arranged to hold the transducer centered and aligned within the bronchial section to be treated. 52. The probe of claim 50, wherein the transducer has an axis, the catheter is constructed and arranged to hold the axis of the transducer generally parallel to the axis of the bronchus, and the transducer is adapted to transmit the ultrasound energy in a 360° cylindrical pattern surrounding the axis of the transducer. 53. The probe of claim 50, wherein the catheter includes a fluid filled centering element surrounding the transducer, the centering element being arranged to hold the transducer substantially centered in the bronchial section. 54. The probe of claim 50, wherein the centering element includes an expansible element adapted to engage the wall of the bronchial section. 55. The method of claim 31 performed with a side-firing transducer inserted into the pulmonary artery. | 2020 | AERWAVE MEDICAL | Medical and Healthcare | Start Up | http://www.aerwavemed.com/ | USA | The patent focuses on bronchial denervation using an ultrasonic transducer in a catheter inserted into secondary bronchial branches, with integrated A-mode signal feedback. It addresses the need for nerve-targeted asthma therapy that uses minimal but well-monitored ultrasound. The A-mode signal helps confirm transducer placement during bronchial nerve deactivation. | הפטנט מתמקד בדה-אינרבציה של הסימפונות באמצעות מתמר אולטרסוני בקטטר המוחדר להסתעפויות סימפונות משניות, עם היזון חוזר משולב של אות A-mode. הוא מתמודד עם הצורך בטיפול אסתמה ממוקד עצבים המשתמש באולטרסאונד מינימלי אך מנוטר היטב. אות ה-A-mode עוזר לאשר את מיקום המתמר במהלך השבתת עצבי הסימפונות. |  |  | 1 |  |  | 1 | [9 empty] | 1 | [10 empty] | 1 | [5 empty] | AERWAVE MEDICAL INC | AERWAVE MEDICAL INC | AERWAVE MEDICAL INC | AERWAVE MEDICAL INC | AERWAVE MEDICAL INC | 2021-01-29 | 2021 | 2023-02-08 | 2023 | 2020-03-31 | 2020 | Application | Utility Invention | active | EP20210779200 | 316231612 | AU2021246864A1¦CA3176126A1¦EP4125612A4¦JP2023520898A¦JP7728015B2¦KR20220162745A¦US11565135B2¦US11607568B2¦US12502559B2¦US12582847B2¦US12594438B2¦US20210316161A1¦US20220008753A1¦US20230057488A1¦US20230135085A1¦US20240123262A1¦US20240123263A1¦US20240366967A1¦WO2021201963A1 | EP
467 | EP4395863A1 | MAGNETICALLY STEERABLE IRRIGATED ABLATION CATHETERS, AND SYSTEMS AND METHODS THEREOF | Embodiments described herein relate to magnetically steerable irrigated ablation catheters and methods of operating the same. A steerable catheter can include a distal tip including a first set of magnets. The steerable catheter further includes a shaft. The shaft includes a flexible section with a distal end coupled to the distal tip. The shaft further includes a second set of magnets spaced along a length of the flexible section and spaced from the first set of magnets such that the first and second sets of magnets collectively enable the flexible section to curve without kinking in response to a magnetic field being applied to the first and second set of magnets. In some embodiments, each magnet from the second set of magnets can be coupled to the flexible section via metallic elements disposed at the ends of each magnet. | 1. A steerable catheter, comprising: a distal tip including a first set of magnets; and a shaft including: a flexible section having a distal end coupled to the distal tip; and a second set of magnets disposed on the flexible section, the second set of magnets being spaced along a length of the flexible section and being spaced from the first set of magnets such that the first and second sets of magnets collectively enable the flexible section to curve without kinking in response to a magnetic field being applied to the first and second sets of magnets. 2. The steerable catheter of claim 1, wherein the flexible section includes a set of coils each being disposed between adjacent magnets of the second set of magnets, each magnet from the second set of magnets is coupled to a coil from the set of coils via metallic elements disposed at the ends of each magnet. 3. The steerable catheter of claim 2, wherein each metallic element is formed of a ferromagnetic material. 4. The steerable catheter of claim 1, wherein the first set of magnets is separated into at least two sections with a flexible element in between, the flexible element configured to allow the distal tip to curve. 5. The steerable catheter of claim 1, wherein the distal tip includes a treatment electrode configured to apply a treatment to a tissue site. 6. The steerable catheter of claim 5, wherein the treatment electrode includes a tip electrode having a set of irrigation openings, the steerable catheter further including: an irrigation lumen extending from a proximal end of the shaft to the tip electrode, the irrigation lumen configured to convey a fluid to the set of irrigation openings such that the fluid can be delivered to the tissue site. 7. The steerable catheter of claim 6, wherein the tip electrode includes a helical distribution member that is configured to distribute the fluid substantially evenly to the set of irrigation openings. 8. The steerable catheter of claim 6, wherein the irrigation lumen has sufficient stretch resistance such that the shaft remains intact with application of tensile forces greater than or equal to about 15 N. 9. The steerable catheter of claim 5, wherein the treatment electrode is configured to apply at least one of: radiofrequency (RF) energy to ablate tissue at the tissue site, or pulsed field energy to irreversibly electroporate tissue at the tissue site. 10. The steerable catheter of claim 1, wherein the distal tip includes a measurement electrode configured to measure electrical signals for at least one of: diagnosis of biological data, localization of the distal tip, or mapping of a region of patient anatomy. 11. The steerable catheter of claim 1, further comprising: a set of electrodes disposed along the flexible section of the shaft. 12. The steerable catheter of claim 11, wherein each electrode from the set of electrodes is disposed adjacent to a magnet from the second set of magnets. 13. The steerable catheter of claim 1, wherein the distal tip includes a temperature sensor configured to sense a temperature of tissue near the distal tip. 14. The steerable catheter of claim 1, wherein the flexible section is configured to curve to fit within a heart chamber. 15. The steerable catheter of claim 1, wherein the second set of magnets is evenly distributed along the length of the flexible section of the shaft. 16. A steerable catheter, comprising: a distal tip including a treatment electrode configured to apply a treatment to a tissue site, the treatment electrode defining a set of irrigation openings; a shaft including: a flexible section having a distal end coupled to the distal tip; and a set of magnets disposed on the flexible section, the set of magnets being evenly distributed along a length of the flexible section such that the set of magnets collectively steer the distal tip in response to a magnetic field being applied to the set of magnets; and an irrigation lumen extending from a proximal end of the shaft to the treatment electrode, the irrigation lumen configured to convey a fluid to the set of irrigation openings such that the fluid can be delivered to the tissue site, the irrigation lumen having greater stretch resistance than the flexible section to prevent breakage of the flexible section with application of tensile forces. 17. The steerable catheter of claim 16, wherein the treatment electrode is configured to apply at least one of: radiofrequency (RF) energy to ablate tissue at the tissue site, or pulsed field energy to irreversibly electroporate tissue at the tissue site. 18. The steerable catheter of claim 16, wherein the treatment electrode is formed of a thermally conductive material, the distal tip further including a temperature sensor communicatively coupled to the treatment electrode such that the temperature sensor can measure a temperature of the tissue site. 19. The steerable catheter of claim 16, wherein the distal tip further includes a measurement electrode configured to measure electrical signals for at least one of: localization of the distal tip or mapping of a region of patient anatomy. 20. The steerable catheter of claim 19, wherein the distal tip further includes a pacing electrode configured to apply pacing to a heart. 21. The steerable catheter of claim 16, wherein the shaft further includes a set of electrodes disposed along the length of the flexible section, the set of electrodes configured to apply a treatment to the tissue site or measure an electrical signal. 22. The steerable catheter of claim 21, wherein each electrode from the set of electrodes is disposed adjacent to a magnet from the set of magnets. 23. The steerable catheter of claim 21, wherein each magnet from the set of magnets is coupled to the flexible section of the shaft via metallic elements disposed at the ends of each magnet, each electrode from the set of electrodes being coupled to one of the metallic elements. 24. The steerable catheter of claim 16, wherein the distal tip includes two rigid sections coupled via a flexible section. 25. A method, comprising: determining a current position and orientation of a distal tip of a catheter disposed within a body of a patient, the catheter including a set of at least five magnets being evenly distributed along a shaft of the catheter; determining a target position and orientation for navigating the distal tip of the catheter; applying, using a magnetic field generator disposed outside of the body of the patient, a magnetic field having a direction orientated for steering the distal tip of the catheter via the set of magnets from the current position and orientation to the target position and orientation; and causing, in response to the application of the magnetic field, one or more sections of the shaft including the set of magnets to curve in a simple or compound configuration based on the direction of the magnetic field and the placement of the set of magnets relative to the direction of the magnetic field. 26. The method of claim 25, wherein the current position and orientation of the distal tip of the catheter is determined based on electrical signals received from at least one electrode disposed on the distal tip or the shaft of the catheter. 27. The method of claim 25, wherein the at least one electrode belongs to a set of electrodes evenly distributed along the shaft of the catheter, each electrode from the set of electrodes being disposed adjacent to a magnet from the set of magnets. 28. The method of claim 25, further comprising at least one of: applying, using a signal generator, a voltage waveform to at least one electrode disposed on the distal tip of the catheter to generate a pulsed electric field for ablating a tissue site adjacent to the distal tip; applying, using a signal generator, radiofrequency energy to at least one electrode disposed on the distal tip of the catheter to ablate a tissue site adjacent to the distal tip; or delivering an ablation medium to a distal portion of the catheter to ablate a tissue site, the ablation medium including at least one of: a thermal agent, a cryogenic agent, or a chemical agent. 29. The method of claim 25, further comprising: delivering a fluid to a tissue site via an irrigation lumen extending through the shaft to the distal tip of the catheter, the distal tip including a tip electrode defining a set of irrigation openings through which the fluid can be delivered. 30. The method of claim 28, further comprising: measuring a temperature of the tissue site using a temperature sensor disposed within the distal tip of the tip electrode; determining whether the temperature is greater than a threshold value; and terminating the ablation of the tissue site in response to determining that the temperature is greater than the threshold value. | 2021 | Stereotaxis | Medical and Healthcare | Small-size company | http://www.stereotaxis.com/ | USA | The patent focuses on a magnetically steerable irrigated ablation catheter with sets of magnets along the shaft and at the distal tip. It addresses the difficulty of steering ablation catheters precisely through complex anatomy without manual manipulation. Magnetic steering plus irrigation enables flexible navigation while preventing tissue overheating. | הפטנט מתמקד בקטטר אבלציה מוצף הניתן לניווט מגנטי עם קבוצות מגנטים לאורך הפיר ובקצה הדיסטלי. הוא מתמודד עם הקושי בניווט קטטרי אבלציה במדויק דרך אנטומיה מורכבת ללא מניפולציה ידנית. ניווט מגנטי בתוספת השקיה מאפשר ניווט גמיש תוך מניעת התחממות יתר של רקמה. | 1 |  |  |  | 1 | [5 empty] | 1 | [13 empty] | 1 | [7 empty] | STEREOTAXIS INC | Stereotaxis | STEREOTAXIS INC | Stereotaxis | Stereotaxis | 2022-08-30 | 2022 | 2024-07-10 | 2024 | 2021-08-30 | 2021 | Application | Utility Invention | active | EP20220777442 | 287400774 | US20230068563A1¦WO2023034319A1 | EP
468 | EP4469152A2 | HISTOTRIPSY SYSTEMS AND METHODS | A histotripsy therapy system configured for the treatment of tissue is provided, which may include any number of features. Provided herein are systems and methods that provide efficacious non-invasive and minimally invasive therapeutic, diagnostic and research procedures. In particular, provided herein are systems and methods for acoustically coupling a histotripsy therapy system to the skin of a patient to provide targeted, efficacious histotripsy in a variety of different regions and under a variety of different conditions without causing undesired tissue damage to intervening/non-target tissues or structures. | 1. Claims of equivalent WO 2023141653 A2 CLAIMSWhat is claimed is: 1. An ultrasound treatment head, comprising: a therapy transducer array configured to deliver ultrasound pulses to a focal location; a bore located within the therapy transducer array; a coupling assembly sized and configured for axial and rotational movement within the bore of the therapy transducer array; an ultrasound imaging probe coupled to the coupling assembly; and at least one user input device configured to control axial movement and/or rotation of the ultrasound imaging probe relative to the therapy transducer array. 2. The ultrasound treatment head of claim 1, further comprising at least one encoder configured to track an axial and/or rotational position of the ultrasound imaging probe. 3. The ultrasound treatment head of claim 2, further comprising one or more processors configured to register the axial and/or rotational position of the ultrasound imaging probe with a digital treatment plan. 4. The ultrasound treatment head of claim 2, wherein the at least one encoder is in electrical communication with one or more remote processors configured to register the axial and/or rotational position of the ultrasound imaging probe with a digital treatment plan. 5. The ultrasound treatment head of claim 4, wherein the digital treatment plan further comprises at least one of a target contour and a margin contour. 6. The ultrasound treatment head of claim 1, wherein the at least one user input device further comprises: an image probe control; a first shaft coupled to the image probe control; a first gear coupled to the first shaft; a second gear coupled to the first gear; a second shaft coupled to the second gear; a shuttle assembly coupled to the coupling assembly and configured to move axially along the second shaft when the second shaft is rotated; wherein manipulation of the image probe control results in a corresponding axial movement of the ultrasound imaging probe along a length of the second shaft. 7. The ultrasound treatment head of claim 1, wherein the at least on user input device further comprises: a rotational sleeve disposed around the coupling assembly, the rotational sleeve being keyed to the coupling assembly, wherein rotation of the rotational sleeve results in a corresponding rotational movement of the ultrasound imaging probe. 8. The ultrasound treatment head of claim 7, further comprising one or more detents positioned at specified rotational angles against the rotational sleeve to provide tactile feedback to a user. 9. The ultrasound treatment head of claim 7, wherein the rotational sleeve further includes one or more tabs which align with an axial position of the ultrasound imaging probe. 10. The ultrasound treatment head of claim 1, further comprising one or more bushings configured to align an outer diameter of the coupling assembly with an inner diameter of the bore. 11. The ultrasound treatment head of claim 1, further comprising a position sensor configured to determine when the ultrasound imaging probe is in a retracted position. 12. The ultrasound treatment head of claim 11, further comprising one or more processors in electrical communication with the position sensor configured to prevent initiation of therapy with the therapy transducer array if the position sensor determines that the ultrasound imaging probe is not in a specified position with respect to the therapy transducer array. 13. The ultrasound treatment head of claim 12, wherein the specified position comprises fully retracted within the ultrasound treatment head. 14. The ultrasound treatment head of claim 12, wherein the specified position comprises positioned so as to not obstruct or substantially obstruct the therapy transducer array. 15. The ultrasound treatment head of claim 1, further comprising a keyed plate configured to rotate with rotation of the ultrasound imaging probe. 16. The ultrasound treatment head of claim 15, further comprising a rotational encoder coupled to the keyed plate, wherein a rotor of the rotational encoder rotates with rotation of the keyed plate and a stator of the rotational encoder remains static during rotation of the keyed plate. 17. The ultrasound treatment head of claim 16, further comprising a grip base disposed over the keyed plate and configured to couple to the therapy transducer array. 18. The ultrasound treatment head of claim 1, further comprising a quick-release connector configured to couple a robotic positioning arm to the ultrasound treatment head. 19. The ultrasound treatment head of claim 18, wherein the quick-release connector includes mating features configured to interface with corresponding mating features on the robotic positioning arm. 20. The ultrasound treatment head of claim 16, further comprising one or more handles, wherein at least one handle is hollow to allow for routing of a wire from the rotational encoder to a processor. 21. The ultrasound treatment head of claim 1, further comprising a pair of handles positioned on opposing sides of the ultrasound treatment head. 22. The ultrasound treatment head of claim 1, further comprising at least one handle, wherein the at least one handle includes a free-drive user input device configured to allow a user to move or manipulate a position and/or orientation of the ultrasound treatment head when the ultrasound treatment head is coupled to a robotic positioning arm. 23. An ultrasound therapy system, comprising: a treatment head that includes an ultrasound therapy transducer and an ultrasound imaging probe, wherein the ultrasound imaging probe is configured to move axially and be rotated independently with respect to the ultrasound therapy transducer; a robotic positioning arm coupled to the treatment head, the robotic positioning arm being configured to move the treatment head according to a digital treatment plan. 24. The ultrasound therapy system of claim 23, further comprising at least one encoder disposed in the treatment head and configured to track an axial and/or rotational position of the ultrasound imaging probe. 25. The ultrasound therapy system of claim 23, further comprising one or more processors configured to register a position of the ultrasound imaging probe with the digital treatment plan. 26. The ultrasound therapy system of claim 25, further comprising an electrical connection between the therapy head and the robotic positioning arm to communicate the position of the ultrasound imaging probe to a controller of the robotic positioning arm. 27. The ultrasound therapy system of claim 23, further comprising a quick-release connector configured to couple the robotic positioning arm to the treatment head. 28. The ultrasound therapy system of claim 1, further comprising at least one encoder configured to track an axial and/or rotational position of the ultrasound imaging probe. 29. The ultrasound therapy system of claim 28, further comprising one or more processors configured to register the axial and/or rotational position of the ultrasound imaging probe with a digital treatment plan. 30. The ultrasound therapy system of claim 28, wherein the at least one encoder is in electrical communication with one or more remote processors configured to register the axial and/or rotational position of the ultrasound imaging probe with a digital treatment plan. 31. The ultrasound therapy system of claim 23, further comprising at least one user input device configured to control axial and/or rotational movement of the ultrasound imaging probe. 32. The ultrasound therapy system of claim 31, wherein the at least one user input device further comprises: an imaging probe control; a first shaft coupled to the imaging probe control; a first gear coupled to the first shaft; a second gear coupled to the first gear; a second shaft coupled to the second gear; a shuttle assembly coupled to the coupling assembly and configured to move axially along the second shaft when the second shaft is rotated; wherein manipulation of the imaging probe control results in a corresponding axial movement of the ultrasound imaging probe along a length of the second shaft. 33. The ultrasound therapy system of claim 31, wherein the at least on user input device further comprises: a rotational sleeve disposed around the coupling assembly, the rotational sleeve being keyed to the coupling assembly, wherein rotation of the rotational sleeve results in a corresponding rotational movement of the ultrasound imaging probe. 34. The ultrasound therapy system of claim 33, further comprising one or more detents positioned at specified rotational angles against the rotational sleeve to provide tactile feedback to a user. 35. The ultrasound therapy system of claim 33, wherein the rotational sleeve further includes one or more tabs which align with an axial position of the ultrasound imaging probe. 36. The ultrasound therapy system of claim 23, further comprising one or more bushings configured to align an outer diameter of the coupling assembly with an inner diameter of the bore. 37. The ultrasound therapy system of claim 23, further comprising a position sensor configured to determine when the ultrasound imaging probe is in a retracted position. 38. The ultrasound therapy system of claim 37, further comprising one or more processors in electrical communication with the position sensor configured to prevent initiation of therapy with the therapy transducer array if the position sensor determines that the ultrasound imaging probe is not in a specified position with respect to the therapy transducer array. 39. The ultrasound therapy system of claim 38, wherein the specified position comprises fully retracted within the ultrasound treatment head. 40. The ultrasound therapy system of claim 38, wherein the specified position comprises positioned so as to not obstruct or substantially obstruct the therapy transducer array. 41. The ultrasound therapy system of claim 23, further comprising a keyed plate configured to rotate with rotation of the ultrasound imaging probe. 42. The ultrasound therapy system of claim 41, further comprising a rotational encoder coupled to the keyed plate, wherein a rotor of the rotational encoder rotates with rotation of the keyed plate and a stator of the rotational encoder remains static during rotation of the keyed plate. 43. The ultrasound therapy system of claim 41, further comprising a grip base disposed over the keyed plate and configured to couple to the therapy transducer array. 44. A method, comprising: submerging an ultrasound treatment head into an acoustic coupling medium in acoustic communication with a patient; positioning the ultrasound treatment head with a robotic arm such that a focus of the ultrasound treatment head is located on a target tissue volume; axially advancing an ultrasound imaging probe of the ultrasound treatment head towards the target tissue volume while maintaining the focus of the ultrasound treatment head on the target tissue volume; obtaining ultrasound images of the target tissue volume with the ultrasound imaging probe for treatment planning; at least partially retracting the ultrasound imaging probe relative to the ultrasound treatment head; and transmitting histotripsy pulses from the ultrasound treatment head into the target tissue volume. 45. The method of claim 44, further comprising rotating the ultrasound imagining probe while maintaining the focus of the ultrasound treatment head on the target tissue volume. 46. The method of claim 44, wherein axially advancing the ultrasound imaging probe further comprises axially advancing the ultrasound imaging probe to make contact with a membrane of a coupling container in acoustic communication with the patient. 47. The method of claim 46, wherein the membrane is in contact with skin of the patient. 48. The method of claim 44, wherein axially advancing the ultrasound imaging probe further comprises axially advancing the ultrasound imaging probe to make contact with skin of the patient. 49. The method of claims 45 or 46, further comprising moving the ultrasound treatment head with the robotic arm while maintaining acoustic coupling between the ultrasound imaging probe and the skin or membrane. 50. The method of claim 44, further comprising, prior to initiating ultrasound therapy: determining if the ultrasound imaging probe is at least partially retracted to a specified position; and initiating ultrasound therapy only if the ultrasound imaging probe is in the specified position. 51. The method of claim 50, wherein the specified position comprises a fully retracted position. 52. The method of claim 50, wherein the specified position comprises a position that does not substantially obstruct transducer elements of the ultrasound treatment head. 53. The method of claim 50, wherein the specified position comprises a position that does not substantially obstruct histotripsy pulses from the ultrasound treatment head. 54. An ultrasound treatment system, comprising: a robotic positioning arm controlled by a robotic positioning system; a treatment head disposed on the robotic positioning arm, the treatment head comprising: a therapy transducer array configured to deliver ultrasound pulses to a focal location; an ultrasound imaging probe sized and configured for axial and rotational movement within a bore of the therapy transducer array; a linear encoder configured to determine an axial position of the ultrasound imaging probe; and, a rotational encoder configured to determine a rotational position of the ultrasound imaging probe; at least one processor configured to communicate the axial position and rotational position of the ultrasound imaging probe to the robotic positioning system; one or more processors in communication with the robotic positioning system configured to receive the axial position and rotational position of the ultrasound imaging probe from the linear encoder and rotational encoder, respectively, and to register the axial position and rotational position of the ultrasound imaging probe with a digital treatment plan. 55. The ultrasound treatment system of claim 54, further comprising a position sensor configured to determine if the ultrasound imaging probe is at a specified axial position with respect to the therapy transducer array. | 2022 | HISTOSONICS | Medical and Healthcare | Small-size company | http://www.histosonics.com/ | USA | The patent focuses on a histotripsy therapy system for treatment of tissue, with features for acoustically coupling the system to the patient's skin and supporting non-invasive procedures. It addresses the gap in non-invasive therapeutic options for many conditions, where conventional ultrasound therapy can have coupling issues. The system enables effective non-invasive and minimally invasive histotripsy therapy. | הפטנט מתמקד במערכת היסטוטריפסיה לטיפול ברקמה, עם תכונות להצמדה אקוסטית של המערכת לעור המטופל ולתמיכה בפרוצדורות לא פולשניות. הוא מתמודד עם הפער באפשרויות טיפול לא פולשניות במצבים רבים, שבהם טיפול אולטרסאונד קונבנציונלי יכול להיות בעל בעיות הצמדה. המערכת מאפשרת טיפול היסטוטריפסיה לא פולשני וזעיר-פולשני יעיל. |  |  | 1 | [12 empty] | 1 |  |  |  |  | 1 | [9 empty] | 1 |  | HISTOSONICS INC | HISTOSONICS INC | HISTOSONICS INC | HISTOSONICS INC | HISTOSONICS INC | 2023-01-24 | 2023 | 2024-12-04 | 2024 | 2022-01-24 | 2022 | Application | Utility Invention | active | EP20230744017 | 321144902 | EP4469152A4¦US20250152135A1¦WO2023141653A2¦WO2023141653A3 | EP
469 | IL302743A | Heart valve ablation catheter | Cardiac annuloplasty methods and devices, based on delivery of tissue-ablating energy to a heart valve annulus, thereby inducing reduction of the valve annulus perimeter. In some embodiments, the reduction is induced by shrinkage of tissue in response to ablation, potentially analogous to tissue shrinkage responsible for pulmonary vein stenosis induced by cardiac ablation to treat atrial fibrillation. In some embodiments, annulus tissue is deformed before ablation energy is applied. This potentially results in plastic deformation apart from tissue shrinkage. Deformation is optionally performed using electrodes that also operate as sharp-tipped jaws of a pliers. They are inserted to tissue in a wider-spaced configuration, reduced in distance to a narrower-spaced tissue-squeezing configuration, and operated to ablate the squeezed tissue. Upon removal of the electrodes, the squeezed tissue retains a new shape as a result of ablation-induced plastic deformation. | 1. A method of performing a cardiac annuloplasty procedure, comprising delivering energy to a perimeter of a heart valve annulus, in an amount sufficient to structurally disrupt tissue and induce shrinkage of the valve annulus perimeter to reduce regurgitation through the heart valve. 2. The method of claim 1, wherein the structural disruption of tissue comprises changes to the fibrotic structure of the tissue. 3. The method of claim 1, comprising mechanically deforming tissue of the heart valve annulus, and wherein the energy disrupts the tissue while the heart valve annulus tissue in its mechanically deformed condition. 4. The method of claim 3, wherein the mechanically deforming comprises compressing on the tissue. 5. The method of claim 4, wherein the compressing the tissue comprises piercing the tissue with at least one sharpened element, and manipulating the sharpened element to exert the compression. 6. The method of claim 5, wherein the at least one sharpened element also comprises an element used to deliver the structurally disruptive energy to the tissue. 7. The method of claim 6, wherein the element is an electrode. 8. The method of claim 7, wherein the electrode disrupts the tissue structure by transmitting RF energy into the tissue. 9. The method of claim 8, wherein the electrode disrupts structure of the tissue by induction of cellular death. 10. The method of any one of claims 8-9, wherein the electrode disrupts structure of the tissue by coagulation. 11. The method of any one of claims 5-8, wherein the compressing comprises pinching the tissue between a plurality of the at least one sharpened element. 12. The method of any one of claims 5-11, wherein the compressing comprises exerting torsion on the tissue using the at least one sharpened element. 13. The method of any one of claims 1-12, wherein the reduction of the valve annulus perimeter comprises tissue shrinkage as a result of the delivery of tissue-ablating energy. 14. The method of any one of claims 1-13, wherein the reduction of the valve annulus perimeter comprises plastic deformation of tissue as a result of the delivery of tissue-ablating energy while the tissue is mechanically deformed. 15. The method of claim 1, comprising piercing tissue with an electrode; wherein the delivering tissue-ablating energy comprises delivery radiofrequency energy through the electrode into the pierced tissue. 16. The method of claim 1, comprising contacting a tissue surface with an electrode; wherein the delivering tissue-ablating energy comprises delivery radiofrequency energy through the electrode into the contacted tissue. 17. The method of claim 1, wherein the tissue-ablating energy is provided by at least one of the group consisting of: radiofrequency energy; focused ultrasound energy; cryogenic cooling. 18. The method of any one of claims 1-17, wherein the tissue ablated comprises at least one of the group consisting of: fibrous tissue of the valve annulus; and tissue of the heart wall adjacent to the fibrous tissue of the valve annulus. 19. The method of any one of claims 1-18, wherein the valve annulus is a valve annulus of mitral valve or a tricuspid valve. 20. The method of any one of claims 1-19, comprising repeating the delivering of tissueablating energy at a plurality of sites along the perimeter of the heart valve annulus. 21. The method of any one of claims 1-20, comprising: selecting a patient with an enlarged heart valve perimeter; planning a targeted reduction in heart valve annulus perimeter, including selection of locations along the heart valve annuls targeted for shrinkage; and performing the delivering energy at each of the selected locations. 22. The method of any one of claims 1-21, wherein at least a portion of the shrinkage occurs during the cardiac annuloplasty procedure. 23. The method of any one of claims 1-22, wherein at least a portion of the shrinkage occurs after the cardiac annuloplasty procedure. 24. A method of performing cardiac annuloplasty, comprising: piercing tissue along a perimeter of a heart valve with at least one electrode; applying mechanical force to the at least one electrode to deform the pierced tissue and reduce the perimeter of the heart valve annulus; delivering tissue-ablating energy through the electrode, thereby inducing plastic deformation of the deformed tissue; and releasing the mechanical force, leaving the heart valve annulus with a reduced perimeter. 25. The method of claim 24, wherein the applying mechanical force comprises placing torsion on the pierced tissue. 26. The method of any one of claims 24-25, wherein the applying mechanical force comprises compressing the pierced tissue. 27. The method of any one of claims 24-26, wherein the tissue-ablating energy is radiofrequency energy. 28. The method of any one of claims 24-27, wherein the delivering tissue-ablating energy induces plastic deformation by coagulation. 29. The method of any one of claims 24-28, wherein the reduced perimeter draws leaflets of the heart valve into positions that reduce regurgitation of the valve. 30. The method of claim 29, wherein the regurgitation reduction comprises restoration of coaptation between the leaflets of the heart valve. 31. The method of any one of claims 24-30, wherein the plastic deformation comprises shrinkage of the deformed tissue. 32. A device for annuloplasty treatment, comprising: a catheter, sized for transvascular insertion to a heart chamber from a percutaneous incision to reach a heart valve annulus thereof; at least one tissue penetrating element at a distal end of the catheter; wherein the at least one penetrating element both moves relative to a body of the catheter, and acts to deliver tissue disrupting energy to penetrated tissue of the heart valve annulus. 33. The device of claim 32, wherein the at least one penetrating element comprises a plurality of penetrating elements, adjustable in their relative distance while inserted to tissue of the heart valve annulus. 34. The device of any one of claims 32-33, wherein each of the at least one tissue penetrating elements is an electrode electrically interconnected to a connection remaining outside the percutaneous incision when the catheter is inserted to the heart chamber. 35. The device of any one of claims 33-34, wherein each of the plurality of penetrating elements operates as an ablation electrode. 36. The device of any one of claims 33-35, wherein the penetrating elements are spaced to insert to the tissue at a relatively wider distance, and adjust to a narrower distance. 37. The device of any one of claims 33-36, wherein the relative distance of the penetrating elements is adjusted by rotation of a gear. 38. The device of claim 37, wherein the gear is rotated by a control element leading to a proximal side of the catheter. 39. The device of claim 38, wherein the control element also acts to provide electrical interconnection between at least one of the tissue penetrating elements and a source of electrical power which remains outside of the percutaneous incision. 40. The device of any one of claims 33-36, wherein the relative distance of the penetrating elements is adjusted by a temperature change of an actuator comprising a shape memory alloy. 41. The device of claim 40, wherein the shape memory allow is positioned so that the temperature change is induced by heating consequent to operation of the penetrating elements as electrodes. 42. The device of any one of claims 40-41, wherein the shape memory alloy is shaped to move the penetrating elements from an initial distance to a relatively narrower distance when heated. 43. The device of any one of claims 40-42, comprising a device reset, actuatable to restore the distance of the penetrating elements before the temperature change, while the device remains inserted to the heart chamber. 44. The device of any one of claims 32-43, comprising: an inner component terminating distally in an energy delivering segment, and housed within an outer tube; the outer tube being sized for insertion to the heart chamber from within a guiding catheter; wherein the outer tube is provided with a predetermined distal bend which it assumes when unconfined by the guiding catheter, and which straightens when the outer tube is withdrawn into the guiding catheter. 45. The device of claim 32, wherein the at least one penetrating element has a noncircular cross-section that engages with and induces torsion in tissue to which it is inserted, upon receiving torque exerted through the catheter. 46. The device of claim 45, comprising: an inner component terminating distally in a tissue ablation segment, and housed within an outer tube; the outer tube being sized for insertion to the heart chamber from within a guiding catheter; wherein the outer tube is provided with a predetermined distal bend which it assumes when unconfined by the guiding catheter, and which straightens when the outer tube is withdrawn into the guiding catheter. 47. The device of any one of claims 45-46, wherein the non-circular cross-section comprises a rectangular blade. 48. The device of any one of claims 45-46, wherein the non-circular cross-section comprises three or more blades radiating from a central axis. | 2020 | BIO REFINE | Medical and Healthcare | Start up | https://www.incentive-il.com/biorefine | Israel | The patent focuses on a heart valve ablation catheter that delivers tissue-ablating energy to a heart valve annulus, inducing reduction of the annulus perimeter via tissue shrinkage. It addresses the limitations of conventional annuloplasty methods, where mechanical devices are often needed. Ablation-induced shrinkage offers a less invasive way to reduce valve annulus size. | הפטנט מתמקד בקטטר אבלציה של שסתום הלב המספק אנרגיה לאבלציית רקמה לטבעת שסתום הלב, וגורם להפחתת היקף הטבעת באמצעות התכווצות רקמה. הוא מתמודד עם המגבלות של שיטות אנולופלסטיקה קונבנציונליות, שבהן לעיתים קרובות נדרשים מכשירים מכניים. התכווצות מושרית-אבלציה מציעה דרך פחות פולשנית להקטין את גודל טבעת השסתום. | [5 empty] | 1 | [18 empty] | 1 | [7 empty] | BIO REFINE LTD | BIO REFINE LTD | BIO REFINE LTD | BIO REFINE LTD | BIO REFINE LTD | 2023-05-07 | 2023 | 2023-07-01 | 2023 | 2020-11-08 | 2020 | Application | Utility Invention | active | IL20230302743 | 322644300 | AU2021373380A1¦CA3195812A1¦CN116669642A¦EP4240265A1¦EP4240265A4¦JP2023548175A¦JP7713014B2¦JP7713014B6¦US20240008915A1¦WO2022097130A1 | IL
470 | WO2023178247A1 | APPARATUSES AND METHODS FOR THE CONTROL AND OPTIMIZATION OF ICE FORMATION DURING CRYOABLATION TREATMENTS | A method of performing a cryoablation treatment may include positioning 804 a plurality of measurement points in predetermined locations relative to a target tissue in a patient and obtaining 808 ice formation measurement information from the plurality of measurement points. The method may also include comparing 810 the ice formation measurement information to a predetermined ice formation plan and adjusting 812 a flow of a cryo-fluid to a cryoprobe if the ice formation measurement information deviates from the predetermined ice formation plan by more than predetermined deviation level. | 1. CLAIMSWhat is claimed is: 1. A system for performing a cryoablation treatment comprising at least one computing device, the at least one computing device configured to: obtain ice formation measurement information from a plurality of measurement points, the plurality of measurement points positioned at predetermined locations relative to a target tissue in a patient; compare the ice formation measurement information to a predetermined ice formation plan; and adjust a flow of a cryo-fluid to a cryoprobe if the ice formation measurement information deviates from the predetermined ice formation plan by more than a predetermined deviation level. 2. The system of claim 1 , wherein the plurality of measurement points are positioned on the cryoprobe. 3. The system of claim 1 or 2, wherein the plurality of measurement points are positioned on a measurement lead, the measurement lead positioned in a desired orientation relative to the cryoprobe. 4. The system of claim 1 , wherein the plurality of measurement points are positioned on both a cryoprobe and a measurement lead. 5. The system of any one of claims 1 to 4, wherein the ice formation measurement information comprises temperatures at the plurality of measurement points, and the at least one computing device compares the temperatures to temperature thresholds. 6. The system of any one of claims 1 to 5, wherein the at least one computing device determines a rate of change of a temperature at each measurement point of the plurality of measurement points and compares each rate of change of the temperature to a temperature profile. 7. The system of any one of claims 1 to 6, wherein the at least one computing device is configured to adjust at least one of a flow speed or a flow volume of the cryofluid flow. 8. The system of any one of claims 1 to 7, wherein the at least one computing device is configured to adjust at least one of a pulse width, a pulse amplitude or pulse frequency of the cryo-fluid flow. 9. The system of any one of claims 1 to 8, wherein the at least one computing device is configured to continuously obtain the ice formation measurement information and continuously adjust the flow of the cryo-fluid to the cryoprobe if the ice formation measurement information deviates from the predetermined ice formation plan until the ice formation measurement information that a desired ice formation has been achieved for predetermined freezing time. 10. The system of any one of claims 1 to 9, further comprising a pump fluidly connected to a cryo-fluid source and a cryo-fluid supply line, the at least one computing device configured to adjust the flow of the cryo-fluid to the cryoprobe using the pump. 11 . A method of performing a cryoablation treatment comprising: positioning a plurality of measurement points in predetermined locations relative to a target tissue in a patient; obtaining ice formation measurement information from the plurality of measurement points; comparing the ice formation measurement information to a predetermined ice formation plan; and adjusting a flow of a cryo-fluid to a cryoprobe if the ice formation measurement information deviates from the predetermined ice formation plan by more than a predetermined deviation level. 12. The method of claim 11 , wherein the plurality of measurement points are located on the cryoprobe. 13. The method of claim 11 , wherein the plurality of measurement points are located on a measurement lead, the measurement lead positioned in a desired orientation relative to the cryoprobe. 14. The method of claim 11 , wherein the plurality of measurement points are located on a cryoprobe and a measurement lead. 15. The method of any one of claims 11 to 14, wherein the ice formation measurement information comprises temperatures at the plurality of measurement points, and the temperatures are compared to temperature thresholds. 16. The method of any one of claims 11 to 15, wherein the step of comparing the ice formation measurement information to the predetermined ice formation plan comprises determining a rate of change of a temperature at each measurement point of the plurality of measurement points and comparing each rate of change of the temperature to a temperature profile. 17. The method of any one of claims 11 to 16, wherein the step of adjusting the flow of the cryo-fluid comprises adjusting at least one of a flow speed and a flow volume. 18. The method of any one of claims 11 to 17, wherein the step of adjusting the flow of the cryo-fluid comprises adjusting at least one of a pulse width, a pulse amplitude or pulse frequency of the cryo-fluid flow. 19. The method of any one of claims 11 to 18, further comprising continuously obtaining the ice formation measurement information and continuously adjusting the flow of the cryo-fluid to the cryoprobe if the ice formation measurement information deviates from the predetermined ice formation plan until the ice formation measurement information that a desired ice formation has been achieved for predetermined freezing time. 20. The method of any one of claims 11 to 19, wherein the step of adjusting the flow of the cryo-fluid is automatically performed by the at least one computing device coupled to pump that is fluidly connected to a cryo-fluid source and a cryo-fluid supply line. | 2022 | Varian Medical Systems | Medical and Healthcare | Large-Size Company | https://www.varian.com | USA | The patent focuses on a cryoablation method that positions multiple measurement points around target tissue, obtains ice formation info, and compares it to a planned ice formation. It addresses the difficulty of optimizing ice growth during cryoablation, where freeze-zone shape can vary unpredictably. Comparing actual to planned ice formation guides the procedure to the intended outcome. | הפטנט מתמקד בשיטת קריו-אבלציה הממקמת מספר נקודות מדידה סביב רקמת המטרה, משיגה מידע על היווצרות קרח, ומשווה אותו להיווצרות קרח מתוכננת. הוא מתמודד עם הקושי באופטימיזציה של גידול קרח במהלך קריו-אבלציה, שבה צורת אזור ההקפאה יכולה להשתנות באופן בלתי צפוי. השוואת היווצרות קרח אמיתית למתוכננת מנחה את הפרוצדורה לתוצאה המיועדת. |  |  |  |  | 1 | [9 empty] | 1 | [6 empty] | 1 | [8 empty] | 1 |  | VARIAN MED SYS INC | Varian Medical Systems | VARIAN MED SYS INC | Varian Medical Systems | Varian Medical Systems | 2023-03-16 | 2023 | 2023-09-21 | 2023 | 2022-03-17 | 2022 | Application | Utility Invention | inactive | WO2023US64541 | 290993261 | CN118871048A¦EP4493087A1¦US12376897B2¦US20230293219A1 | WO
471 | US20240285252A1 | MEDICAL SYSTEMS, DEVICES, AND RELATED METHODS FOR IMAGING, ASSESSING, AND/OR TREATING A PROSTATE | Medical systems are described, including a medical system that includes a sheath and a catheter. The sheath is configured to be positioned within a urethra of a patient. A proximal portion of the sheath includes a port configured to receive one or more fluids. The catheter is insertable into the sheath and configured to be movably positioned within the sheath in the urethra. A distal portion of the catheter includes one or more ultrasound transducers configured to perform ultrasound imaging of at least a portion of a prostate of the patient. | 1. A medical system, comprising: a sheath configured to be positioned within a urethra of a patient, wherein a proximal portion of the sheath includes a port configured to receive one or more fluids; and a catheter insertable into the sheath and configured to be movably positioned within the sheath in the urethra; wherein a distal portion of the catheter includes one or more ultrasound transducers configured to perform ultrasound imaging of at least a portion of a prostate of the patient. 2. The medical system of claim 1, further comprising an insertion device with first and second lumens, one of the first lumen or second lumen being configured to receive the sheath or the catheter. 3. The medical system of claim 2, further comprising a needle, wherein at least a portion of the needle is configured to be movably positioned within the other of the first lumen or the second lumen of the insertion device. 4. The medical system of claim 3, wherein a distal portion of the needle includes a negative electrode portion and a positive electrode portion, the needle being configured to be energized for irreversible electroporation therapy. 5. The medical system of claim 4, wherein the distal portion of the needle includes one or more echo markings. 6. The medical system of claim 2, wherein a distal portion of one or both of the first lumen or the second lumen is sharpened or tapered. 7. The medical system of claim 3 wherein a proximal portion of the insertion device includes a female luer connection, and wherein a proximal portion of the needle includes a male luer connection. 8. The medical system of claim 2, wherein the insertion device further includes a connection portion, an umbilicus, and an electronic connector to operably couple insertion device and the catheter to a console. 9. The medical system of claim 3, wherein the needle is an ultrasonic needle coupled to a piezoelectric transducer. 10. The medical system of claim 9, wherein the piezoelectric transducer is configured to convey one or more of longitudinal movement, lateral movement, or rotational movement to the ultrasonic needle. 11. The medical system of claim 1, wherein the sheath is flexible, and wherein a distal end of the sheath is closed. 12. The medical system of claim 9, wherein the sheath is translucent to ultrasound waves. 13. The medical system of claim 1, further comprising a needle configured to be delivered to the prostate of the patient transperineally, the needle comprising at least one light source. 14. The medical system of claim 1, further comprising one or more light sources positioned on the distal portion of the catheter. 15. The medical system of claim 1, wherein the distal portion of the catheter includes two transducers. 16. A medical system, comprising: an ultrasound probe insertable into a rectum of a patient, the ultrasound probe including a distal end comprising at least one transducer, wherein the at least one transducer, when energized, forms an ultrasound field of view that at least partially overlaps with a prostate of the patient; an illumination device insertable into the rectum or a perineum of the patient, the illumination device including at least one light source that, when energized, forms an illumination area that at least partially overlaps with the prostate of the patient; and a needle, wherein a distal end of the needle is configured to biopsy or treat the prostate of the patient. 17. The medical system of claim 16, wherein the illumination device is coupled to the ultrasound probe, and wherein the illumination device includes two light sources coupled to a distal end of the ultrasound probe, and wherein the two light sources are adjacent to the transducer. 18. The medical system of claim 16, wherein the illumination device includes the needle, the illumination device including two light sources coupled to a distal portion of the needle, wherein the ultrasound probe is configured to be delivered transrectally, and wherein the needle is configured to be delivered transperineally. 19. A method of treating a patient, comprising: inserting a sheath and a catheter through a urethra of the patient, wherein the sheath is flexible and includes a port in a proximal portion of the sheath, the sheath having a closed distal end, and wherein a distal portion of the catheter includes at least one ultrasound transducer; delivering a fluid to the sheath through the port such that the fluid is between the catheter and the sheath; imaging a prostate of the patient by activating or energizing the at least one ultrasound transducer of the catheter and longitudinally or rotatably moving the catheter; and removing the catheter and the sheath from the urethra. 20. The method of claim 19, wherein imaging the prostate includes obtaining multiple images by advancing the catheter distally, activating or energizing the at least one ultrasound transducer multiple times, and longitudinally or rotatably moving the catheter to ultrasonically image different portions of the prostate; and wherein the method further comprises applying pressure to the prostate transrectally and performing one or more elastography procedures. | 2023 | Boston Scientific Corporation (BSC) | Medical and Healthcare | Large-Size Company | https://www.bostonscientific.com | USA | The patent focuses on a urethral medical system with a sheath positioned in the urethra and a catheter with ultrasound transducers for prostate imaging. It addresses the limited integration of ultrasound imaging with prostate procedures from the urethral approach. The ultrasound transducers enable real-time imaging of the prostate during minimally invasive procedures. | הפטנט מתמקד במערכת רפואית שופכתית עם נדן הממוקם בשופכה וקטטר עם מתמרי אולטרסאונד להדמיית הערמונית. הוא מתמודד עם האינטגרציה המוגבלת של הדמיית אולטרסאונד עם פרוצדורות ערמונית מהגישה השופכתית. מתמרי האולטרסאונד מאפשרים הדמיה בזמן אמת של הערמונית במהלך פרוצדורות זעיר-פולשניות. |  | 1 | 1 | [17 empty] | 1 | [10 empty] | 1 | BOSTON SCIENT SCIMED INC | Boston Scientific Corporation (BSC) | BOSTON SCIENT SCIMED INC | Boston Scientific Corporation (BSC) | Boston Scientific Corporation (BSC) | 2024-02-23 | 2024 | 2024-08-29 | 2024 | 2023-02-28 | 2023 | Application | Utility Invention | active | US202418585493 | 314526758 | CN120787140A¦EP4673057A1¦WO2024182227A1 | US
472 | WO2025117508A1 | PHOTO OR IMAGE ASSISTED MEDICAL DEVICES FOR DELIVERING ENERGY AND/OR FLUID TO TISSUE AND RELATED METHODS | A medical device includes an insertion portion, a light-emitting optical fiber, and a light-receiving optical fiber. The insertion portion includes a shaft, and the shaft includes a sheath and an electrode. The sheath includes a sheath lumen, and the electrode is positioned within at least a distal portion of the sheath. The electrode includes an electrode lumen, and the electrode is configured to deliver energy to a target site. The light-emitting optical fiber is configured to deliver light to the target site. At least a portion of the light-emitting optical fiber is positioned within the sheath or the electrode lumen. The light-receiving optical fiber is configured to receive light from the target site. At least a portion of the light-receiving optical fiber is positioned within the electrode lumen. | 1. CLAIMSWe claim: 1 . A medical device, comprising: an insertion portion, including a shaft that includes a sheath and an electrode, wherein the sheath includes a sheath lumen, wherein the electrode is positioned within at least a distal portion of the sheath, wherein the electrode includes an electrode lumen, and wherein the electrode is configured to deliver energy to a target site; a light-emitting optical fiber configured to deliver light to the target site, wherein at least a portion of the light-emitting optical fiber is positioned within the sheath or the electrode lumen; and a light-receiving optical fiber configured to receive light from the target site, wherein at least a portion of the light-receiving optical fiber is positioned within the electrode lumen. 2. The medical device of claim 1 , wherein the electrode has a “T” shaped cross-section with an electrode shaft and a widened distal portion, wherein a portion of the electrode extends distally of the sheath lumen, and wherein the widened distal portion is wider than an opening of the sheath lumen. 3. The medical device of claim 2, wherein at least a portion of the sheath is configured to convey light and form the light-emitting optical fiber. 4. The medical device of any one of claims 1 -3, a distal end of the sheath includes a first width, wherein a distal end of the electrode includes a second width, and wherein the second width is less than the first width. 5. The medical device of claim 4, wherein a width of the electrode lumen is greater than a width of the light-receiving optical fiber, such that a portion of the electrode lumen forms a fluid delivery portion that occupies a space adjacent to the light-receiving optical fiber. 6. The medical device of any one of claims 1-5, wherein the light-emitting optical fiber is positioned within the electrode lumen. 7. The medical device of claim 6, wherein the sheath includes one or more fluid delivery or suction application lumens. 8. The medical device of claim 7, wherein the one or more fluid delivery or suction application lumens includes at least two fluid delivery or suction application lumens. 9. The medical device of claim 7 or claim 8, wherein respective distal openings of the one or more fluid delivery or suction application lumens are at least partially radially outside of a distal end of the electrode. 10. The medical device of any one of claims 7-9, wherein the sheath includes the light-emitting optical fiber and the one or more fluid delivery or suction application lumens. 11 . The medical device of any one of claims 1 -10, wherein the medical device includes a handle, wherein the light-emitting optical fiber extends from the handle through a portion of the shaft to deliver light to the target site. 12. The medical device of any one of claims 1-11 , wherein the handle of the medical device includes a port, wherein the port is coupled to a light source that provides light to the light-emitting optical fiber. 13. The medical device of any one of claims 1 -12, wherein the lightreceiving optical fiber extends through the electrode lumen, through a shaft lumen, to the handle, and to a processor, wherein the processor analyzes light received by the light-receiving optical fiber to generate a light profile. 14. The medical device of claim 13, wherein based on the generated light profile, the processor is configured to deactivate the electrode. 15. The medical device of claims 13 or 14, wherein when the processor detects a red light the processor sends a signal to disable energization of the electrode. | 2023 | Boston Scientific Corporation (BSC) | Medical and Healthcare | Large-Size Company | https://www.bostonscientific.com | USA | The patent focuses on a photo or image-assisted medical device with an insertion shaft, electrode in the sheath, light-emitting and light-receiving optical fibers for tissue energy/fluid delivery. It addresses the limitation of energy/fluid devices that lack integrated optical guidance. The integrated optics enable photo-assisted, image-guided energy and fluid delivery to tissue. | הפטנט מתמקד במכשיר רפואי בעזרת תמונה או צילום עם פיר החדרה, אלקטרודה בנדן, סיבים אופטיים פולטי אור ומקבלי אור למסירת אנרגיה/נוזל לרקמה. הוא מתמודד עם המגבלה של מכשירי אנרגיה/נוזל החסרים הנחיה אופטית משולבת. האופטיקה המשולבת מאפשרת מסירת אנרגיה ונוזל לרקמה בעזרת תמונה ומונחית הדמיה. |  | 1 | [17 empty] | 1 | [10 empty] | 1 |  | BOSTON SCIENT SCIMED INC | Boston Scientific Corporation (BSC) | BOSTON SCIENT SCIMED INC | Boston Scientific Corporation (BSC) | Boston Scientific Corporation (BSC) | 2024-11-26 | 2024 | 2025-06-05 | 2025 | 2023-11-28 | 2023 | Application | Utility Invention | active | WO2024US57404 | 321465898 | US20250169675A1 | WO
473 | US20220323142A1 | METHODS, SYSTEMS AND DEVICES FOR ENDOVASCULAR ELECTROPORATION OF A GREATER SPLANCHNIC NERVE | Systems, devices, and methods for transvascular ablation of target tissue. The devices and methods may, in some examples, be used for splanchnic nerve ablation to increase splanchnic venous blood capacitance to treat at least one of heart failure and hypertension. For example, the devices disclosed herein may be advanced endovascularly to a target vessel in the region of a thoracic splanchnic nerve (TSN), such as a greater splanchnic nerve (GSN) or a TSN nerve root. Also disclosed are methods of treating heart failure, such as HFpEF, by endovascularly ablating a thoracic splanchnic nerve to increase venous capacitance and reduce pulmonary blood pressure. | 1. A method of deactivating a GSN or GSN root of a patient, the method comprising: delivering an ablation catheter to a target zone in one of the patient's T9, T10 or T11 intercostal veins; delivering electroporation energy from the ablation catheter; deactivating the GSN or GSN root; and removing the ablation catheter from the patient. 2. The method of claim 1, wherein delivering electroporation energy temporarily deactivates one or more pain nerves in the target zone. 3. The method of claim 1 wherein deactivating the GSN or GSN root comprises delivering ablation energy from the ablation catheter. 4. The method of claim 3, wherein the ablation catheter comprises one or more energy delivery elements carried by an elongate shaft, wherein the electroporation energy and the ablation energy are delivered using the one or more energy delivery elements. 5. The method of claim 4, wherein the one or more energy delivery elements comprise one or more electrodes carried by the elongate shaft. 6. The method of claim 1, wherein deactivating the GSN or GSN root comprises ablating the GSN or GSN root. 7. The method of claim 6, wherein ablating the GSN or GSN root comprises delivering RF ablation energy from the ablation catheter. 8. The method of claim 6, wherein delivering electroporation energy from the ablation catheter causes the GSN or GSN root to be ablated. 9. The method of claim 1, further comprising administering fentanyl or midazolam to the patient. 10. The method of claim 1, further comprising delivering saline from the ablation catheter. 11. The method of claim 1, further comprising: delivering a nerve stimulation signal from the ablation catheter and monitoring a first physiological reaction prior to delivering the electroporation energy. 12. The method of claim 11, further comprising delivering a second nerve stimulation signal from the ablation catheter and monitoring a second physiological reaction after delivering the electroporation energy. 13. The method of claim 12, further comprising repeating the delivering electroporation energy step if the second physiological reaction is not reduced compared to the first physiological reaction. 14. The method of claim 1, wherein delivering an ablation catheter to a target zone in one of the patient's T9, T10 or T11 intercostal veins comprises delivering the ablation catheter to a segment of the T9, T10 or T11 intercostal vein that is adjacent to an adjoining azygos vein. 15. The method of claim 14, wherein the segment is from 5 mm to 20 mm long. 16. The method of claim 14, wherein the segment is from 10 mm to 20 mm long. 17. The method of claim 1, wherein delivering electroporation energy from the ablation catheter comprises delivering energy from an electroporation console with delivery parameters comprising: a maximum power in a range of 30 W to 35 W, a current that is one of DC or RF, a pulse frequency in a range of 2 Hz to 4 Hz, and a pulse on duration in a range of 0.5 ms to 2 ms. 18. The method of claim 1, wherein delivering the electroporation energy comprises delivering electroporation energy in a duration in a range of 1 s to 120 s. 19. The method of claim 1, wherein delivering the ablation catheter to a target zone in one of the patient's T9, T10 or T11 intercostal veins comprises delivering an ablation catheter that includes distal and proximal electrically conductive flexible and coiled ablation electrodes carried by a distal linear section of the ablation catheter. 20. The method of claim 19, wherein the distal and proximal electrically conductive flexible and coiled ablation electrodes together having an axial length from 5 mm-25 mm and an axial spacing therebetween that is not more than 2.0 mm. 21. The method of claim 19, wherein the distal and proximal electrically conductive flexible and coiled ablation electrodes each comprise a dielectric coating. | 2020 | AXON VASCULAR INC | Medical and Healthcare | Start up | http://www.axontherapies.com/ | USA | The patent focuses on endovascular electroporation of the greater splanchnic nerve through a catheter advanced into a thoracic intercostal vein (T9, T10, or T11) to treat heart failure and hypertension. It addresses the difficulty of reaching and modulating the splanchnic nerves with conventional approaches for heart failure or pain. Endovascular electroporation deactivates the GSN to increase splanchnic venous capacitance. | הפטנט מתמקד באלקטרופורציה אנדו-וסקולרית של העצב הספלנכני הגדול דרך קטטר המקודם לוריד תוך-צלעי חזי (T9, T10 או T11) לטיפול באי ספיקת לב ויתר לחץ דם. הוא מתמודד עם הקושי בהגעה ובמודולציה של העצבים הספלנכניים בגישות קונבנציונליות לאי ספיקת לב או כאב. אלקטרופורציה אנדו-וסקולרית משביתה את ה-GSN להגדלת קיבולת ורידית ספלנכנית. | 1 | [10 empty] | 1 | [17 empty] | 1 |  |  | AXON THERAPIES INC | AXON THERAPIES INC | AXON VASCULAR, INC. | AXON VASCULAR INC | AXON VASCULAR INC | 2021-12-17 | 2021 | 2022-10-13 | 2022 | 2020-12-23 | 2020 | Application | Utility Invention | active | US202117644998 | 285113294 | [NA] | US
474 | WO2026035967A1 | METHODS AND SYSTEMS FOR TARGETING & ASSESSING UROLOGY ORGAN TREATMENT | Systems and methods are provided for treating indications associated with a urological organ (e.g., bladder, kidney, prostate, ureter, or urethra, etc.). | 1. Atorney Docket No. 58496-0002W01WHAT IS CLAIMED IS: 1. A method for treating a prostate, bladder, kidney, ureter, or urethra of a patient, comprising the steps of: inserting a perfusion visualization system into a bladder, kidney, ureter, or urethra of the patient; inserting a non-mucosal-disrupting ablation system into the patient, ablating a portion of a prostate, bladder, kidney, ureter, or urethra of the patient using the non-mucosal-disrupting ablation system; and using the perfusion visualization system to detect perfusion or lack of perfusion in real time in the prostate, bladder, kidney, ureter, or urethra to determine efficacy of the ablating step based on the visualization system; and stopping or continuing treatment according to the perfusion or lack of perfusion seen in real time. 2. The method of claim 1, wherein the perfusion visualization system is delivered trans-urethrally. 3. The method of claim 2, wherein the visualization system is an endoscope or a cystoscope. preferably a flexible or rigid cystoscope, or a flexible or rigid ureteroscope. 4. The method of claim 1, wherein using the perfusion visualization system comprises detecting a dye-based reagent, a dye-free tool, or a combination of both a dye-based reagent and a dye-free tool. 5. The method of claim 4, wherein using the visualization system comprises detecting a combination of both a dye-based reagent and a dye-free tool. 6. The method of claim 5, wherein the dye-free tool is laser speckle. 7. The method of claim 5, further comprising intravenously administering the dye-based reagent to the patient. Atorney Docket No. 58496-0002W018. The method of claim 7, wherein the dye-based reagent is administered immediately before or immediately after the ablating step. 9. The method of claim 8, wherein the dye-based reagent is administered 30 seconds to 30 minutes prior to the ablating step. 10. The method of claim 8. wherein the dye-based reagent is administered 30 seconds to 30 minutes after the ablating step. 11. The method of any one of claims 7 to 10, wherein the dye-based reagent appears in the prostate, bladder, kidney, ureter, or urethra within 15 seconds to 5 minutes of administering the dye-based reagent to the patient, and the perfusion visualization system is used to detect perfusion or lack of perfusion in real time in the prostate, bladder, kidney, ureter, or urethra within 15 seconds to 5 minutes of the administration. 12 The method of any one of claims 7 to 11, wherein the dye-based reagent clears from the patient's prostate, bladder, kidney, ureter, or urethra within 10 minutes to 30 minutes of administering the dye-based reagent to the patient. 13. The method of any one of claims 7 to 12, wherein the dye-based reagent is indocyanine green (ICG) angiography, riboflavin, or fluoresceine. 14. The method of claim 13, wherein the dye is ICG. 15. The method of claim any one of claims 7 to 14, wherein the dye is administered to the patient at a concentration range of about 1 mg/ml to about 5 mg/ml. 16. The method of claim any one of claims 7 to 14, wherein the dose of the dye that is administered to the patient is about 0.5 to 10 mg, e.g., 0.5 mg, 1.0 mg, 1.5 mg, 2.0 mg, 2.5 mg, 3.0 mg, 3.5 mg, 4.0 mg, 4.5 mg, 5.0 mg, 5.5 mg, 6.0 mg, 6.5 mg, 7.0 mg. 7.5 mg, 8.0 mg, 8.5 mg, 9.0 mg. 9.5 mg, 10.0 mg (e.g., 5.0 mg). Attorney Docket No. 58496-0002W0117. The method of any one of claims 1 to 16, wherein the non-mucosal- disrupting ablation system is delivered trans-urethrally, trans-perineally, trans- abdominally, or trans-rectally. 18. The method of any one of claims 1 to 17, wherein the non-mucosal disrupting ablation system utilizes one of the following methods: water vapor/steam. laser, electroporation, cryoablation, high intensity focal ultrasound, or microwave methods. 19. The method of claim 18, wherein the non-mucosal disrupting ablation system is a water vapor/steam system and is delivered trans-urethrally. 20. The method of any one of the preceding claims, wherein the patient has benign prostatic hyperplasia or prostate cancer. 21. A system for ablating prostate, bladder, kidney, ureter, or urethra tissue, the system comprising: a visualization system for trans-urethral delivery into a bladder, kidney, ureter, or urethra collecting system, and anon-mucosal disrupting ablation system. 22. The system of claim 21, wherein the visualization system is an endoscope, a cystoscope, preferably a flexible or rigid cystoscope, or a flexible or rigid ureteroscope. 23. The system of claim 22, wherein the visualization system comprises using a dye-based reagent, a dye-free tool, or a combination of both a dye-based reagent and a dye-free tool. 24. The system of any one of claims 21 to 23, wherein the non-mucosal- disrupting ablation system utilizes one of the following methods: water vapor/steam, laser, electroporation, cryoablation, high intensity focal ultrasound, or microwave methods. Atorney Docket No. 58496-0002W0125. The system of any one of claims 21 to 24, wherein the non-mucosal- disrupting ablation system is for trans-urethral, trans-perineal, trans-abdominal, or trans-rectal delivery. | 2024 | AOA UROLOGICAL | Research & development | NA | http://www.auanet.org/ | USA | The patent focuses on systems and methods for treating urological organs (bladder, kidney, prostate, ureter, urethra) with targeting and assessment capabilities. It addresses the limitations of conventional treatments for benign prostatic hyperplasia and other urological conditions, where procedures lack integrated targeting and assessment. The system provides targeted treatment with real-time assessment of urological tissue. | הפטנט מתמקד במערכות ושיטות לטיפול באיברים אורולוגיים (שלפוחית, כליה, ערמונית, שופכן, שופכה) עם יכולות מיקוד והערכה. הוא מתמודד עם המגבלות של טיפולים קונבנציונליים להיפרפלזיה שפירה של הערמונית ומצבים אורולוגיים אחרים, שבהם פרוצדורות חסרות מיקוד והערכה משולבים. המערכת מספקת טיפול ממוקד עם הערכה בזמן אמת של רקמה אורולוגית. |  | 1 | [15 empty] | 1 | [13 empty] | 1 | AOA UROLOGICAL ASS LLC | AOA UROLOGICAL ASS LLC | AOA UROLOGICAL ASS LLC | AOA UROLOGICAL ASS LLC | AOA UROLOGICAL ASS LLC | 2025-08-07 | 2025 | 2026-02-12 | 2026 | 2024-08-08 | 2024 | Application | Utility Invention | active | WO2025US41142 | 326546892 | [NA] | WO
475 | WO2025072715A2 | NON-THERMAL TREATMENT OF ENDOMETRIOSIS | A device, system and methods for the non-thermal treatment of endometriosis. | 1. Arnold & Porter Ref. P35370WO00 CLAIMS: 1. A device for non-thermal treatment of one or more lesions in a subject in need thereof comprising: a generator 110 comprising a first output connector and a second output connector; an applicator 200, wherein applicator 200 is a articulating tip applicator 202 comprising a first electrode 210 and a second electrode 220; a dispersive pad 300; and a switch 410 in electrical connection with the generator, wherein the switch is capable of changing between a first position to apply an electric field of a first mode through the applicator and a second position to apply an electric field of a second mode through the applicator. 2. The device of claim 1, wherein the first mode is a monopolar mode that generates an electric field between the applicator and the dispersive pad. 3. The device of claim 2, wherein the switch 410 in the first position is configured to provide an electrical connection between the generator 110 through the first output connector with the first electrode 210 and the second electrode 220 of the applicator and provide an electrical connection between the generator through the second output connector with the dispersive pad 300. 4. The device of claim 3, wherein the first electrode 210 and the second electrode 220 are in electrical connection with each other. 5. The device of claim 3, wherein the dispersive pad 300 is configured to be in contact with the patient when the electric field is applied. 6. The device of claim 5, wherein the applicator 200 is configured to operate as a blunt tip for the treatment of the one or more lesions. 7. The device of claim 5, wherein the dispersive pad 300 is a large surface electrode. 8. The device of claim 1, wherein the second mode is a bipolar mode that generates an electric field between the first electrode 210 and the second electrode 220 of the applicator. Arnold & Porter Ref. P35370WO00 9. The device of claim 1, wherein the switch in the second position is configured to provide an electronic connection between the generator 110 through the first output connector with the first electrode 210 and provide an electronic connection between the generator 110 through the second output connector with the second electrode 220. 10. The device of claim 8, wherein there is no electrical connection between the generator 110 and the dispersive pad 300. 11. The device of claim 1, wherein the electric field of the first mode through the applicator 300 is capable of forming a treatment zone comprising the one or more lesions. 12. The device of claim 1, wherein the electric field of the second mode through the applicator is capable of forming a treatment zone comprising the one or more lesions. 13. The device of claim 1, wherein the first electrode 210 comprises a first stainless steel rod 212 surrounded by a first insulating layer 212i, a distal end 213, and a proximal end 214, wherein the second electrode 220 comprises a second stainless steel tube 222 surrounded by a second insulating layer 222i, a distal end 223, and a proximal end 224, wherein the first electrode 210 and the second electrode 220 are concentric, and wherein the proximal end 214 of the first electrode and the distal end 223 of the second electrode are separated by a distance in length 230w. 14. The device of claim 13, further comprising an insulating layer 230 between the first electrode 210 and the second electrode 220. 15. The device of claim 14, wherein the insulating layer 212 extends beyond the distal end of the first electrode by between 0 and 2 mm. 16. The device of claim 13, wherein the first electrode comprises a first outer diameter 216 of between 2 and 4 millimeters (mm). 17. The device of claim 13, wherein the first outer diameter 216 is approximately 3 mm. 18. Device of claim 13, wherein the second electrode 220 comprises a second outer diameter 226 of between 6 and 10 mm. 19. The device of claim 13, wherein the first electrode 210 comprises a wall thickness between 0 mm and 0.8 mm. Arnold & Porter Ref. P35370WO00 20. The device of claim 13, wherein the distance in length between the proximal end 214 of the first electrode and the distal end 223 of the second electrode is between 0 mm and 6 mm. 21. A device for an application of a non-thermal treatment dose comprising: a generator comprising a first output connector and a second output connector; an applicator comprising a first electrode and a second electrode; a first switch in electrical connection with the generator, wherein the switch is capable of changing between a first position to apply a therapeutic electric field of a first mode through the applicator and a second position to apply a therapeutic electric field of a second mode through the applicator; a dispersive pad in an electrical connection with the second output connector; a sensor configured to receive a first signal before the application of a non-thermal treatment dose and a second signal after the application of the non-thermal treatment dose; a processor electronically coupled to the switch and the sensor, a non-transitory computer readable medium electronically coupled to the processor, wherein the non-transitory computer readable medium comprises instructions stored thereon that, when executed on the processor, perform the steps of receiving the first signal from the sensor and the second signal from the sensor; and determining a difference between the first signal and the second signal, comparing the difference against a predetermined threshold; and instructing the application to end if the difference exceeds the predetermined threshold. 22. The device of claim 21, wherein the first switch is in a first electrical connection with the first output connector and in a second electrical connection with the first electrode 210. 23. The device of claim 22, wherein the instructing the application to end comprises a third signal to a second switch 420 upstream of the first switch and in a first electrical connection with the first output connector and in a second electrical connection with the first electrode, wherein the second switch is configured to disrupt the first electrical connection to the first output connector and the second electrical connection with the first electrode. 24. A device for an application of a non-thermal treatment dose comprising: a generator 110 comprising a first output connector and a second output connector; Arnold & Porter Ref. P35370WO00 an applicator 200 comprising a first electrode 210 in a first electrical connection with the first output and a second electrode 220 in a second electrical connection with the second output; a switch in a first electrical connection with the first output connector and in a second electrical connection with the first electrode, wherein the switch in a first position is configured to apply a therapeutic electric field through the first electrode, and wherein the switch in a second position is configured to disrupt the first electrical connection to the first output connector and the second electrical connection with the first electrode; a sensor configured to receive a first signal before the application of a non-thermal treatment dose and a second signal after the application of the non-thermal treatment dose; a processor electronically coupled to the switch and the sensor; a non-transitory computer readable medium electronically coupled to the processor, wherein the non-transitory computer readable medium comprises instructions stored thereon that, when executed on the processor, perform the steps of receiving the first signal from the sensor and the second signal from the sensor; and determining a difference between the first signal and the second signal, comparing the difference against a predetermined threshold; and 25. instructing the switch to configure to the second position. A device for non-thermal treatment comprising: a generator comprising a first output connector and a second output connector; an applicator comprising a first electrode and a second electrode; a dispersive pad; and a first switch 410 in electrical connection with the generator, wherein the switch is capable of changing between a first position to apply an electric field of a first mode through the applicator and a second position to apply an electric field of a second mode through the applicator; a second switch 420 upstream of the first switch and in a first electrical connection with the first output connector and in a second electrical connection with the first electrode, wherein the second switch is configured to disrupt the first electrical connection to the first output connector and the second electrical connection with the first electrode; a sensor configured to recognize a signal; a processor electronically coupled to the switch and configured to receive a second signal from the sensor; and Arnold & Porter Ref. P35370WO00 a non-transitory computer readable medium electronically coupled to the processor, wherein the non-transitory computer readable medium comprises instructions stored thereon activated, after the receipt of the signal, to perform the steps of disconnecting the electrical connection between the first electrical connection and the first output. 26. The device of claim 21, 24, or 25, wherein the signal comprises a measurement for electrical magnitude and phase. 27. A method of non-thermal treatment of a lesion in a subject in need thereof comprising: identifying the location, size and volume of the lesion in the subject; determining an electric field useful in the treatment of the identified lesion; and applying the electric field to the lesion for an appropriate time and appropriate voltage using a device comprising: a generator comprising a first output connector and a second output connector; an applicator comprising a first electrode and a second electrode; a dispersive pad; and a switch in electrical connection with the generator, wherein the switch is capable of changing between a first position to apply an electric field of a first mode through the applicator and a second position to apply an electric field of a second mode through the applicator. 28. The method of claim 27, wherein the first mode is a monopolar mode that generates aa electric field between the applicator and the dispersive pad. 29. The method of claim 28, wherein the dispersive pad is in contact with the subject in need thereof. 30. A method of non-thermal treatment of a lesion in a subject in need thereof comprising: determining an electric field useful in the treatment of the lesion to form a determined electric field; applying the determined electric field of a first mode or an electric field of a second mode to the lesion with a device comprising: a generator comprising a first output and a second output; an applicator comprising a first electrode and a second electrode; a dispersive pad; and Arnold & Porter Ref. P35370WO00 a switch in electrical connection with the generator, wherein the switch is capable of changing between a first position to apply the electric field of the first mode through the applicator and a second position to apply the electric field of the second mode through the applicator, wherein the electric field is within a treatment area in the subject in need thereof. 31. The method of claim 30, wherein the lesion is an endometrial lesion, a malignant skin lesion, a non-malignant skin lesion, a bladder lesion, a ductal breast cancer lesion, a lesion of the gastrointestinal system, or a combination thereof. 32. The method of claim 30, wherein the electric field does not cause damage outside of the treatment area. 33. The method of claim 30, wherein the applying the determined electric field comprises applying the determined electric field in cycles separated by one or more rest periods to reduce thermal damage. 34. The method of claim 30, wherein the electric field comprises 500 to 2000 volts (V). 35. The method of claim 30, wherein the first mode is a monopolar mode that generates an electric field between the applicator and the dispersive pad. 36. The method of claim 35, wherein the dispersive pad is configured to be in contact with the patient when the electric field is applied. 37. The method of claim 30, wherein the second mode is a bipolar mode that generates an electric field between the first electrode and the second electrode of the applicator. 38. A method of causing controlled cell necrosis in an endometrium of a subject in need thereof comprising: determining an electric field useful in the treatment of an endometrial lesion to form a determined electric field; applying the determined electric field of a first mode or an electric field of a second mode to the endometrial lesion with a device comprising: a generator comprising a first output and a second output, an applicator comprising a first electrode and a second electrode; a dispersive pad; and Arnold & Porter Ref. P35370WO00 a switch in electrical connection with the generator, wherein the switch is capable of changing between a first position to apply the electric field of the first mode through the applicator and a second position to apply the electric field of the second mode through the applicator, wherein the electric field of the first mode or the electric field of the second mode is within a treatment area of the subject in need thereof and causes cell necrosis within the treatment area. 39. The method of claim 38, wherein the first mode is a monopolar mode that generates an electric field between the applicator and the dispersive pad. 40. The method of claim 39, wherein the dispersive pad is configured to be in contact with the patient when the electric field is applied. 41. The method of claim 38, wherein the second mode is a bipolar mode that generates an electric field between the first electrode and the second electrode of the applicator. 42. The device of claim 1, wherein the applicator is an articulating tip catheter comprising a shaft comprising: a distal end comprising a first electrode 210 comprises a first stainless steel rod 212, a distal end 213, and a proximal end 214; a second electrode 220 comprises a second stainless steel tube 222, a distal end 223, and a proximal end 224, and an insulating layer; a flexible section; and a proximal end comprising mechanical levers connected to the flexible section with cables, wherein the flexible section is situated between the proximal end and the distal end, and wherein the first electrode 210 and the second electrode 220 are concentric, and the proximal end 214 of the first electrode and the distal end 223 of the second electrode are separated by a distance in length 230w. 43. An articulating tip applicator comprising: a shaft comprising: a distal end comprising a first electrode 210 comprises a first stainless steel rod 212, a distal end 213, and a proximal end 214; a second electrode 220 comprises a second stainless steel tube 222, a distal end 223, and a proximal end 224, and an insulating layer; Arnold & Porter Ref. P35370WO00 a flexible section; and a proximal end comprising mechanical levers connected to the flexible section with cables, wherein the flexible section is situated between the proximal end and the distal end, and wherein the first electrode 210 and the second electrode 220 are concentric, and the proximal end 214 of the first electrode and the distal end 223 of the second electrode are separated by a distance in length 230w. 44. The articulating tip applicator of claim 43, wherein the proximal end of the shaft further comprises a motor. 45. A device for non-thermal treatment of one or more lesions in a subject in need thereof comprising: a generator 110 comprising a first output connector and a second output connector; an applicator 1200 comprising a first electrode 1210 and a second electrode 1220; a dispersive pad 300; and a switch 410 in electrical connection with the generator, wherein the switch is capable of changing between a first position to apply an electric field of a first mode through the applicator and a second position to apply an electric field of a second mode through the applicator. 46. The device of claim 45, wherein the switch 410 is selected from the group consisting of a mechanical switch, an electronic switch, a latch switch, a momentary switch, a double throw switch, a single pole switch, or a changeover switch. 47. The device of claim 46, wherein the first mode is a monopolar mode that generates an electric field between the applicator and the dispersive pad. 48. The device of claim 46, wherein the switch 410 in the first position is configured to provide an electrical connection between the generator 110 through the first output connector with the first electrode 1210 and the second electrode 1220 of the applicator and provide an electrical connection between the generator through the second output connector with the dispersive pad 300. 49. The device of claim 46, wherein the dispersive pad 300 is configured to be in contact with the patient when the electric field is applied. Arnold & Porter Ref. P35370WO00 50. A device for non-thermal treatment of one or more lesions in a subject in need thereof comprising an applicator 1200 comprising: a proximal end 1201; a shaft 1203; applicator tip 1205 comprising one or more first electrode 1210 and one or more second electrode 1220; and distal end 1209. 51. The device of claim 50, further comprising two or more of the first electrode 1210 and two or more of the second electrode 1220. 52. The device of claim 51, wherein the two or more of the first electrode 1210 and two or more of the second electrode 1220 are separated by insulating layer 1290. 53. A method of non-thermal treatment of a lesion in a subject in need thereof comprising: determining an electric field useful in the treatment of the lesion to form a determined electric field; applying the determined electric field of a first mode or an electric field of a second mode to the lesion with a device comprising: a generator comprising a first output and a second output; an clamping probe comprising a first probe and a second probe; and a switch in electrical connection with the generator, wherein the switch is capable of changing between a first position to apply the electric field of the first mode through the applicator and a second position to apply the electric field of the second mode through the applicator, wherein the electric field is within a treatment area in the subject in need thereof. 54. The method of claim 53, wherein the device further comprises a dispersive pad. 55. The method of claim 53, wherein the clamping probe is clamping probe 201, 301, or 401. | 2023 | J3 MEDICAL | Medical and Healthcare | NA | https://j3medical.com/ | USA | The patent focuses on a non-thermal treatment device for endometriosis using electric-field treatment to remove endometrial lesions in a subject. It addresses the limited treatment options for endometriosis, where surgery and hormonal therapy have side effects. Non-thermal electric-field therapy aims to remove endometrial lesions with reduced collateral damage. | הפטנט מתמקד במכשיר טיפול לא תרמי לאנדומטריוזיס באמצעות טיפול בשדה חשמלי להסרת נגעים אנדומטריאליים בנבדק. הוא מתמודד עם אפשרויות הטיפול המוגבלות לאנדומטריוזיס, שבהן לניתוח ולטיפול הורמונלי יש תופעות לוואי. טיפול לא תרמי בשדה חשמלי שואף להסיר נגעים אנדומטריאליים עם נזק נלווה מופחת. | 1 | [29 empty] | 1 |  | J3 MEDICAL LLC | J3 MEDICAL LLC | J3 MEDICAL LLC | J3 MEDICAL LLC | J3 MEDICAL LLC | 2024-09-27 | 2024 | 2025-04-03 | 2025 | 2023-09-28 | 2023 | Application | Utility Invention | inactive | WO2024US48916 | 320414599 | WO2025072715A3 | WO
476 | WO2025222085A1 | MEDICAL DEVICE GUIDE SHEATHS WITH RETURN ELECTRODES FOR BIPOLAR THERAPEUTIC ELECTROPORATION | A medical device includes a guide sheath and one or more return electrodes disposed on an outer surface of the guide sheath. The one or more return electrodes are configured, with one or more radiofrequency ablation electrodes disposed on a distal lip of an elongate member inserted through the guide sheath, to perform bipolar radiofrequency ablation. The one or more return electrodes may comprise a first return electrode disposed on a first location of the outer surface of the guide sheath and at least a second return electrode disposed on a second location of the outer surface of the guide sheath. | 1. CLAIMSWhat is claimed is: 1. A medical device, comprising: a guide sheath; and one or more return electrodes disposed on an outer surface of the guide sheath, the one or more return electrodes being configured, with one or more radiofrequency ablation electrodes disposed on a distal tip of an elongate member inserted through the guide sheath, to perform bipolar radiofrequency ablation. 2. The medical device of claim 1, wherein the one or more return electrodes comprises a first return electrode disposed on a first location of the outer surface of the guide sheath and at least a second return electrode disposed on a second location of the outer surface of the guide sheath. 3. The medical device of claim 2, wherein the guide sheath is configured such that a distal tip thereof is deployed in a first lumen and a portion of a length of the guide sheath is deployed in a second lumen different than the first lumen. 4. The medical device of claim 3, wherein the first location is proximate the distal tip of the guide sheath deployed in the first lumen and the second location is on the portion of the length of the guide sheath deployed in the second lumen. 5. The medical device of claim 4, wherein the one or more return electrodes further comprises a third return electrode disposed on a third location of the outer surface of the guide sheath, the third location being configured to contact at least a portion of a wall of the first lumen and at least a portion of a wall of the second lumen. 6. The medical device of claim 3, wherein the first lumen comprises a vessel branching from a celiac trunk, and the second lumen comprises the celiac trunk. 7. The medical device of claim 2, wherein the first location is disposed on a first portion of a circumference of the outer surface of the guide sheath along a first length of the guide sheath, and wherein the second location is disposed on a second portion of the circumference of the outer surface of the guide sheath along a second length of the guide sheath. 8. The medical device of claim 2, wherein the first location is disposed on a first portion of a circumference of the outer surface of the guide sheath along a first length of the guide sheath, and wherein the second location is disposed on the first portion of the circumference of the outer surface of the guide sheath along a second length of the guide sheath. 9. The medical device of claim 1, wherein the one or more return electrodes are further configured, with the one or more radiofrequency ablation electrodes, to monitor a tissue impedance of target tissues in a vicinity of a wall of a lumen in which the guide sheath is deployed. 10. The medical device of claim 1, further comprising the elongate member inserted through the guide sheath. 11. The medical device of claim 10, wherein the elongate member comprises a delivery catheter. 12. The medical device of claim 10, further comprising: one or more expandable elements positioned on the elongate member, the one or more expandable elements being configured to expand radially outward from the elongate member, the one or more expandable elements having inner surfaces facing the elongate member and outer surfaces opposite the inner surfaces, at least a portion of the one or more expandable elements providing a radiofrequency ablation electrode; and a set of temperature sensors attached to the inner surfaces of at least one of the one or more expandable elements. 13. The medical device of claim 10, further comprising: one or more expandable elements positioned on the elongate member, the one or more expandable elements being configured to expand radially outward from the elongate member, the one or more expandable elements having inner surfaces facing the elongate member and outer surfaces opposite the inner surfaces, at least a portion of the one or more expandable elements providing a radiofrequency ablation electrode; and a flexible circuit surrounding at least a portion of the outer surfaces of the one or more expandable elements, the flexible circuit comprising a set of temperature sensors. 14. The medical device of claim 10, wherein the elongate member comprises a distal tip section, a first inflection section, a first elongate section, and a second inflection section, and at least a second elongate section, the first inflection section defining a pre-bent tip angle for the distal tip section of the elongate member, the second inflection section comprising an articulable bend of the first elongate section relative to the second elongate section, and further comprising one or more radiofrequency ablation electrodes disposed at the distal tip section of the elongate member. 15. A method, comprising: delivering, via a guide sheath of a medical device, an elongate member to a target site within a lumen; actuating at least a portion of the elongate member to bring at least a portion of a distal tip of the elongate member into contact with a wall of the lumen; and performing bipolar radiofrequency ablation between one or more radiofrequency ablation electrodes disposed on a distal tip of the elongate member and one or more return electrodes disposed on an outer surface of the guide sheath to able target tissues in a vicinity of the wall of the lumen. 16. The method of claim 15, further comprising monitoring an impedance between the one or more radiofrequency ablation electrodes disposed on the distal tip of the elongate member and the one or more return electrodes disposed on the outer surface of the guide sheath. 17. The method of claim 16, further comprising controlling a current applied to the one or more radiofrequency ablation electrodes disposed on the distal tip of the elongate member based at least in part on the monitored impedance. 18. The method of claim 17, further comprising monitoring a temperature of target tissues in the vicinity of the wall of the lumen utilizing one or more temperature sensors disposed proximate the one or more radiofrequency ablation electrodes, wherein the current applied to the one or more radiofrequency ablation electrodes disposed on the distal tip of the elongate member is further based at least in part on the monitored temperature. 19. The method of claim 15, wherein the one or more return electrodes comprises a first return electrode disposed on a first location of the outer surface of the guide sheath and at least a second return electrode disposed on a second location of the outer surface of the guide sheath. 20. The method of claim 19, further comprising positioning the guide sheath such that a distal tip of the guide sheath is deployed in the lumen and a portion of a length of the guide sheath is deployed in another lumen. 21. The method of claim 20, wherein the first location is proximate the distal tip of the guide sheath positioned in the lumen and the second location is on the portion of the length of the guide sheath positioned in the other lumen. 22. The method of claim 21, wherein the one or more return electrodes further comprises a third return electrode disposed on a third location of the outer surface of the guide sheath, wherein the third location is positioned to contact at least a portion of the wall of the lumen and at least a portion of a wall of the other lumen. 23. The method of claim 20, wherein the lumen comprises a vessel branching from a celiac trunk, and wherein the other lumen comprises the celiac trunk. 24. The method of claim 19, wherein the first location is disposed on a first portion of a circumference of the outer surface of the guide sheath along a first length of the guide sheath, and wherein the second location is disposed on a second portion of the circumference of the outer surface of the guide sheath along a second length of the guide sheath. 25. The method of claim 19, wherein the first location is disposed on a first portion of a circumference of the outer surface of the guide sheath along a first length of the guide sheath, and wherein the second location is disposed on the first portion of the circumference of the outer surface of the guide sheath along a second length of the guide sheath. 26. The method of claim 15, wherein actuating at least a portion of the elongate member to bring at least a portion of the distal tip of the elongate member into contact with the wall of the lumen comprises expanding one or more expandable elements positioned on the elongate member radially outward from the elongate member such that outer surfaces of the one or more expandable elements contact the wall of the lumen, at least a portion of the one or more expandable elements providing at least one of the one or more radiofrequency ablation electrodes. 27. The method of claim 15, wherein the elongate member comprises a distal tip section, a first inflection section, a first elongate section, and a second inflection section, and at least a second elongate section, the first inflection section defining a pre-bent tip angle for the distal tip section of the elongate member, the second inflection section comprising an articulable bend of the first elongate section relative to the second elongate section, and wherein actuating at least a portion of the elongate member to bring at least a portion of the distal tip of the elongate member into contact with the wall of the lumen comprises actuating the second inflection section to provide the articulable bend of the first elongate section relative to the second elongate section to bring the distal tip section into contact with a wall of the lumen. | 2024 | AUTONOMIX MEDICAL | Medical and Healthcare | Small-size company | https://www.autonomix.com/ | USA | The patent focuses on a medical guide sheath with one or more return electrodes on its outer surface, paired with RF ablation electrodes on a distal lip of an elongate member inserted through the sheath. It addresses the workflow burden of needing separate return electrodes during bipolar therapeutic electroporation. Sheath-mounted return electrodes simplify bipolar RF ablation in one device. | הפטנט מתמקד בנדן הנחיה רפואי עם אלקטרודת חזרה אחת או יותר על משטחו החיצוני, מותאם עם אלקטרודות אבלציית RF בשפה דיסטלית של אלמנט מאורך המוחדר דרך הנדן. הוא מתמודד עם נטל זרימת העבודה של צורך באלקטרודות חזרה נפרדות במהלך אלקטרופורציה טיפולית דו-קוטבית. אלקטרודות חזרה מותקנות-נדן מפשטות אבלציית RF דו-קוטבית במכשיר אחד. | 1 |  |  |  | 1 | [5 empty] | 1 | [19 empty] | 1 |  | AUTONOMIX MEDICAL INC | AUTONOMIX MEDICAL INC | AUTONOMIX MEDICAL INC | AUTONOMIX MEDICAL INC | AUTONOMIX MEDICAL INC | 2025-04-18 | 2025 | 2025-10-23 | 2025 | 2024-04-19 | 2024 | Application | Utility Invention | active | WO2025US25307 | 324042680 | [NA] | WO
477 | US20220287767A1 | Targeted Lung Denervation with Directionally-Adjustable Perfusion | A lung denervation method comprises advancing a catheter along the airway; deploying an open loop in contact with the epithelium; simultaneously delivering radio frequency energy from a plurality of discrete spaced-apart locations along the loop to target regions according to a set of ablation parameters sufficient to heat and interrupt nerve functionality; and forming a liquid film between the loop and the epithelium for minimizing collateral damage. The method serves to destroy motor axons of the peripheral bronchial nerve, blocks parasympathetic transmission in the pulmonary and reduces acetylcholine release, reducing airway smooth muscle tension and mucus production. Related systems are described. | 1. A method for lung denervation along an airway in the lung comprising: advancing a catheter along the airway; deploying an open loop in contact with the epithelium of the airway; simultaneously delivering radio frequency energy from a plurality of discrete spaced-apart locations along the loop to target regions along the epithelium according to a set of ablation parameters sufficient to heat and interrupt the bronchial nerve functionality; and forming a liquid film between the loop and the epithelium for protecting damage to the epithelium. 2. The method of claim 1, wherein forming is performed by flowing a cooling agent from the discrete spaced-apart locations onto the epithelium. 3. The method of claim 2, wherein the loop comprises an electrode at each discrete spaced-apart location for delivering radio frequency energy. 4. The method of claim 3, wherein each electrode comprises an array of egress ports, through which the cooling agent is ejected. 5. The method of claim 1, wherein non-targeted regions of the epithelium between the electrodes are protected by the cooling film. 6. The method of claim 1, further comprising retracting the loop, moving the catheter to a new location, deploying the open loop at the new location, and repeating the delivering and forming steps. 7. The method of claim 6, wherein the moving comprises advancing and/or rotating. 8. The method of claim 1, wherein the ablation parameters comprise a single electrode output energy of 1000-1500 J, and a power limit not to exceed 20 W. 9. The method of claim 1, wherein the ablation parameters comprise a single electrode output energy of 1080 J˜1360 J and power limit of 12 W˜16 W. 10. The method of claim 2, wherein the flowing is performed using iced saline. 11. The method of claim 2, wherein the flowing comprises adjusting the flowrate of the cooling agent from high to low. 12. The method of claim 2, further comprising monitoring each location, and independently adjusting the flowrate of the cooling agent to each location based on the monitoring. 13. The method of claim 12, wherein the monitoring comprises monitoring temperature. 14. The method of claim 12, further comprising displaying ablation progress based on the monitoring. 15. The method of claim 2, further comprising independently adjusting the flowrate of the cooling agent to each location such that the coolant is controllably directed to one or more desired areas of the airway and excludes one or more undesired areas. 16. The method of claim 15, wherein each area is monitored for the presence of the coolant, and the controlling is based on the monitoring. 17. An electrosurgical method of treating chronic bronchitis comprising: destroying motor axons of a peripheral bronchial nerve, blocking parasympathetic transmission in the pulmonary nerve and reducing acetylcholine release, thereby reducing mucus production, thereby improving airway obstruction; and simultaneously, during the destroying step, ejecting a cooling agent to a plurality of regions along the inner wall of the airway according to a plurality of customized flowrates based on temperature of each region. 18. The method of claim 17, wherein the destroying is performed by applying radiofrequency energy to discrete circumferential locations. 19. The method of claim 18, further comprising displaying ablation progress based on monitoring the temperature of each region. 20. The method of claim 19, wherein the ejecting step is performed at a sufficient rate and geography to protect epithelial tissue yet allow heat penetration to the bronchial nerve. | 2020 | BRONCUS MEDICAL | Medical and Healthcare | Small-size company | http://www.broncus.com/ | USA | The patent focuses on targeted lung denervation using an open-loop catheter that delivers RF from multiple discrete locations along the loop while forming a protective liquid film between the loop and epithelium. It addresses the risk of damaging airway epithelium during bronchial nerve ablation. The liquid film protects the surface tissue while RF reaches the deeper bronchial nerves. | הפטנט מתמקד בדה-אינרבציה ריאתית ממוקדת באמצעות קטטר בלולאה פתוחה המספק RF ממספר מיקומים נפרדים לאורך הלולאה תוך יצירת סרט נוזלי מגן בין הלולאה לאפיתל. הוא מתמודד עם הסיכון לפגיעה באפיתל דרכי הנשימה במהלך אבלציית עצבי הסימפונות. סרט הנוזל מגן על רקמת המשטח בעוד RF מגיע לעצבי הסימפונות העמוקים יותר. |  |  |  |  | 1 | [5 empty] | 1 | [15 empty] | 1 | [5 empty] | BRONCUS MEDICAL INC | BRONCUS MEDICAL INC | BRONCUS MEDICAL INC. | BRONCUS MEDICAL INC | BRONCUS MEDICAL INC | 2022-04-12 | 2022 | 2022-09-15 | 2022 | 2020-09-17 | 2020 | Application | Utility Invention | inactive | US202217658947 | 292203583 | CN110960314A¦CN110960314B¦CN111012481A¦CN111012481B¦CN112402008A¦CN112402008B¦CN112641501A¦CN112641501B¦CN112641504A¦CN112641504B¦CN112712884A¦CN112712884B¦CN112734909A¦CN112734909B¦CN112741682A¦CN112741682B¦CN112741683A¦CN112773498A¦CN112773498B¦CN112807071A¦CN112807071B¦CN112842514A¦CN112842514B¦CN112865774A¦CN112865774B¦CN114942710A¦EP4085858A1¦EP4085858A4¦EP4085861A1¦EP4085861A4¦EP4109755A1¦EP4109755A4¦EP4113532A1¦EP4113532A4¦EP4115827A1¦EP4115827A4¦EP4119079A1¦EP4119079A4¦EP4272667A1¦EP4272667A4¦EP4272668A1¦EP4272668A4¦EP4272669A1¦EP4272669A4¦EP4272670A1¦EP4272670A4¦EP4273810A1¦EP4273810A4¦JP2023509659A¦JP2023509660A¦JP7416959B2¦JP7490063B2¦KR102776055B1¦KR102782596B1¦KR20220156519A¦KR20220156520A¦US12395867B2¦US12433661B2¦US12484960B2¦US20220241001A1¦US20220323149A1¦US20220331007A1¦US20220409272A1¦US20230016496A1¦US20230016753A1¦US20230017640A1¦US20230021076A1¦US20230022310A1¦US20230025054A1¦US20230029056A1¦US20230029103A1¦US20230144706A1¦WO2021135460A1¦WO2021135461A1¦WO2022057949A1¦WO2022121017A1¦WO2022141685A1¦WO2022141686A1¦WO2022141687A1¦WO2022141688A1¦WO2022141689A1¦WO2022141690A1¦WO2022141691A1¦WO2022141769A1 | US
478 | JP7774074B2 | Tumor Ablation Instruments and Techniques | The present invention relates to determining the boundaries of a tumor. A method for ablating a bordered tumor (15) in tissue of a subject includes advancing a distal region (200) of a tumor ablation device (102) into the tissue, the distal region having electrodes (2, 4) attached thereto. The bioimpedance of the tissue is then sensed using the distal region of the device. In response to the sensed bioimpedance, the electrodes are positioned within the border of the tumor. While the electrodes remain within the border, an electroporation pulse is applied to the tumor using the electrodes. Other embodiments are also described. [Selection diagram] Figure 6C | 1.1. A device for use on a tumor within a subject's body, the device comprising: a tumor ablation device having a distal region, the tumor ablation device comprising: an inner shaft; and an outer shaft coaxially disposed with the inner shaft, the inner shaft comprising a first bioimpedance sensing electrode and a first electroporation electrode mounted at a fixed location proximal to the first bioimpedance sensing electrode; the outer shaft comprising a second bioimpedance sensing electrode and a second electroporation electrode mounted at a fixed location distal to the second bioimpedance sensing electrode; the inner shaft and the outer shaft being telescopically slidable relative to each other to vary a first axial distance between the first electroporation electrode and the second electroporation electrode. 2.The device of claim 1 , wherein the tumor is a tumor in the subject's lung and the distal region is transbronchially deliverable to the lung. 3.2. The device of claim 1, wherein the tumor has a boundary, and the distal region is advanceable to a position within the lung by telescopically sliding the inner shaft and the outer shaft relative to each other, where the first electroporation electrode and the second electroporation electrode are both positioned within the boundary of the tumor, and the first bioimpedance sensing electrode and the second bioimpedance sensing electrode are both positioned outside the boundary of the tumor, on opposite sides of the tumor. 4.10. The device of claim 1, configured to sense bioimpedance in the tumor by sensing bioimpedance between the first electroporation electrode and the first bioimpedance sensing electrode, and bioimpedance between the second electroporation electrode and the second bioimpedance sensing electrode. 5.10. The device of claim 1, configured to sense bioimpedance at the tumor by sensing bioimpedance between the first bioimpedance sensing electrode and the second bioimpedance sensing electrode. 6.The device of claim 1 , wherein the inner shaft is advanceable from a distal end of the outer shaft. 7.10. The device of claim 1, further comprising a control unit including a generator electrically connectable to both the first electroporation electrode and the second electroporation electrode and adapted to drive electroporation pulses therebetween. 8.The device of claim 1 , wherein the outer shaft is formed from an electrical insulator. 9.9. The device of claim 8, wherein the effective length of the first electroporation electrode can be adjusted by sliding the outer shaft over the first electroporation electrode. 10.10. The device of claim 1, comprising a control unit configured to receive a signal from the first bioimpedance sensing electrode and responsively provide an output indicative of the bioimpedance of tissue adjacent to the first electroporation electrode, and to receive a signal from the second bioimpedance sensing electrode and responsively provide an output indicative of the bioimpedance of tissue adjacent to the second electroporation electrode. 11.11. The apparatus of claim 10, wherein the control unit is configured to output information indicative of a position of each of the first and second electroporation electrodes relative to the tumor in response to a signal. 12.11. The device of claim 10, wherein the control unit is configured to: in response to the signal, determine whether a state of the distal region is such that both the first electroporation electrode and the second electroporation electrode are disposed within a tumor; and, in response to determining such a state of the distal region, provide an indication that both the first electroporation electrode and the second electroporation electrode are disposed within a tumor. 13.11. The apparatus of claim 10, wherein the control unit is configured to generate a signal by driving an electrical pulse between the first bioimpedance sensing electrode and the second bioimpedance sensing electrode. 14.11. The device of claim 10, wherein the control unit is configured to receive a signal by driving a first electrical pulse between the first electroporation electrode and the first bioimpedance sensing electrode and driving a second electrical pulse between the second electroporation electrode and the second bioimpedance sensing electrode. 15.11. The apparatus of claim 10, wherein the control unit is configured to output information indicative of positions of the first bioimpedance sensing electrode and the second bioimpedance sensing electrode relative to the tumor in response to a signal. 16.16. The device of claim 15, wherein the control unit is configured to, in response to the signal, distinguish a state of the distal region in which both the first bioimpedance sensing electrode and the second bioimpedance sensing electrode are located outside a tumor from one or more other states of the distal region in which at least one of the first bioimpedance sensing electrode and the second bioimpedance sensing electrode is located inside a tumor, and, in response to the distinction, provide an indication that the distal region is in that state. 17.16. The device of claim 15, wherein the control unit is adapted to, in response to the signal, distinguish between (i) a condition of the distal region in which both the first bioimpedance sensing electrode and the second bioimpedance sensing electrode are located outside a tumor and (ii) one or more other conditions of the distal region in which at least one of the first bioimpedance sensing electrode and the second bioimpedance sensing electrode is located inside a tumor, and, in response to the distinction, provide an indication of whether the first electroporation electrode and the second electroporation electrode are positioned within the tumor. 18.10. The device of claim 1, wherein the tumor is located within a lung of the subject, and the device is configured to determine a boundary of the tumor by sensing bioimpedance via at least one of the first bioimpedance sensing electrode and the second bioimpedance sensing electrode. 19.20. The apparatus of claim 18, comprising a bronchoscope, the bronchoscope being transbronchially advanceable to the lung, and the distal region being deliverable to the tumor via the bronchoscope. | 2021 | Snipe Medical | Medical and Healthcare | Start Up | https://snipemedical.com/ | Israel | The patent focuses on a tumor ablation device with electrodes at the distal region that senses bioimpedance to determine tumor boundaries. It addresses the difficulty of accurately delineating tumor boundaries during ablation, where blind ablation may leave residual disease. Bioimpedance-based boundary detection guides more complete tumor ablation. | הפטנט מתמקד במכשיר אבלציית גידולים עם אלקטרודות באזור הדיסטלי החש ביו-אימפדנס לקביעת גבולות הגידול. הוא מתמודד עם הקושי בהגדרה מדויקת של גבולות גידול במהלך אבלציה, שבה אבלציה עיוורת עלולה להשאיר מחלה שיורית. גילוי גבולות מבוסס ביו-אימפדנס מנחה אבלציית גידול שלמה יותר. | 1 | [10 empty] | 1 | [16 empty] | 1 |  |  |  | [NA] | [NA] | [NA] | [NA] | SNIPE MEDICAL LTD | 2022-05-10 | 2022 | 2025-11-20 | 2025 | 2021-05-22 | 2021 | Grant | Utility Invention | active | JP20230561062 | 319706079 | AU2022282840A1¦CA3216185A1¦CN117241751A¦EP4340759A1¦EP4340759A4¦JP2024520990A¦JP2026027412A¦US20240206937A1¦WO2022248957A1 | JP
479 | JP2025516498A | DETECTING NERVE ACTIVITY | An interventional device (110) for detecting neural activity is provided. The interventional device includes a magnetic sensor (120) coupled to an insertable portion of the interventional device. The magnetic sensor (120) is configured for generating a signal in response to a magnetic field generated by neural activity. The invention also provides a system (200). The system includes an interventional device (110) and a controller (140) configured to receive a signal generated by the magnetic sensor (120) and output a detection result indicative of neural activity in response to the received signal. | 1.1. An interventional device for detecting neural activity, the interventional device having a magnetic sensor coupled to an insertable portion of the interventional device, the magnetic sensor configured to generate a signal in response to a magnetic field generated by neural activity. 2.2. The interventional device of claim 1, wherein the magnetic sensor comprises an optically pumped magnetometer (OPM), the OPM having an optical cell containing an alkali metal in a liquid and/or gas phase, the OPM configured to generate the signal in response to a magnetic field generated by neural activity by measuring a magnetic field-dependent optical property of the alkali metal. 3.3. The interventional device of claim 2, wherein the optical cell further comprises a background gas, and the OPM is further configured to generate the signal in response to a magnetic field generated by neural activity by measuring a magnetic field-dependent optical property of the background gas. 4.4. The interventional device of claim 2 or claim 3, wherein the OPM further comprises a heater, the heater configured to provide heat to the optical cell. 5.5. A system having an interventional device according to any one of claims 1 to 4 and a controller, the controller configured to receive the signal generated by the magnetic sensor and output a detection result indicative of neural activity in response to the received signal. 6.The system of claim 5 when dependent on claim 4, wherein the controller is configured to selectively operate the OPM in a relatively low sensitivity mode in which a relatively low power level is supplied by the heater to the optical cell to generate a signal in response to a magnetic field generated by neural activity, and in a relatively high sensitivity mode in which a relatively high power level is supplied by the heater to the optical cell. 7.7. The system of claim 6, wherein the optical cell comprises an alkali metal and a background gas in a liquid and/or gas phase, and the OPM is configured to generate the signal in response to a magnetic field generated by neural activity by measuring magnetic field-dependent optical properties of the alkali metal and the background gas, and in the relatively low sensitivity mode, optical properties of the background gas are sensed, and in the relatively high sensitivity mode, optical properties of the alkali metal are sensed; or the optical cell comprises a first alkali metal in a liquid and/or gas phase and a second alkali metal in a liquid and/or gas phase, and the OPM is configured to generate the signal in response to a magnetic field generated by neural activity by measuring magnetic field-dependent optical properties of the first alkali metal and the second alkali metal, and in the relatively low sensitivity mode, optical properties of the first alkali metal are sensed, and in the relatively high sensitivity mode, optical properties of the second alkali metal are sensed. 8.The system of claim 5 when dependent on claim 4, wherein the interventional device further comprises a temperature sensor, the temperature sensor being in thermal contact with i) an outer surface of the interventional device or ii) the optical cell, and the controller is configured to operate the heater in response to a temperature measured by the temperature sensor. 9.The system of claim 5 when dependent on claim 2, wherein the OPM has a light source configured to provide an optical pump beam, the OPM is configured to measure the magnetic field-dependent optical property of the alkali metal in response to excitation of the alkali metal in the optical cell by the optical pump beam, the controller is configured to modulate the intensity of the optical pump beam between a relatively low intensity and a relatively high intensity, and the OPM is configured to measure the magnetic field-dependent optical property of the alkali metal by determining the transmittance of the optical pump beam through the optical cell while the optical pump beam excites the alkali metal at the relatively high intensity. 10.6. The system of claim 5, wherein the magnetic sensor comprises a plurality of magnetic sensor elements, each magnetic sensor element configured to detect a magnetic field across a different volume of the interventional device, and the controller is further configured to measure a signal-to-noise ratio of the signals generated by the magnetic sensor elements and to generate the detection result using a signal from one or more of the magnetic sensor elements having a highest signal-to-noise ratio. 11.11. The system of claim 10, wherein the sensor elements are distributed around an axis of the interventional device, the sensor elements generating signals in response to magnetic fields generated at different orientations around the axis, and the controller is further configured to identify an orientation around the axis of the interventional device in which the measured signal has a maximum signal-to-noise ratio. 12.6. The system of claim 5, wherein the controller is further configured to output the detection result indicative of neural activity in response to the received signal by identifying a signature within the received signal and outputting a measure of renal sympathetic overdrive based on the identified signature. 13.6. The system of claim 5, wherein the system further comprises one or more coils configured to generate a magnetic field to compensate for a background magnetic field detected by the magnetic sensor, the interventional device further comprises a motion sensor, the motion sensor being mechanically coupled to the interventional device to detect movement of the magnetic sensor, and the controller is further configured to: measure the background magnetic field based on a signal generated by the magnetic sensor in the absence of a magnetic field generated by neural activity; apply a signal to the one or more coils to compensate for the background magnetic field during measurement of the magnetic field generated by neural activity; and, in response to detection of movement of the interventional device by the motion sensor, repeat the measurement of the background magnetic field and the corresponding application of a signal to the one or more coils. 14.14. The system of claim 13, wherein the signal is generated in response to a magnetic field generated by neural activity of the subject, the background magnetic field being generated at least in part by cardiac activity of the subject, and the controller is further configured to receive an electrocardiogram ECG signal for the subject, and the controller is configured to apply a signal to the one or more coils to compensate for the background magnetic field during measurement of the magnetic field generated by the neural activity by applying a time-dependent signal to the one or more coils in synchronization with the received ECG signal to compensate for the cardiac activity. 15.The system of claim 13 or claim 14, wherein the controller is further configured to receive tracking data representing a position of the magnetic sensor within a sensing area, receive a 3D background magnetic field map representing a background magnetic field distribution within the sensing area, and estimate a magnitude of the background magnetic field distribution at the position of the magnetic sensor within the sensing area based on the received tracking data and the received 3D background magnetic field map, and the controller is configured to apply a signal to the one or more coils to compensate for the background magnetic field during measurement of a magnetic field generated by neural activity based at least in part on the estimated magnitude of the background magnetic field distribution at the position of the magnetic sensor. | 2022 | Philips | Medical and Healthcare | Large-Size Company | https://www.philips.com | Netherlands | The patent focuses on an interventional device with a magnetic sensor coupled to its insertable portion that detects neural activity by sensing magnetic fields generated by the nerves. It addresses the challenge of detecting nerve activity during interventional procedures, where electrical sensing has noise and selectivity limitations. Magnetic sensing offers a different signal pathway for nerve detection. | הפטנט מתמקד במכשיר התערבותי עם חיישן מגנטי המחובר לחלקו הניתן להחדרה המזהה פעילות עצבית באמצעות חישת שדות מגנטיים שנוצרים על ידי העצבים. הוא מתמודד עם האתגר בזיהוי פעילות עצבית במהלך פרוצדורות התערבותיות, שבה לחישה חשמלית יש מגבלות רעש וסלקטיביות. חישה מגנטית מציעה נתיב אות שונה לזיהוי עצבים. | [7 empty] | 1 | [17 empty] | 1 | [6 empty] | [NA] | [NA] | [NA] | [NA] | Philips | 2023-05-04 | 2023 | 2025-05-30 | 2025 | 2022-05-12 | 2022 | Application | Utility Invention | active | JP20240564701 | 324689404 | CN119300760A¦EP4285821A1¦EP4522027A1¦US20250359795A1¦WO2023217617A1 | JP
480 | JP2025527304A | Lung treatment | The invention relates to methods and system for locally treating a wound (for example, ARDS) of a lung tissue, or mediastinum condition, comprising choosing a target in said tissue requiring treatment; insertion trans-bronchial catheter with mini-ultrasonic therapeutic transducers as close as possible to the target; calibrating contact, calibrating position and irradiating said target with therapeutic ultrasonic waves. Optionally, also providing treatment using phonophoresis and iontophoresis for local administration of drugs into the wound. | 1.1. A catheter system for treating a wound, the catheter system comprising: a. at least one catheter having a proximal end and a distal end, the distal end comprising: i. one or more ultrasound transmitters; ii. an inflatable balloon covering the one or more ultrasound transmitters, the inflatable balloon comprising a first electrode located proximally of the inflatable balloon and a second electrode located distally of the inflatable balloon; iii. a plurality of first type openings located on a surface of the balloon, the positions of the first type openings coinciding with the positions of the one or more ultrasound transmitters on the catheter; and iv. a plurality of second type openings located on a surface of the balloon, the positions of the second type openings coinciding with the positions of the first electrodes. 2.10. The catheter system of claim 1, further comprising at least one liquid configured to inflate the inflatable balloon. 3.The catheter system of claim 2 , wherein the at least one liquid includes at least one drug. 4.The catheter system of claim 2 , wherein the at least one liquid is saline. 5.The catheter system according to claim 1 , further comprising one or more controllers disposed at the proximal end of the catheter and configured to control movement of the distal end of the catheter. 6.6. The catheter system according to claim 1, further comprising a video camera located at the most distal end of the catheter. 7.The catheter system according to claim 1 , wherein the at least one catheter is configured to perform phonophoresis therapy. 8.7. The catheter system according to claim 1, wherein the at least one catheter is configured to perform iontophoresis therapy. 9.The method further includes an external unit to which the at least one catheter is connected, the external unit comprising: a. a video unit having a navigation system configured to operate the at least one catheter and provide visualization to a user from the video camera; b. an electric field unit configured to generate an interference field; and c. an ultrasonic wave unit configured to deliver ultrasound waves of about 0.5 MHz to about 3 MHz at a frequency of 2 watts/cm.29. The catheter system of claim 1, further comprising: an ultrasound unit configured to generate ultrasound waves at a power of up to 100 W; d. a calibration unit configured to specify attenuation coefficients at different positions and further configured to optimally locate the one or more ultrasound transmitters; e. a pressure unit configured to maintain and monitor pressure in the balloon; f. a phonophoresis drug unit configured to monitor and administer the at least one drug during phonophoresis treatment; and g. an iontophoresis drug unit configured to monitor and administer the at least one drug during iontophoresis treatment. 10.10. The catheter system of claim 9, wherein the electric field unit is configured to supply current to a first catheter at a first frequency and to supply current to a second catheter at a second frequency. 11.11. The catheter system according to claim 10, wherein the first frequency is about 4000 Hz, and the second frequency is about 4100 Hz to about 4500 Hz. 12.10. The catheter system of claim 9, wherein the electric field unit is configured to provide a current not exceeding 13 milliamps. 13.10. A method for treating a pulmonary tissue condition, a pulmonary embolism, and/or a mediastinal thromboembolic condition, comprising: a. selecting a target site in the tissue requiring treatment; b. advancing at least one catheter of claim 1 toward a location near the target site; c. performing contact calibration; d. performing position calibration; and e. performing ultrasound treatment. 14.14. The method of claim 13, wherein performing the contact calibration comprises: a. inflating the inflatable balloon; b. receiving signals from the first electrode and the second electrode; and c. stopping inflation of the balloon. 15.15. The method of claim 13 or 14, wherein performing the position calibration comprises: a. delivering a first amount of ultrasonic energy; b. determining an attenuation coefficient; and c. determining whether a position near the target site is an optimal position according to the result of the determination. 16.The method of claim 15, wherein the optimum position is a position where the damping coefficient results in a low damping coefficient. 17.17. The method of claim 16, wherein the low attenuation coefficient is about 0.18 dB/mHz*cm. 18.18. The method of claim 13, further comprising repositioning the catheter to a new location near the target site if the position calibration fails. 19.18. The method of claim 13, wherein performing the ultrasound treatment comprises irradiating the target area with ultrasound, the irradiation being characterized by time, intensity, and number of irradiation fields. 20.20. The method of claim 19, wherein the irradiating is characterized to obtain maximum therapeutic effect with minimum side effects and/or to avoid induction of cell death at the target site. 21. twenty one.20. The method of claim 19, wherein the irradiating comprises irradiating at an intensity between about 1 MHz and about 3 MHz. 22. twenty two.20. The method of claim 19, wherein the irradiating comprises irradiating for about 10 seconds to about 15 minutes. 23. twenty three.The method of claim 19 , wherein the illuminating comprises illuminating with a focused beam of light. 24. twenty four.24. The method of claim 13, further comprising monitoring the temperature of the target area where treatment is being performed. 25. twenty five.25. The method according to claim 13, further comprising preventing the temperature of the target site to be treated from rising above a predetermined temperature. 26.26. The method of claim 25, wherein the predetermined temperature is between about 40 degrees Celsius and about 45 degrees Celsius. 27.27. The method of any one of claims 13 to 26, further comprising actively controlling the temperature of the target area where treatment is performed. 28.28. The method of claim 27, wherein the controlling comprises varying the temperature of the target site where treatment is performed. 29.28. The method of claim 27, wherein the controlling comprises flowing a cold liquid to the target site when the target site reaches a predetermined temperature. 30.30. The method of any one of claims 13 to 29, further comprising visually monitoring the tissue at the target site. 31.20. The method of claim 19, wherein the irradiating is performed by two different catheters operated simultaneously. 32.The method according to any one of claims 13 to 31, further comprising administering phonophoresis treatment. 33.33. The method of claim 32, wherein administering the phonophoresis treatment comprises further inflating the inflatable balloon to a second inflated state, wherein the second inflated state opens the plurality of first type openings. 34.33. The method of claim 32, wherein administering the phonophoresis treatment comprises delivering at least one agent during administration of the phonophoresis treatment. 35.The method according to any one of claims 13 to 34, further comprising performing iontophoresis treatment. 36.36. The method of claim 35, wherein administering the iontophoresis treatment comprises further inflating the inflatable balloon to a third inflation state, wherein the third inflation state opens the plurality of second type openings. 37.36. The method of claim 35, wherein performing the iontophoretic treatment comprises delivering at least one agent during the iontophoretic treatment. 38.38. The method of any one of claims 13 to 37, further comprising generating an interference current by actuating two separate catheters towards the location. 39.39. The method according to any one of claims 13 to 38, wherein the treatment is used to treat one or more of ARDS, pulmonary parenchymal disease, and mediastinal disease. 40.40. The method of any one of claims 13 to 39, wherein the advancing is performed via the esophagus and/or trachea. | 2022 | Skybeam Ltd | NA | NA | NA | NA | The patent focuses on a method for locally treating lung wounds (e.g., ARDS) or mediastinum conditions using trans-bronchial catheters with mini-ultrasonic therapeutic transducers, after target selection, contact calibration, and position calibration. It addresses the limited targeted treatment options for ARDS and similar lung-tissue injuries. Trans-bronchial focused ultrasound enables localized therapy with calibrated targeting. | הפטנט מתמקד בשיטה לטיפול מקומי בפצעי ריאה (למשל ARDS) או מצבי מדיאסטינום באמצעות קטטרים טרנס-ברונכיאליים עם מיקרו-מתמרי אולטרסאונד טיפוליים, לאחר בחירת מטרה, כיול מגע וכיול מיקום. הוא מתמודד עם אפשרויות הטיפול המוגבלות הממוקדות ל-ARDS ופציעות רקמת ריאה דומות. אולטרסאונד ממוקד טרנס-ברונכיאלי מאפשר טיפול מקומי עם מיקוד מכויל. |  |  |  |  | 1 | [21 empty] | 1 | [5 empty] | [NA] | [NA] | [NA] | [NA] | SKYBEAM LTD | 2023-08-04 | 2023 | 2025-08-20 | 2025 | 2022-08-05 | 2022 | Application | Utility Invention | active | JP20250506964 | 324994869 | CN119836272A¦EP4565165A1¦KR20250047780A¦US12427293B2¦US20250256083A1¦US20250381380A1¦WO2024028825A1 | JP

## DEFINITION & ASSUMPTIONS
BAYSLOPE
TAXONOMY DEFINITION
Sr. No. | Taxonomy Nodes | Taxonomy Definitions
1 | Sensing Modality | The category includes patents/published applications disclosing different sensing modalities.
1.1 | Electrical / Impedance | This category includes patents/published applications that measure electrical properties, such as voltage, current, and impedance, of tissue or the electrode-tissue interface to evaluate contact quality, tissue condition, or lesion formation. Examples include impedance changes during RF ablation, multi-frequency bioimpedance analysis, electrogram (EGM)-based assessment, and voltage-current monitoring in pulsed field ablation (PEF).
1.2 | Optical / Photonic | This category includes patents/published applications that use how light interacts with tissue, such as reflection, absorption, scattering, or emission, to assess tissue type, composition, or ablation status. Examples include OCT for tissue structure, fiber-optic and Raman spectroscopy, ICG fluorescence for perfusion, and plasma light emission in lithotripsy.
1.3 | Ultrasound / Acoustic | This category includes patents/published applications that use transmitted and received ultrasound waves to visualize tissue structure or measure acoustic properties before, during, or after ablation. Examples include IVUS for vessel wall imaging, A-mode echo for cartilage or bronchial measurements, acoustic pressure monitoring in balloon lithotripsy, and microbubble detection during HIFU.
1.4 | Photoacoustic/Thermoacoustic | This category includes patents/published applications that use photoacoustic or thermoacoustic methods for assessment.
1.5 | Thermal / Infrared | This category covers patents/published applications that use direct tissue temperature monitoring to evaluate thermal damage or regulate energy delivery. Examples include MRI thermometry for 3D temperature mapping during laser ablation, thermocouples or thermistors on RF catheter tips, and temperature-sensing wires in pulmonary artery ablation systems.
1.6 | Mechanical / Contact Force | This category includes patents/published applications that measure force, pressure, or deformation between a device and tissue to confirm contact or evaluate tissue properties. Examples include contact force sensors on ablation catheter tips, pressure sensors in cryoballoons for pulmonary vein occlusion, and elastography techniques for assessing tissue stiffness.
1.7 | Chemical / Biochemical | This category includes patents/published applications that detect chemical or molecular signals in tissue or nearby fluids to assess tissue composition or cell viability. Examples include mid-infrared spectroscopy to measure tissue hydration and electrochemical sensing of metabolites such as lactate or pH released during ablation.
1.8 | Magnetic / Electromagnetic | This category includes patents/published applications describing the sensing of magnetic or electromagnetic properties of tissue, mainly for detecting nerve activity or tracking devices. For example, an optically pumped magnetometer (OPM) on an intravascular catheter can detect extremely weak (Pico Tesla-level) magnetic fields generated by renal sympathetic nerve signals.
1.9 | Imaging Based | This category includes patents/published applications that describe how tissue or lesions are characterized by analyzing 2D or 3D medical images from different imaging methods (such as MRI, CT, ultrasound, fluoroscopy, or endoscopy). Examples include: MRI-based thermal imaging to create real-time temperature maps, AI-based lesion detection and segmentation from CT or MRI, Intracardiac ultrasound (ICE) image analysis to assess cardiac structures during ablation procedures.
1.1 | Others | This category includes patents/published applications that describe alternative sensing approaches not covered above, such as strain sensing, nuclear magnetic resonance relaxometry, and other novel or hybrid methods used for tissue characterization during ablation.
2 | Ablation Energy Modality | The category includes patents/published applications disclosing different ablation modalities.
2.1 | Radiofrequency (RF) Ablation | This category includes patents/published applications that involve the use of radio frequency-based tissue ablation.
2.2 | Pulsed Electric Field (PEF) / Electroporation | The category includes patents/published applications disclosing use of Pulsed Electric Field/ Electroporation type ablation of tissues.
2.3 | Microwave Ablation | The category includes patents/published applications disclosing use of microwave electromagnetic energy for tissue ablation.
2.4 | Laser Ablation | The category includes patents/published applications disclosing use of laser for tissue ablation.
2.5 | Cryoablation | This category covers patents/published applications that use cryoprobes or cryoballoons to destroy tissue by freezing it, causing cell damage through ice crystal formation and osmotic stress. Examples include cryoballoons for pulmonary vein isolation in atrial fibrillation and percutaneous cryoablation for renal or prostate tumors.
2.6 | Ultrasound Ablation (HIFU) | This category covers patents/published applications that use high-intensity ultrasound to precisely destroy tissue, either by heating (thermal ablation) or mechanical disruption—while sparing surrounding tissue. Examples include intraluminal HIFU catheters for bronchial nerve denervation (asthma/COPD) and histotripsy techniques for non-thermal tumor destruction.
2.7 | Phototherapy / Photothermal | This category includes patents/published applications that use light, typically non-coherent or broadband near-infrared, to generate controlled heating in tissue through natural light absorption. Examples include photothermal ablation of gastrointestinal mucosal tissue using vacuum-assisted light delivery, and photodynamic therapy using photosensitizers.
2.8 | Steam Ablation | The category includes patents/published applications disclosing use of water vapor (steam) for tissue ablation.
2.9 | Others | This category covers patents/published applications related to ablation energy types not listed above. Examples include non-thermal plasma ablation using reactive oxygen species, hybrid or combined energy sources, direct current electrolysis used alone for ablation, and radioactive seed implantation.
3 | Data Computational Framework/Data Processing (Sensor Data Processing) | The category includes patents/published applications disclosing different types of data processing frameworks/models.
3.1 | Signal Processing Algorithms | This category includes patents/published applications that apply physics-based mathematical techniques to raw sensor signals to derive tissue characterization metrics. Typical approaches include A-mode echo integration for measuring bronchial diameter, FFT or wavelet analysis of electrograms, impedance ratio calculations across frequencies, energy estimation from real-time voltage–current data, and adaptive template filtering following pulsed-field ablation (PFA).
3.2 | Image Processing Algorithms | This category includes patents/published applications that extract tissue characterization information using spatial analysis of 2D or 3D medical images. Examples include CNN-based lesion segmentation from CT/MRI, Gaussian splatting for cardiac mapping, conformity index calculation between planned and actual ablation zones, and 3D temperature mapping using MRI thermometry.
3.3 | Statistical / Empirical Models | This category includes patents/published applications that use experimentally derived thresholds, reference curves, or mathematical formulas to interpret sensor data and classify tissue state. For example, impedance change curves may be compared with tissue-thickness libraries, formulas may estimate PFA lesion depth based on parameters like voltage, temperature, pulse width, and contact force, and impedance-time profiles may be used to classify tissue as normal, dehydrated, or fused.
3.4 | AI/ML | This category includes patents/published applications that use trained machine learning models to analyze sensor or imaging data for tissue characterization and lesion assessment. Examples include CNN-based segmentation of ablation zones, fuzzy C-means clustering of RF ultrasound backscatter for real-time detection, and machine learning approaches that generate electro-anatomical maps by combining imaging and electrogram data.
3.5 | Multimodal Data Fusion | This category includes patents/published applications that combine data from two or more different sensing modalities to generate a unified tissue characterization output. Examples include integrating photoacoustic, OCT, ultrasound, and temperature sensing on a single catheter, fusing electrical mapping with CT or MRI for electro-anatomical mapping, and combining impedance with contact force to assess the electrode-tissue interface.
4 | Tissue / Anatomical Target | The category includes patents/published applications disclosing different tissues of anatomical target getting treated by different ablation procedures.
4.1 | Cardiac | This category includes patents/published applications related to heart tissue and intracardiac structures, such as the myocardium (e.g., AF and VT ablation), pulmonary vein ostia, His bundle, cardiac conduction system, heart valve tissue, and intracardiac imaging used during ablation procedures.
4.2 | Renal / Perirenal | This category covers patents/published applications focused on kidney tissue and surrounding renal nerves, including procedures like renal sympathetic denervation using catheter-based approaches, percutaneous kidney tumor ablation, and sensing of renal nerve activity.
4.3 | Pulmonary / Airway | This category includes patents/published applications related to the lung parenchyma, bronchial walls, and airway structures. Examples include bronchial nerve denervation (for asthma or COPD), lung tumor ablation (via transbronchial or percutaneous approaches), pulmonary artery ablation, and nasal nerve treatments.
4.4 | Gastrointestinal | This category covers patents/published applications related to tissues of the gastrointestinal (GI) tract, including the esophagus, stomach, intestine, colon, and rectum. Examples include Barrett’s esophagus RF ablation, GI mucosal photothermal ablation, small intestine electroporation, and esophageal temperature monitoring during cardiac ablation.
4.5 | Oncology / Tumor | This category covers patents/published applications focused on solid tumor ablation across organs such as the liver, lung, kidney, prostate, and brain. The primary goal is tumor destruction using techniques like RF, microwave, laser, cryoablation, HIFU, and PEF. It also includes technologies for tumor boundary detection (e.g., bioimpedance for PEF electrode placement) and image-guided approaches such as MRI-guided laser interstitial thermal therapy (LITT).
4.6 | Neurological / Nerve | This category includes patents/published applications related to the use of peripheral nerves, nerve ganglia, or nerve trunks for ablation or functional assessment. It covers approaches such as greater splanchnic nerve pulsed field ablation, posterior nasal nerve RF ablation for rhinitis, His bundle pacing and mapping, pre-ablation sensory or motor nerve stimulation testing, and OPM-based detection of renal nerve activity.
4.7 | Generic (No Organ Specified) | This category includes patents/published applications covering platform technologies that can be used across multiple organs, without focusing on a specific anatomical target. Examples include broad impedance-based contact or thickness sensing systems, universal waveform generation for pulsed electric field (PEF), and general RF energy control algorithms.
4.8 | Other / Non-Specific | This category covers patents/published applications targeting anatomical areas not included in the main groups. Examples include the vascular or arterial wall (e.g., lithotripsy), skin and subcutaneous fat (e.g., RF skin tightening, fat treatment), endometrial lesions (e.g., PEF), urological organs (bladder, urethra, ureter), and mediastinal structures.
ASSUMPTIONS
Sr. No. | Assumptions ¦ Taxonomy & Assignee Categorization
1 | We have considered patents/published applications published having priority from 2020 onwards and published from 2022 onwards.
2 | We have analyzed one patent per INPADOC family.
3 | We have extracted the data using paid database, i.e., Relecura.
4 | For the analysis of non-English patents and published applications, we relied on machine-translated texts.
5 | In the current study, we have included the following patents/published applications:
6 | Patents/publications focused on in-vivo characterization of ablated tissue and/or lesions.
7 | Patents/publications disclosing catheter based solution for ablated tissue characterization.
8 | In the current study, we have not included the following patents/published applications:
9 | Patents/publications related to pre-procedural planning of ablation procedures.
10 | Patents/publications related to skin treatment in ablation procedures.
11 | Patents/publications related to electrosurgical instrument used to cut and coagulate tissue in ablation procedures.
12 | Patents/publications disclosing only tissue assessment or characterization without integration into an ablation system.
13 | Patents/publications related to non-invasive tissue ablation procedures.
14 | For universities, hospitals, and medical centers, revenue and size information has not been included.
15 | Startups are defined as companies founded after 2018.
16 | Companies are categorized by size as follows: • $0–$50M: Small • $50M–$1B: Medium • Above $1B: Large

## SEARCH STRINGS & STRATEGY
BAYSLOPE
SEARCH STRINGS
Sr. No. | Filed-1 | String-1 | Filed-2 | String-2 | Filed-3 | String-3 | Filed-4 | String-4 | Count
1 | TAC | ((tissue*) NEAR3 (characterizing OR characterization OR characterising OR characterisation OR assess* OR monitor OR monitors OR monitored OR monitoring OR evaluation OR evaluates OR evaluating OR assessment OR assessing OR assesses OR assess OR viability OR viable)) | TAC | (ablation OR ablated OR ablating OR ablates OR necrosis OR necrotic OR Necrotization OR transmurality OR "radiofrequency ablation" OR "RF ablation" OR "cryoablation" OR "cryogenic ablation" OR "freeze ablation" OR "pulsed field ablation" OR "irreversible electroporation" OR "electroporation ablation" OR "electroablation" OR "microwave ablation" OR "laser ablation" OR "laser-induced thermotherapy" OR "thermal ablation" OR "heat ablation" OR "ultrasound ablation" OR denervation OR Cauterization OR Cauterize OR Cauterizes OR Cauterizing OR Coagulation OR coblation) |  |  | FT | (catheter OR probe OR cannula OR sheath OR "tubular member" OR "elongate member" OR "elongated shaft" OR "intraluminal device" OR "intravascular device" OR "endoluminal device" OR “In-vivo” OR “In vivo” OR Intracorporeal OR Intraoperative OR “Real-time” OR “Real time”) | 365
2 | TAC | ((tissue*) NEAR3 (characterizing OR characterization OR characterising OR characterisation OR assess* OR monitor OR monitors OR monitored OR monitoring OR evaluation OR evaluates OR evaluating OR assessment OR assessing OR assesses OR assess OR viability OR viable)) | TAC | ((Tissue OR cell) NEAR3 (treatment OR treated OR viability OR viable OR "non-viable" OR "nonviable" OR damage OR injury OR scar* OR strain*)) | FT | (ablation OR ablated OR ablating OR ablates OR necrosis OR necrotic OR Necrotization OR transmurality OR "radiofrequency ablation" OR "RF ablation" OR "cryoablation" OR "cryogenic ablation" OR "freeze ablation" OR "pulsed field ablation" OR "irreversible electroporation" OR "electroporation ablation" OR "electroablation" OR "microwave ablation" OR "laser ablation" OR "laser-induced thermotherapy" OR "thermal ablation" OR "heat ablation" OR "ultrasound ablation" OR denervation OR Cauterization OR Cauterize OR Cauterizes OR Cauterizing OR Coagulation OR coblation) | FT | (catheter OR probe OR cannula OR sheath OR "tubular member" OR "elongate member" OR "elongated shaft" OR "intraluminal device" OR "intravascular device" OR "endoluminal device" OR “In-vivo” OR “In vivo” OR Intracorporeal OR Intraoperative OR “Real-time” OR “Real time”) | 319
3 | TAC | ((lesion*) NEAR3 (characterizing OR characterization OR characterising OR characterisation OR assess* OR monitor OR monitors OR monitored OR monitoring OR evaluation OR evaluates OR Verification* OR verifying OR verifies OR completeness OR depth OR size OR volume OR quality OR adequacy OR confirmation OR quality OR prediction OR predicting OR diagnosis OR measure OR measuring OR measurement)) | TAC | (ablation OR ablated OR ablating OR ablates OR necrosis OR necrotic OR Necrotization OR transmurality OR "radiofrequency ablation" OR "RF ablation" OR "cryoablation" OR "cryogenic ablation" OR "freeze ablation" OR "pulsed field ablation" OR "irreversible electroporation" OR "electroporation ablation" OR "electroablation" OR "microwave ablation" OR "laser ablation" OR "laser-induced thermotherapy" OR "thermal ablation" OR "heat ablation" OR "ultrasound ablation" OR denervation OR Cauterization OR Cauterize OR Cauterizes OR Cauterizing OR Coagulation OR coblation) |  |  | FT | (catheter OR probe OR cannula OR sheath OR "tubular member" OR "elongate member" OR "elongated shaft" OR "intraluminal device" OR "intravascular device" OR "endoluminal device" OR “In-vivo” OR “In vivo” OR Intracorporeal OR Intraoperative OR “Real-time” OR “Real time”) | 248
4 | TAC | ((lesion*) NEAR3 (characterizing OR characterization OR characterising OR characterisation OR assess* OR monitor OR monitors OR monitored OR monitoring OR evaluation OR evaluates OR Verification* OR verifying OR verifies OR completeness OR depth OR size OR volume OR quality OR adequacy OR confirmation OR quality OR prediction OR predicting OR diagnosis OR measure OR measuring OR measurement)) | TAC | ((Tissue OR cell) NEAR3 (treatment OR treated OR viability OR viable OR "non-viable" OR "nonviable" OR damage OR injury OR scar* OR strain*)) | FT | (ablation OR ablated OR ablating OR ablates OR necrosis OR necrotic OR Necrotization OR transmurality OR "radiofrequency ablation" OR "RF ablation" OR "cryoablation" OR "cryogenic ablation" OR "freeze ablation" OR "pulsed field ablation" OR "irreversible electroporation" OR "electroporation ablation" OR "electroablation" OR "microwave ablation" OR "laser ablation" OR "laser-induced thermotherapy" OR "thermal ablation" OR "heat ablation" OR "ultrasound ablation" OR denervation OR Cauterization OR Cauterize OR Cauterizes OR Cauterizing OR Coagulation OR coblation) | FT | (catheter OR probe OR cannula OR sheath OR "tubular member" OR "elongate member" OR "elongated shaft" OR "intraluminal device" OR "intravascular device" OR "endoluminal device" OR “In-vivo” OR “In vivo” OR Intracorporeal OR Intraoperative OR “Real-time” OR “Real time”) | 74
5 | TAC | ((ablation OR ablated OR ablating OR ablates OR necrosis OR necrotic OR Necrotization OR transmurality OR "radiofrequency ablation" OR "RF ablation" OR "cryoablation" OR "cryogenic ablation" OR "freeze ablation" OR "pulsed field ablation" OR "irreversible electroporation" OR "electroporation ablation" OR "electroablation" OR "microwave ablation" OR "laser ablation" OR "laser-induced thermotherapy" OR "thermal ablation" OR "heat ablation" OR "ultrasound ablation" OR denervation OR Cauterization OR Cauterize OR Cauterizes OR Cauterizing OR Coagulation OR coblation) NEAR5 (characterizing OR characterization OR characterising OR characterisation OR assess* OR monitor OR monitors OR monitored OR monitoring OR evaluation OR evaluates OR evaluating OR assessment OR assessing OR assesses OR assess OR viability OR viable)) | Class: Treatment (A61B 18/00) | A61B0018000000 |  |  | FT | (catheter OR probe OR cannula OR sheath OR "tubular member" OR "elongate member" OR "elongated shaft" OR "intraluminal device" OR "intravascular device" OR "endoluminal device" OR “In-vivo” OR “In vivo” OR Intracorporeal OR Intraoperative OR “Real-time” OR “Real time”) | 558
6 | TAC | (((Tissue OR cell) NEAR3 (treatment OR treated OR viability OR viable OR "non-viable" OR "nonviable" OR damage OR injury OR scar* OR strain*)) NEAR5 (characterizing OR characterization OR characterising OR characterisation OR assess* OR monitor OR monitors OR monitored OR monitoring OR evaluation OR evaluates OR evaluating OR assessment OR assessing OR assesses OR assess OR viability OR viable)) | Class: Treatment (A61B 18/00) | A61B0018000000 |  |  | FT | (catheter OR probe OR cannula OR sheath OR "tubular member" OR "elongate member" OR "elongated shaft" OR "intraluminal device" OR "intravascular device" OR "endoluminal device" OR “In-vivo” OR “In vivo” OR Intracorporeal OR Intraoperative OR “Real-time” OR “Real time”) | 84
7 | TAC | ((tissue*) NEAR3 (characterizing OR characterization OR characterising OR characterisation OR assess* OR monitor OR monitors OR monitored OR monitoring OR evaluation OR evaluates OR evaluating OR assessment OR assessing OR assesses OR assess OR viability OR viable)) | FT | (ablation OR ablated OR ablating OR ablates OR necrosis OR necrotic OR Necrotization OR transmurality OR "radiofrequency ablation" OR "RF ablation" OR "cryoablation" OR "cryogenic ablation" OR "freeze ablation" OR "pulsed field ablation" OR "irreversible electroporation" OR "electroporation ablation" OR "electroablation" OR "microwave ablation" OR "laser ablation" OR "laser-induced thermotherapy" OR "thermal ablation" OR "heat ablation" OR "ultrasound ablation" OR denervation OR Cauterization OR Cauterize OR Cauterizes OR Cauterizing OR Coagulation OR coblation) | Class: Treatment (A61B 18/00) | A61B0018000000 | FT | (catheter OR probe OR cannula OR sheath OR "tubular member" OR "elongate member" OR "elongated shaft" OR "intraluminal device" OR "intravascular device" OR "endoluminal device" OR “In-vivo” OR “In vivo” OR Intracorporeal OR Intraoperative OR “Real-time” OR “Real time”) | 331
8 | TAC | ((lesion*) NEAR3 (characterizing OR characterization OR characterising OR characterisation OR assess* OR monitor OR monitors OR monitored OR monitoring OR evaluation OR evaluates OR Verification* OR verifying OR verifies OR completeness OR depth OR size OR volume OR quality OR adequacy OR confirmation OR quality OR prediction OR predicting OR diagnosis OR measure OR measuring OR measurement)) | FT | (ablation OR ablated OR ablating OR ablates OR necrosis OR necrotic OR Necrotization OR transmurality OR "radiofrequency ablation" OR "RF ablation" OR "cryoablation" OR "cryogenic ablation" OR "freeze ablation" OR "pulsed field ablation" OR "irreversible electroporation" OR "electroporation ablation" OR "electroablation" OR "microwave ablation" OR "laser ablation" OR "laser-induced thermotherapy" OR "thermal ablation" OR "heat ablation" OR "ultrasound ablation" OR denervation OR Cauterization OR Cauterize OR Cauterizes OR Cauterizing OR Coagulation OR coblation) | Class: Treatment (A61B 18/00) | A61B0018000000 | FT | (catheter OR probe OR cannula OR sheath OR "tubular member" OR "elongate member" OR "elongated shaft" OR "intraluminal device" OR "intravascular device" OR "endoluminal device" OR “In-vivo” OR “In vivo” OR Intracorporeal OR Intraoperative OR “Real-time” OR “Real time”) | 196
9 | TAC | ((tissue*) NEAR3 (characterizing OR characterization OR characterising OR characterisation OR assess* OR monitor OR monitors OR monitored OR monitoring OR evaluation OR evaluates OR evaluating OR assessment OR assessing OR assesses OR assess OR viability OR viable)) | FT | ((Tissue OR cell) NEAR3 (treatment OR treated OR viability OR viable OR "non-viable" OR "nonviable" OR damage OR injury OR scar* OR strain*)) | Class: Treatment (A61B 18/00) | A61B0018000000 | FT | (catheter OR probe OR cannula OR sheath OR "tubular member" OR "elongate member" OR "elongated shaft" OR "intraluminal device" OR "intravascular device" OR "endoluminal device" OR “In-vivo” OR “In vivo” OR Intracorporeal OR Intraoperative OR “Real-time” OR “Real time”) | 309
10 | TAC | ((lesion*) NEAR3 (characterizing OR characterization OR characterising OR characterisation OR assess* OR monitor OR monitors OR monitored OR monitoring OR evaluation OR evaluates OR Verification* OR verifying OR verifies OR completeness OR depth OR size OR volume OR quality OR adequacy OR confirmation OR quality OR prediction OR predicting OR diagnosis OR measure OR measuring OR measurement)) | FT | ((Tissue OR cell) NEAR3 (treatment OR treated OR viability OR viable OR "non-viable" OR "nonviable" OR damage OR injury OR scar* OR strain*)) | Class: Treatment (A61B 18/00) | A61B0018000000 | FT | (catheter OR probe OR cannula OR sheath OR "tubular member" OR "elongate member" OR "elongated shaft" OR "intraluminal device" OR "intravascular device" OR "endoluminal device" OR “In-vivo” OR “In vivo” OR Intracorporeal OR Intraoperative OR “Real-time” OR “Real time”) | 154
11 | TAC | ((energy OR microwave OR "radio frequency" OR ultrasound OR laser) NEAR4 (probe OR probes)) | TAC | ((lesion OR tissue) NEAR3 (size OR shape OR volume OR dimension OR depth OR temperature)) |  |  |  |  | 78
12 | Assignee | "Boston scientific", "Delft University of Technology", Abbott, DIRECTSENSE, MedLumics, "Columbia University", "Siemens Healthineers", Erasmus, "TU Delft", "Imricor Medical Systems", Farapulse, LuxCath, "480 Biomedical", CardioFocus, SoundCath, Philips, "Conavi Medical", "Galvanize Therapeutics", "Pulse Biosciences", "Field Medical", "Adagio Medical", Kardium, CardioFocus, "Acutus Medical", Affera | Abstract | ((catheter OR probes OR probe) NEAR10 (ablation OR ablated)) AND ((lesion OR tissue) NEAR7 (size OR shape OR dimension OR depth OR temperature OR volume OR characterization OR characterisation OR assessment)) |  |  |  |  | 6
13 | Assignee | "Boston scientific", "Delft University of Technology", Abbott, DIRECTSENSE, MedLumics, "Columbia University", "Siemens Healthineers", Erasmus, "TU Delft", "Imricor Medical Systems", Farapulse, LuxCath, "480 Biomedical", CardioFocus, SoundCath, Philips, "Conavi Medical", "Galvanize Therapeutics", "Pulse Biosciences", "Field Medical", "Adagio Medical", Kardium, CardioFocus, "Acutus Medical", Affera | Title | ((catheter OR probes OR probe) NEAR10 (ablation OR ablated)) AND ((lesion OR tissue) NEAR7 (size OR shape OR dimension OR depth OR temperature OR volume OR characterization OR characterisation OR assessment)) |  |  |  |  | 3
14 | Class: Diagnosis (A61B 1/00, A61B 5/00, A61B 6/00, A61B 8/00) | A61B0001000000,A61B0005000000,A61B0006000000,A61B0008000000 | Class: Treatment (A61B 18/00) | A61B0018000000 | FT (Tissue Assessment & Ablation) | ((tissue*) NEAR3 (characterizing OR characterization OR characterising OR characterisation OR assess* OR monitor OR monitors OR monitored OR monitoring OR evaluation OR evaluates OR evaluating OR assessment OR assessing OR assesses OR assess OR viability OR viable)) AND (ablation OR ablated OR ablating OR ablates OR necrosis OR necrotic OR Necrotization OR transmurality OR "radiofrequency ablation" OR "RF ablation" OR "cryoablation" OR "cryogenic ablation" OR "freeze ablation" OR "pulsed field ablation" OR "irreversible electroporation" OR "electroporation ablation" OR "electroablation" OR "microwave ablation" OR "laser ablation" OR "laser-induced thermotherapy" OR "thermal ablation" OR "heat ablation" OR "ultrasound ablation" OR denervation OR Cauterization OR Cauterize OR Cauterizes OR Cauterizing OR Coagulation OR coblation) | FT | (catheter OR probe OR cannula OR sheath OR "tubular member" OR "elongate member" OR "elongated shaft" OR "intraluminal device" OR "intravascular device" OR "endoluminal device" OR “In-vivo” OR “In vivo” OR Intracorporeal OR Intraoperative OR “Real-time” OR “Real time”) | 728
15 | Class: Diagnosis (A61B 1/00, A61B 5/00, A61B 6/00, A61B 8/00) | A61B0001000000,A61B0005000000,A61B0006000000,A61B0008000000 | Class: Treatment (A61B 18/00) | A61B0018000000 | FT (Tissue Assessment & Tissue Treatment) | ((tissue*) NEAR3 (characterizing OR characterization OR characterising OR characterisation OR assess* OR monitor OR monitors OR monitored OR monitoring OR evaluation OR evaluates OR evaluating OR assessment OR assessing OR assesses OR assess OR viability OR viable)) AND ((Tissue OR cell) NEAR3 (treatment OR treated OR viability OR viable OR "non-viable" OR "nonviable" OR damage OR injury OR scar* OR strain*)) | FT | (catheter OR probe OR cannula OR sheath OR "tubular member" OR "elongate member" OR "elongated shaft" OR "intraluminal device" OR "intravascular device" OR "endoluminal device" OR “In-vivo” OR “In vivo” OR Intracorporeal OR Intraoperative OR “Real-time” OR “Real time”) | 645
16 | Class: Diagnosis (A61B 1/00, A61B 5/00, A61B 6/00, A61B 8/00) | A61B0001000000,A61B0005000000,A61B0006000000,A61B0008000000 | Class: Treatment (A61B 18/00) | A61B0018000000 | FT (Lesion Assessment & Ablation) | ((lesion*) NEAR3 (characterizing OR characterization OR characterising OR characterisation OR assess* OR monitor OR monitors OR monitored OR monitoring OR evaluation OR evaluates OR Verification* OR verifying OR verifies OR completeness OR depth OR size OR volume OR quality OR adequacy OR confirmation OR quality OR prediction OR predicting OR diagnosis OR measure OR measuring OR measurement)) AND (ablation OR ablated OR ablating OR ablates OR necrosis OR necrotic OR Necrotization OR transmurality OR "radiofrequency ablation" OR "RF ablation" OR "cryoablation" OR "cryogenic ablation" OR "freeze ablation" OR "pulsed field ablation" OR "irreversible electroporation" OR "electroporation ablation" OR "electroablation" OR "microwave ablation" OR "laser ablation" OR "laser-induced thermotherapy" OR "thermal ablation" OR "heat ablation" OR "ultrasound ablation" OR denervation OR Cauterization OR Cauterize OR Cauterizes OR Cauterizing OR Coagulation OR coblation) | FT | (catheter OR probe OR cannula OR sheath OR "tubular member" OR "elongate member" OR "elongated shaft" OR "intraluminal device" OR "intravascular device" OR "endoluminal device" OR “In-vivo” OR “In vivo” OR Intracorporeal OR Intraoperative OR “Real-time” OR “Real time”) | 467
17 | Class: Diagnosis (A61B 1/00, A61B 5/00, A61B 6/00, A61B 8/00) | A61B0001000000,A61B0005000000,A61B0006000000,A61B0008000000 | Class: Treatment (A61B 18/00) | A61B0018000000 | FT (Lesion Assessment & Ablation) | ((ablation OR ablated OR ablating OR ablates OR necrosis OR necrotic OR Necrotization OR transmurality OR "radiofrequency ablation" OR "RF ablation" OR "cryoablation" OR "cryogenic ablation" OR "freeze ablation" OR "pulsed field ablation" OR "irreversible electroporation" OR "electroporation ablation" OR "electroablation" OR "microwave ablation" OR "laser ablation" OR "laser-induced thermotherapy" OR "thermal ablation" OR "heat ablation" OR "ultrasound ablation" OR denervation OR Cauterization OR Cauterize OR Cauterizes OR Cauterizing OR Coagulation OR coblation) NEAR3 (characterizing OR characterization OR characterising OR characterisation OR assess* OR monitor OR monitors OR monitored OR monitoring OR evaluation OR evaluates OR evaluating OR assessment OR assessing OR assesses OR assess OR viability OR viable)) | FT | (catheter OR probe OR cannula OR sheath OR "tubular member" OR "elongate member" OR "elongated shaft" OR "intraluminal device" OR "intravascular device" OR "endoluminal device" OR “In-vivo” OR “In vivo” OR Intracorporeal OR Intraoperative OR “Real-time” OR “Real time”) | 712
18 | Class: Diagnosis (A61B 1/00, A61B 5/00, A61B 6/00, A61B 8/00) | A61B0001000000,A61B0005000000,A61B0006000000,A61B0008000000 | Class: Treatment (A61B 18/00) | A61B0018000000 | FT (Lesion Assessment & Ablation) | (((Tissue OR cell) NEAR3 (treatment OR treated OR viability OR viable OR "non-viable" OR "nonviable" OR damage OR injury OR scar* OR strain*)) NEAR3 (characterizing OR characterization OR characterising OR characterisation OR assess* OR monitor OR monitors OR monitored OR monitoring OR evaluation OR evaluates OR evaluating OR assessment OR assessing OR assesses OR assess OR viability OR viable)) | FT | (catheter OR probe OR cannula OR sheath OR "tubular member" OR "elongate member" OR "elongated shaft" OR "intraluminal device" OR "intravascular device" OR "endoluminal device" OR “In-vivo” OR “In vivo” OR Intracorporeal OR Intraoperative OR “Real-time” OR “Real time”) | 265
| TAC | Title, Abstract & Claims
| FT | Full Text
| Class | Classifications
SEARCH STRATEGIES
| Classes | Definitions |  | Assignees
| A61B0001000000 | Instruments for performing medical examinations of the interior of cavities or tubes of the body by visual or photographical inspection, e.g. endoscopes |  | Boston scientific |  | CardioFocus
| A61B0005000000 | Measuring for diagnostic purposes |  | Delft University of Technology |  | SoundCath
| A61B0006000000 | Apparatus or devices for radiation diagnosis; Apparatus or devices for radiation diagnosis combined with radiation therapy equipment |  | Abbott |  | Philips
| A61B0008000000 | Diagnosis using ultrasonic, sonic or infrasonic waves |  | DIRECTSENSE |  | Conavi Medical
| A61B0018000000 | Surgical instruments, devices or methods for transferring non-mechanical forms of energy to or from the body |  | MedLumics |  | Galvanize Therapeutics
|  |  |  | Columbia University |  | Pulse Biosciences
|  |  |  | Siemens Healthineers |  | Field Medical
|  |  |  | Erasmus |  | Adagio Medical
|  |  |  | Imricor Medical Systems |  | Kardium
|  |  |  | Farapulse |  | CardioFocus
|  |  |  | LuxCath |  | Acutus Medical
|  |  |  | 480 Biomedical |  | Affera

## DISCLAIMER
BAYSLOPE
DISCLAIMER
The report contains no legal opinion and only seeks to provide a technical analysis. However, the report may include personal, non-legal judgment of technical researcher(s) involved in the search, analysis, and preparation of this report. Therefore, no content of this report should be construed as having any legal weight or being legally dispositive in any manner.