381 | WO2024081158A3 | METHOD AND APPARATUS FOR TREATMENT OF PULMONARY CONDITIONS | Apparatus and methods for deactivating bronchial nerves and ablating smooth muscle extending along a bronchial branch of a mammalian subject to treat asthma and related conditions. An ultrasonic transducer (11) is inserted into the bronchus as, for example, by advancing the distal end of a catheter (10) bearing the transducer into the bronchial section to be treated. The ultrasonic transducer emits ultrasound so as to heat tissues throughout a relatively large impact volume (13) to a temperature sufficient to inactivate nerve conduction and ablate smooth muscle but insufficient to cause rapid ablation or necrosis of the surrounding tissues. The treatment can be performed without locating or focusing on individual bronchial nerves or smooth muscle. | 1. IN THE CLAIMS: 1. System for simultaneously inactivating bronchial nerve conduction and ablating smooth muscle in a mammalian subject, comprising: an elongated member supporting an ultrasound transducer adapted for insertion into a bronchial branch of a bronchial tree of the mammalian subject and for transmitting ultrasound energy; and an actuator or control unit electrically connected to the transducer, the actuator or control unit adapted to control the ultrasound transducer to i) transmit ultrasound energy at a sub-therapeutic level in a diagnostic mode; and ii) transmit ultrasound energy in an impact volume encompassing the bronchial branch in a therapeutic mode so that the ultrasound energy is applied at a therapeutic level sufficient to inactivate conduction of bronchial nerves and ablate smooth muscle throughout the impact volume in a therapeutic mode. 2. The system of claim 1, wherein the actuator or control unit transmits ultrasound energy in the diagnostic mode to generate a signal and integrate ultrasound echoes to measure bronchial diameter. 3. The system of claim 1, wherein the actuator or control unit is configured to analyze the impedance measurements to determine locations of bronchial cartilage rings and the actuator or control unit is additionally configured to activate the ultrasound transducer to transmit ultrasound therapeutic waveform energy between adjacent cartilaginous bronchial tissue. 4. The system of claim 1 , wherein the actuator or control unit is configured to transmit ultrasound energy in the diagnostic mode to generate a signal and analyze ultrasound echoes to ensure circumferential coupling. 5. The system of claim 1 , wherein the actuator or control unit is configured to transmit ultrasound energy in the diagnostic mode to generate a signal and analyze ultrasound echoes to ensure inter-cartilage positioning. The system of claim 1 , wherein analysis of the return signal of the transmitted pulse in the diagnostic mode determines one or both of a type or state of tissue. The system of claim 1, wherein analysis of the return signal of the transmitted in the diagnostic mode provides centering of the transducer in the bronchial branch. The system of claim 1 , wherein the actuator or control unit is configured to transmit ultrasound energy at the therapeutic level interleaved with transmitting in the diagnostic mode. The system of claim 1, wherein the actuator or control unit is configured to transmit pulsed signals at the therapeutic level. The system of claim 1, wherein the actuator or control unit further transmits ultrasound energy at the sub-therapeutic level after transmitting ultrasound energy at the therapeutic level has ablated peri bronchial tissues. The system of claim 1, further comprising a balloon, the balloon containing cooling fluid. The system of claim 11, wherein pulsating fluid within the balloon enables volumetric A mode diameter measurements by analyzing amplitude and signal width fluctuations. The system of claim 11, wherein the cooling fluid is pulsated to provide pulsating flow to measure bronchial compliance and degree of smooth muscle ablation. The system of claim 11, wherein the transducer emits ultrasonic energy at a sub-therapeutic level to detect heating of the fluid and/or efficacy of cooling. The system of claim 1, further comprising at least one electrode for making impedance measurements along a wall of the bronchial branch to locate cartilaginous tissue and spaces or gaps between cartilaginous tissue. The system of claim 1, wherein the transducer transmits ultrasound energy in the therapeutic mode to direct the ultrasound energy between cartilage rings. System for conducting pulmonary treatment in a mammalian subject, comprising: an elongated member having a balloon and an energy transducer within the balloon, the transducer adapted for insertion into a bronchial branch of a bronchial tree of the mammalian subject and for transmitting energy into surrounding tissue, and an actuator or control unit electrically connected to the transducer, the actuator or control unit being adapted to control the transducer to emit ultrasonic energy into an impact volume encompassing the bronchial branch so that the energy is applied at a therapeutic level sufficient to inactivate conduction of bronchial nerves throughout the impact volume; wherein the balloon contains cooling fluid and the cooling fluid is pulsated to measure bronchial compliance. The system of claim 17, wherein the transducer emits ultrasound energy at a sub-therapeutic level in a diagnostic mode to measure bronchial diameter to optimize dosing. The system of claim 17, wherein the cooling fluid is pulsated to generate an A-mode amplitude modulation to enable echo identification for bronchial diameter measurements. The system of claim 19, wherein the pulsating fluid within the balloon enables volumetric A mode diameter measurements by analyzing amplitude and signal width fluctuations. The system of claim 20, wherein the transducer emits ultrasonic energy at a sub-therapeutic level to detect heating of the fluid and/or efficacy of cooling. The system of claim 17, wherein the transducer emits ultrasonic energy at a sub-therapeutic level to detects heating of the fluid or efficacy of cooling. System for simultaneously inactivating bronchial nerve conduction and ablating smooth muscle in a mammalian subject, comprising: an ultrasound transducer adapted for insertion into a bronchial branch of a bronchial tree of the mammalian subject and for transmitting circumferential ultrasound energy; and an actuator or control unit electrically connected to the transducer, the actuator being adapted to control the ultrasound transducer to transmit ultrasound energy into an impact volume encompassing the bronchial branch so that the ultrasound energy is applied at a therapeutic level sufficient to inactivate conduction of bronchial nerves and ablate smooth muscle throughout the impact volume, the actuator or control unit being configured to energize the ultrasound transducer to emit a short pulse at a sub-therapeutic level, the actuator or control unit being further configured to process a volume integrated A- mode signal, which represents an accumulated intensity of the circumferential ultrasound echoes, to enable a precise dosing of ultrasound energy in the impact volume. The system of claim 23, wherein the bronchial branch is a 3rd or 4th generation bronchial branch and the actuator or control unit is adapted to control the ultrasound transducer depending on volumetric A mode bronchial diameter measurement to transmit ultrasound energy at an acoustic power level of approximately 10 to approximately 30 watts for approximately 10 to approximately 30 seconds to provide an absorbed dose of approximately 100 to approximately 900 joules in the impact volume depending on bronchial diameter to optimize simultaneous ablation of both smooth muscle and bronchial nerves in the impact volume surrounding the 3rd or 4th generation bronchial branch. The system of claim 24, further comprising a balloon filled with a coupling fluid, said ultrasound transducer being disposed inside said balloon, said balloon being expandable into contact with a wall of the bronchial branch, wherein the actuator or control unit is adapted to control the temperature of the coupling fluid together with dose parameters of acoustic power level and duration so as to maintain a temperature of the bronchial wall below 40°C while achieving a temperature above 60°C throughout the impact volume including the smooth muscle and or bronchial nerves surrounding 3rd or 4th generation bronchial branch. The system of claim 23, wherein the actuator or control unit is adapted to control the ultrasound transducer depending on volumetric A mode bronchial diameter measurement to transmit ultrasound energy at an acoustic power level of approximately 10 to approximately 30 watts for approximately 10 to approximately 30 seconds to provide an absorbed dose of approximately 100 to approximately 900 joules in the impact volume, the acoustic power level and duration of energy application depending on bronchial diameter to optimize the separate ablation of both smooth muscle and bronchial nerves in impact volumes surrounding bronchial branches of larger diameters. The system of claim 23, wherein the actuator or control unit is adapted to control the ultrasound transducer to transmit the ultrasound energy in a pulsed function interleaved with volumetric A mode diagnostic acquisitions to enable a quasi-simultaneous therapeutic/diagnostic mode. The system of claim 23, further comprising a catheter with a distal end and a proximal end, the transducer being mounted to the catheter adjacent the distal end inside a compliant balloon, filled with circulating fluid in order to cool the ultrasound transducer, the ultrasound transducer being disposed in alignment with and centered on an axis of the bronchial branch. The system of claim 23, further comprising a catheter with a distal end and a proximal end, the transducer being mounted to the catheter adjacent the distal end inside a compliant balloon filled with circulating, pulsating fluid in order to enable volumetric A mode diameter measurements by analyzing amplitude and signal width fluctuations. The system of claim 23, further comprising a catheter with a distal end and a proximal end, the transducer being mounted to the catheter adjacent the distal end inside a compliant balloon filled with circulating, pulsating fluid, the actuator or control unit being configured for measuring bronchial compliance and therewith ASM ablation by analyzing resistance to balloon pressure pulsations. The system of claim 30, wherein the catheter is configured to hold the ultrasound transducer out of contact with the wall of the bronchial branch. The system of claim 30, wherein the ultrasound transducer has an axis, the catheter being constructed and arranged to hold the ultrasound transducer so that the axis of the ultrasound transducer is generally parallel to a longitudinal axis of the bronchial branch, the ultrasound transducer being adapted to transmit the ultrasound energy in a 360° cylindrical pattern surrounding the axis of the ultrasound transducer. The system of claim 23, further comprising a catheter including a compliant balloon surrounding the ultrasound transducer and arranged to hold the ultrasound transducer substantially centered in the bronchial branch, the compliant balloon including a porous membrane engageable with the wall of the bronchial branch and filled with a coupling fluid for improved acoustic coupling, the fluid being exudable through the porous membrane enabling sliding of the balloon axially while in engagement with the wall of the bronchial branch to achieve inter cartilage positioning. The system of claim 33, wherein the compliant balloon is provided with electrodes, said actuator or control unit being operatively connected to said electrodes for making impedance measurements along a wall of the bronchial branch to enable a locating of cartilaginous tissue and spaces or gaps between cartilaginous tissue. The system of claim 34, wherein the actuator or control unit is configured to analyze the impedance measurements to determine locations of bronchial cartilage rings and the actuator or control unit being additionally configured to activate the ultrasound transducer to transmit ultrasound therapeutic waveform energy between adjacent cartilaginous bronchial tissue. The system of claim 23, wherein the ultrasound transducer is further adapted to receive volumetric A mode ultrasound signals representing or encoding bronchial geometry, the actuator or control unit being further adapted to: control the ultrasound transducer to transmit measurement ultrasound energy at a level below therapeutic levels, receive volumetric A mode echo signals from the ultrasound transducer representing reflected measurement ultrasonic energy; analyze the received volumetric A mode echo signals; and determine a size of the bronchial branch based on the received echo volumetric A mode signals. The system of claim 23, further comprising an axially movable circular loop electrode operatively connected to said actuator or control unit, said actuator or control unit being adapted to receive electrical impedance measurements representing bronchial cartilage geometry from said circular loop electrode. The system of claim 23, further comprising a compliant balloon surrounding the ultrasound transducer and provided with multiple circular electrodes on an outer surface, said actuator or control unit being operatively connected to said circular electrodes to receive therefrom electrical impedance measurements representing bronchial cartilage geometry. The system of claim 38, wherein the actuator or control unit is adapted to control the ultrasound transducer to vary the acoustic power used to transmit the ultrasound energy depending on the determined size of the bronchial branch. A method for pulmonary treatment, comprising the steps of: inserting an ultrasound transducer into a bronchial section of the mammalian subject; and actuating the transducer to transmit therapeutically effective ultrasound energy into an impact volume of at least approximately 1 cm3, encompassing the bronchial section so that the therapeutically effective ultrasound energy inactivates all bronchial nerves and smooth muscle in the impact volume. The method of claim 40, wherein the ultrasound energy is transmitted at an acoustic power level of approximately 10 to approximately 30 watts for approximately 10 to approximately 30 seconds to provide an absorbed dose of approximately 100 to approximately 900 joules throughout the impact volume, further comprising conducting a volumetric A mode ultrasound signal scan, analyzing resulting volumetric A mode echo signals, and selecting intensity and duration parameters of the ultrasound energy in accordance with a geometry of the bronchial section and intended ablation of nerves and or smooth muscle. The method of claim 41, wherein the ultrasound transducer is disposed inside a compliant balloon attached proximate a distal end of a catheter and filled with coupling fluid, further comprising controlling a temperature of the coupling fluid in the balloon to prevent or reduce thermal damage to a wall of the bronchial section while enabling therapeutically effective increase in temperature of the bronchial nerves and or smooth muscle in the impact volume. The method of claim 40, wherein the steps of inserting the ultrasound transducer and actuating the transducer to transmit ultrasound energy are performed without determining actual locations of the bronchial nerves and the smooth muscle. The method of claim 40, wherein the bronchial section is a third-generation branch of a bronchial tree of the mammalian subject, a major source of resistance to airflow, the step of actuating the ultrasound transducer being performed so that a single application of the ultrasound energy at a treatment site in the third-generation branch is effective to inactivate nerve conduction distal of the treatment site and to ablate all smooth muscle throughout the impact volume. The method of claim 44, further comprising the steps of: repositioning the ultrasound transducer distal into a next or fourth generation bronchial branch, another major source of airflow resistance, after the step of actuating the transducer; and then repeating the step of actuating the transducer. The method of claim 40, wherein the step of actuating the transducer is performed so that the ultrasound energy is transmitted in a pulsed function. The method of claim 46, wherein the ultrasound transducer is operated in a pulsed mode performing quasi-simultaneous volumetric A mode measurements and treatment. . The method of claim 40, wherein the step of inserting the ultrasound transducer into the bronchial section is performed through the working channel of a bronchoscope under visual guidance. . The method of claim 40, wherein the step of inserting the ultrasound transducer is performed through a steerable sheath. . The method of claim 40, wherein the step of inserting the ultrasound transducer is performed without a sheath or bronchoscope and with a steerable ultrasound catheter. 1. The method of claim 40, wherein the step of inserting the ultrasound treatment catheter is performed over a guide wire which has been placed through a working channel of a bronchoscope. 2. System for conducting pulmonary treatment in a mammalian subject, comprising: an energy transducer adapted for insertion into a bronchial branch of a bronchial tree of the mammalian subject and for transmitting energy into surrounding tissue; an actuator or control unit electrically connected to the transducer, the actuator being adapted to control the transducer to emit energy into an impact volume encompassing the bronchial branch so that the energy is applied at a therapeutic level sufficient to inactivate conduction of bronchial nerves throughout the impact volume; and at least one electrode, the actuator or control unit operatively connected to the at least one electrode for making impedance measurements along a wall of the bronchial branch to enable locating of cartilaginous tissue and spaces or gaps between cartilaginous tissue. 3. The system of claim 52, wherein the at least one electrode comprises a plurality of electrodes and the actuator or control unit is configured to direct RF current through selected ones of said plurality of electrodes or portions thereof positioned between cartilage rings based on the impedance measurements. 54. The system of claim 53, wherein the electrodes are supported by a compliant balloon. 55. The system of claim 55, wherein the transducer is positioned within the balloon. 56. The system of claim 54, wherein the balloon contains cooling fluid. | Apparatus and methods for treating asthma by deactivating bronchial nerves and ablating bronchial smooth muscle using an ultrasound transducer inserted through a catheter. Focused ultrasound heats a relatively large volume of bronchial wall tissue at once. | R | Tissue Ablation ¦ Monitor tissue ¦ impedance measurements ¦ impedance measurements along a wall of the bronchial branch to enable a locating of cartilaginous tissue "Additionally, further features of the system, device and method can include one or more of the following i) pulsating cooling fluid within the balloon for diameter measurements; ii) pulsating cooling fluid within the balloon to measure bronchial compliance and ablation; iii) utilization of ultrasound energy at a sub-therapeutic level to determine tissue type or state; and/or iv) utilization of ultrasound energy at a sub-therapeutic level to assess treatment. Each of these are discussed in more detail below." "Apparatus and methods for deactivating bronchial nerves and ablating smooth muscle extending along a bronchial branch of a mammalian subject to treat asthma and related conditions. An ultrasonic transducer (11) is inserted into the bronchus as, for example, by advancing the distal end of a catheter (10) bearing the transducer into the bronchial section to be treated. The ultrasonic transducer emits ultrasound so as to heat tissues throughout a relatively large impact volume (13) to a temperature sufficient to inactivate nerve conduction and ablate smooth muscle but insufficient to cause rapid ablation or necrosis of the surrounding tissues. The treatment can be performed without locating or focusing on individual bronchial nerves or smooth muscle." |  |  | 1 | [12 empty] | 1 |  |  | 1 | [7 empty] | 1 |  |  |  |  | AERWAVE MEDICAL INC | AERWAVE MEDICAL INC | 2023-10-06 | 2023 | 2024-05-23 | 2024 | 2022-10-14 | 2022 | WO2023US34606 | EP4601549A2¦WO2024081158A2
382 | CN119836310A | System and method for subclavian periarterial denervation | A method of treating a heart disease in a subject includes percutaneously introducing a catheter into the vasculature of the subject. The catheter may include a neuromodulation element. A neuromodulation element may be positioned on the expansion element. A neuromodulation element of the catheter may be positioned in a subclavian artery of a subject. The catheter may be used to electrically stimulate at least one of a dorsal subclavian loop or a ventral subclavian loop. The method may include confirming stimulation to a dorsal subclavian loop and/or a ventral subclavian loop by monitoring cardiac parameters. After the stimulation to the dorsal subclavian loop and/or the ventral subclavian loop is confirmed, ablation energy can be provided to the dorsal subclavian loop and/or the ventral subclavian loop. | 1.A method for treating a cardiac disease in a subject, the method comprising: percutaneously introducing a catheter into the vasculature of the subject, wherein the catheter includes a neuromodulatory element; positioning the neuromodulatory element in the left subclavian artery of the subject; electrically stimulating the left dorsal subclavian loop; confirming the stimulation of the left dorsal subclavian loop by monitoring a first cardiac parameter; providing ablative energy to the left dorsal subclavian loop after confirming the stimulation of the left dorsal subclavian loop; electrically stimulating the left ventral subclavian loop; confirming the stimulation of the left ventral subclavian loop by monitoring a second cardiac parameter; providing ablative energy to the left ventral subclavian loop after confirming the stimulation of the left ventral subclavian loop; positioning the neuromodulatory element in the right subclavian artery of the subject; electrically stimulating the right dorsal subclavian loop; confirming the stimulation of the right dorsal subclavian loop by monitoring a third cardiac parameter; After confirming stimulation of the right dorsal subclavian loop, providing ablative energy to the right dorsal subclavian loop; electrically stimulating the right ventral subclavian loop; confirming stimulation of the right ventral subclavian loop by monitoring a fourth cardiac parameter; and after confirming stimulation of the right ventral subclavian loop, providing ablative energy to the right ventral subclavian loop. 2.The method of claim 1, wherein the first cardiac parameter is different from at least one of the second cardiac parameter, the third cardiac parameter, or the fourth cardiac parameter. 3.The method of claim 2, wherein the second cardiac parameter is different from at least one of the first cardiac parameter, the third cardiac parameter, or the fourth cardiac parameter. 4.The method of claim 3, wherein the third cardiac parameter is different from at least one of the first cardiac parameter, the second cardiac parameter, or the fourth cardiac parameter. 5.The method of claim 4, wherein the fourth cardiac parameter is different from at least one of the first cardiac parameter, the second cardiac parameter, or the third cardiac parameter. 6.The method according to claim 1, wherein the first cardiac parameter is the same as at least one of the second cardiac parameter, the third cardiac parameter or the fourth cardiac parameter. 7.The method according to claim 1, wherein the second cardiac parameter is the same as at least one of the first cardiac parameter, the third cardiac parameter or the fourth cardiac parameter. 8.The method of claim 1, wherein the third cardiac parameter is the same as at least one of the first cardiac parameter, the second cardiac parameter, or the fourth cardiac parameter. 9.The method of claim 1, wherein the fourth cardiac parameter is the same as at least one of the first cardiac parameter, the second cardiac parameter, or the third cardiac parameter. 10.The method according to any one of claims 1 to 9, wherein at least one of the first cardiac parameter, the second cardiac parameter, the third cardiac parameter or the fourth cardiac parameter includes at least one of arterial micro-blood pressure changes, whether AF is induced, whether arrhythmia is induced, whether cardiac cycle length changes are induced, repetitive recovery curves, right atrial ERP, left atrial ERP, dERP, heart rate, IACT, or a multipolar catheter unipolar electrogram of short DI under or after steady-state RV pacing. 11. The method according to any one of claims 1 to 9 further comprises re-stimulating the subclavian loop after providing the ablation energy, and further providing ablation energy to the subclavian loop if the cardiac parameters confirm the stimulation. 12.A method according to claim 11, comprising repeating said restimulation and providing further ablation energy until said cardiac parameters no longer confirm stimulation. 13. The method according to any one of claims 1 to 9, wherein the catheter includes an expansion element and the neuromodulation element is positioned on the expansion element. 14. The method of claim 13, comprising expanding the expansion element to dilate the vasculature. 15.A method for treating heart disease in a subject, the method comprising: positioning a neuromodulatory element in the subclavian artery of the subject; electrically stimulating at least one of the dorsal subclavian loop or the ventral subclavian loop; confirming the stimulation of the dorsal subclavian loop and/or the ventral subclavian loop by monitoring cardiac parameters; and providing ablation energy to the dorsal subclavian loop and/or the ventral subclavian loop after confirming the stimulation of the dorsal subclavian loop and/or the ventral subclavian loop. 16.The method of claim 15, wherein the ablation energy comprises radiofrequency ablation energy. 17.The method according to claim 15, wherein the ablation energy includes cryogenic ablation energy. 18. The method of claim 15, wherein the electrical stimulation comprises using a combination of electrodes of the neuromodulatory element, and wherein providing ablation energy comprises using the same combination of electrodes of the neuromodulatory element. 19.The method of claim 15, further comprising assessing whether sympathetic nerve drive drives the refractory arrhythmia. 20. A method according to any one of claims 15 to 19, wherein the heart disease comprises atrial fibrillation. 21. twenty one. The method according to any one of claims 15 to 19, wherein the heart disease comprises a refractory arrhythmia or a ventricular arrhythmia. 22. twenty two.A catheter for treating heart disease in a subject, the catheter comprising: a slender element configured to be positioned in at least one of a left subclavian artery or a right subclavian artery, wherein a distal portion of the slender element comprises: an expansion element, a neuromodulatory element positioned on the expansion element and configured to stimulate at least one of a dorsal subclavian loop or a ventral subclavian loop and configured to ablate at least one of the dorsal subclavian loop or the ventral subclavian loop, the neuromodulatory element of the catheter being in the subclavian artery of the subject, the neuromodulatory element comprising an array of electrodes spaced circumferentially about the circumference of the catheter and longitudinally along the length of the catheter. 23. twenty three.A catheter according to claim 22, wherein the neuromodulatory element comprises a plurality of electrodes connected to a plurality of struts. 24. twenty four.A catheter according to claim 22, wherein the neuromodulatory element comprises a plurality of electrodes connected to a ring. 25.The catheter according to claim 22 includes a distal protection device located distal to the neuromodulatory element. 26. The catheter of any one of claims 23 to 25, wherein at least one of the plurality of electrodes comprises a triangular base comprising a tip configured to press into a blood vessel wall. 27. A catheter for treating a heart disease in a subject, the catheter comprising: an elongated element configured to be positioned in at least one of a left subclavian artery or a right subclavian artery, wherein a distal portion of the elongated element comprises: a neuromodulatory element configured to stimulate at least one of a dorsal subclavian haptic or a ventral subclavian haptic and configured to ablate at least one of the dorsal subclavian haptic or the ventral subclavian haptic, the neuromodulatory element comprising: An array of electrodes is spaced circumferentially about the circumference of the catheter and longitudinally along the length of the catheter. 28.A catheter according to claim 27, wherein the electrode array comprises a plurality of electrodes connected to a plurality of struts. 29.A catheter according to claim 27, wherein the electrode array comprises a plurality of electrodes connected to a ring. 30.A catheter according to claim 27, wherein the electrode array comprises a plurality of electrodes connected to a helix. 31.A catheter according to any one of claims 27 to 30, wherein at least one electrode of the electrode array comprises a triangular base comprising a tip configured to press into a blood vessel wall. 32.A catheter according to any one of claims 27 to 30, further comprising an expansion element on which the electrode array is located. 33.A catheter according to claim 32, wherein the expansion element comprises a balloon. 34.A method for treating heart disease in a subject, the method comprising: percutaneously introducing a catheter into the vascular system of the subject, wherein the catheter comprises a neuromodulatory element; positioning the neuromodulatory element in the subclavian artery of the subject; electrically stimulating at least one of the dorsal subclavian loop or the ventral subclavian loop; confirming the stimulation of the dorsal subclavian loop and/or the ventral subclavian loop by monitoring cardiac parameters; and after confirming the stimulation of the dorsal subclavian loop and/or the ventral subclavian loop, providing ablative energy to the dorsal subclavian loop and/or the ventral subclavian loop. 35.The method of claim 34, further comprising re-stimulating the subclavian loop after providing the ablative energy, and further providing ablative energy to the subclavian loop if the cardiac parameters confirm the stimulation. 36.A method according to claim 35, comprising repeating said restimulation and providing further ablation energy until said cardiac parameters no longer confirm stimulation. 37. The method of any one of claims 34 to 36, wherein percutaneously introducing the catheter into the vasculature comprises inserting the catheter into the femoral artery of the subject. 38. The method of any one of claims 34 to 36, wherein percutaneously introducing the catheter into the vasculature comprises inserting the catheter into a radial artery of the subject. 39. The method of any one of claims 34 to 36, wherein percutaneously introducing the catheter into the vasculature comprises inserting the catheter into the subject's carotid artery. 40.The method of any one of claims 34 to 36, wherein percutaneously introducing the catheter into the vasculature comprises inserting the catheter into a femoral vein of the subject, and at least one of positioning the neuromodulatory element in the left subclavian artery or positioning the neuromodulatory element in the right left subclavian artery comprises traversing from a venous vasculature to an arterial vasculature. 41.A method for treating heart disease in a subject, the method comprising: percutaneously introducing a catheter into the vascular system of the subject, wherein the catheter includes a neuromodulatory element; positioning the neuromodulatory element in the subclavian artery of the subject; electrically stimulating the dorsal subclavian loop; confirming the stimulation of the dorsal subclavian loop by monitoring a first cardiac parameter; after confirming the stimulation of the dorsal subclavian loop, providing ablative energy to the dorsal subclavian loop; electrically stimulating the ventral subclavian loop; confirming the stimulation of the ventral subclavian loop by monitoring a second cardiac parameter; and after confirming the stimulation of the ventral subclavian loop, providing ablative energy to the ventral subclavian loop. 42.The method of claim 41, wherein the first cardiac parameter is different from the second cardiac parameter. 43.The method of claim 41, wherein the first cardiac parameter is the same as the second cardiac parameter. 44.The method according to any one of claims 41 to 43, wherein at least one of the first cardiac parameter or the second cardiac parameter includes at least one of arterial micro-blood pressure changes, whether AF is induced, whether arrhythmia is induced, whether cardiac cycle length changes are induced, repetitive recovery curves, right atrial ERP, left atrial ERP, dERP, heart rate, IACT, or a multipolar catheter unipolar electrogram of short DI under or after steady-state RV pacing. 45. A method according to any one of claims 41 to 43, further comprising re-stimulating the subclavian haptic after providing the ablation energy, and further providing ablation energy to the subclavian haptic if the cardiac parameters confirm the stimulation. 46. A method according to claim 45, comprising repeating said restimulation and providing further ablation energy until said cardiac parameters no longer confirm stimulation. 47. A method according to any one of claims 41 to 43, wherein the catheter includes an expansion element, the neuromodulation element is positioned thereon, and the expansion element is optionally expanded an amount to dilate the vasculature. 48.A catheter for treating heart disease in a subject, the catheter comprising: a slender element configured to be positioned in at least one of a left subclavian artery or a right subclavian artery, wherein a distal portion of the slender element comprises: a first neuromodulatory element configured to stimulate at least one of a dorsal subclavian haptic or a ventral subclavian haptic; a second neuromodulatory element configured to ablate at least one of the dorsal subclavian haptic or the ventral subclavian haptic; and a distal protection device distal to the first neuromodulatory element and the second neuromodulatory element. 49.A catheter according to claim 48, wherein the first neuromodulatory element comprises a first plurality of electrodes directly or indirectly coupled to one or more struts, rings and/or spirals. 50.A catheter according to claim 49, wherein at least one of the plurality of electrodes comprises a triangular base comprising a tip configured to press into a blood vessel wall, wherein, optionally, the at least one electrode further comprises a plurality of electrode protrusions located on different sides of the tip of the triangular base. 51.A catheter according to any one of claims 48 to 50, wherein the second neuromodulatory element comprises the first plurality of electrodes and a second plurality of electrodes directly or indirectly coupled to a second plurality of struts, rings and/or spirals or other elements of similar functionality. 52.The catheter of claim 48, further comprising an expansion element, and wherein the first neuromodulation element and the second neuromodulation element are positioned on the expansion element, wherein the expansion element optionally comprises a balloon. 53.A catheter for treating heart disease in a subject, the catheter comprising: a slender element configured to be positioned in at least one of a left subclavian artery or a right subclavian artery, wherein a distal portion of the slender element comprises: a neuromodulatory element configured to stimulate at least one of a dorsal subclavian loop or a ventral subclavian loop, and configured to ablate at least one of the dorsal subclavian loop or the ventral subclavian loop; and a distal protection device at the distal end of the neuromodulatory element. 54.A catheter according to claim 53, wherein the neuromodulatory element includes a plurality of electrodes directly or indirectly connected to a plurality of struts, rings and/or spirals or other elements of similar functionality. 55.A catheter according to any one of claims 53 to 54, wherein at least one of the plurality of electrodes comprises a triangular base comprising a tip configured to press into a blood vessel wall, and optionally comprising a plurality of electrode protrusions located on different sides of the tip of the triangular base. 56.The catheter of claim 55, further comprising an expansion element, and wherein the neuromodulation element is positioned on the expansion element, wherein the expansion element optionally comprises a balloon. 57.A method for treating a subject, the method comprising: stimulating specific electrodes of a neuromodulatory device; providing the results to an interface computer, the computer being configured to: perform real-time data analysis; and superimpose the response to the stimulation on an anatomical image, including marking the location where the response is greatest; and performing ablation at said location. 58.The method according to claim 57 further comprises one or more of the following: after ablation, re-stimulation to verify the effect of the ablation; if the ablation is ineffective, further ablation, and optionally repeating the verification and further ablation until the ablation is effective. 59.A method for treating a subject experiencing a refractory arrhythmia, the method comprising: stabilizing the subject; assessing the cause of the refractory arrhythmia; assessing whether sympathetic nerve drive is driving the refractory arrhythmia; coordinating an assessment of an arrhythmogenic lesion if sympathetic nerve drive is not driving the refractory arrhythmia; and performing subclavian loop ablation if sympathetic nerve drive is driving the refractory arrhythmia. 60.The method of claim 59, wherein assessing whether sympathetic nerve drive is driving the refractory arrhythmia comprises infusing epidural anesthesia and/or wherein assessing whether sympathetic nerve drive is driving the refractory arrhythmia comprises shutting down an implantable cardioverter-defibrillator of the subject. 61.A method for treating heart disease in a subject, the method comprising: positioning a neuromodulatory element in the subclavian artery of the subject; and providing ablative energy to the dorsal subclavian loop and/or the ventral subclavian loop. 62.The method of claim 61, wherein providing the ablation energy comprises selectively providing the ablation energy. 63.A method according to claim 62, wherein selectively providing the ablation energy includes one or more of: non-therapeutic stimulation to determine the location of the dorsal subclavian loop and/or the ventral subclavian loop; imaging the subclavian artery to determine the possible location of the dorsal subclavian loop and/or the ventral subclavian loop. 64.The method of claim 63, wherein providing the ablation energy comprises providing the ablation energy around the subclavian artery. 65A method of treating a subject, the method comprising: positioning a neuromodulatory element of a catheter in a subclavian artery of the subject, the neuromodulatory element comprising an array of electrodes spaced circumferentially about the circumference of the catheter and longitudinally along the length of the catheter; stimulating specific electrodes in the electrode array; providing the results to an interface computer configured to: perform real-time data analysis; and superimpose the response to the stimulation on an anatomical image, including marking the location where the response is greatest; and perform ablation at said location. 66.The method according to claim 65 further includes, after ablation, restimulation to verify the effectiveness of the ablation, and if the ablation is ineffective, further ablation, and optionally includes repeating restimulation and further ablation until it is verified that the ablation is effective. 67.A method according to claim 65 or 66, wherein the neuromodulatory element is positioned on an expansion element. 68. The method of claim 67, comprising expanding the expansion element to dilate the subclavian artery. 69.The method of claim 68, comprising performing impedance measurements when the expansion element is expanded. 70.A method for treating heart disease in a subject, the method comprising: percutaneously introducing a catheter into the vascular system of the subject, wherein the catheter comprises a neuromodulatory element; positioning the neuromodulatory element in the subclavian artery of the subject; electrically stimulating at least one of a dorsal subclavian loop or a ventral subclavian loop; confirming the stimulation of the dorsal subclavian loop and/or the ventral subclavian loop by monitoring biomarkers, wherein the biomarkers comprise at least one of an inotropic biomarker, a chronotropic biomarker, a conduction biomarker, a comfort biomarker, or an inflammatory biomarker; after confirming the stimulation of the dorsal subclavian loop and/or the ventral subclavian loop, providing ablative energy to the dorsal subclavian loop and/or the ventral subclavian loop. 71.The method of claim 70, further comprising re-stimulating the subclavian loop after providing the ablative energy, and further providing ablative energy to the subclavian loop if the biomarker confirms the stimulation. 72.A method according to claim 71, comprising repeating said restimulation and providing further ablation energy until said biomarker no longer confirms stimulation. 73.The method of claim 71, wherein the biomarker comprises at least one of increased dP/dt, arterial blood pressure (systolic, diastolic, mean), pulse pressure variation, increased systemic blood pressure, or increased left ventricular dP/dt. 74.The method of claim 71, wherein the biomarker comprises at least one of a change in heart rate (a change in the R_R or Q_Q interval on a surface ECG), an increase in heart rate (shortening of the RR interval), or a change in the length of the basal cycle on an intracardiac electrogram. 75.The method of claim 71, wherein the biomarker comprises a rheotropic biomarker. 76.A method according to any one of claims 70 to 75, wherein the catheter comprises a balloon, and wherein the neuromodulatory element is positioned on the balloon, and optionally the balloon catheter is inflated an amount to dilate the vasculature of the subject. 77. A method of treating a heart condition in a subject, the method comprising: positioning a neuromodulatory element on an expansion element in a subclavian artery of the subject; inflating the expansion element; electrically stimulating at least one of a dorsal subclavian haptic or a ventral subclavian haptic; Confirming stimulation of the dorsal subclavian loop and/or the ventral subclavian loop by monitoring cardiac parameters; and providing ablation energy to the dorsal subclavian loop and/or the ventral subclavian loop after confirming stimulation of the dorsal subclavian loop and/or the ventral subclavian loop. 78.A method according to any of the preceding claims, wherein the ablation or other neuromodulation is performed using radiofrequency, ultrasound, microwaves, steam, laser, heat, alcohol or other chemicals, cryogenics, or a combination thereof, and wherein the ablation is reversible or irreversible. 79.A method according to claim 78, wherein the electrical stimulation includes using an electrode combination of the neuromodulatory element, and wherein providing ablation energy includes using the same electrode combination of the neuromodulatory element. 80.A method according to any one of claims 77 to 79, wherein the heart disease comprises atrial fibrillation and/or refractory arrhythmias. 81.A signal generator for generating ablation energy according to any one of the preceding claims, having one or more of the features described in the preceding description. 82.A method for regulating and/or ablating the subclavian haptic, having one or more of the features described in the preceding description. 83.A therapeutic system having one or more of the features described in the preceding description. 84.A tissue treatment system having one or more of the features described in the preceding description. | A system to treat cardiac disease by introducing a catheter with a neuromodulatory element on an expansion element into the subclavian artery and stimulating the dorsal or ventral subclavian loop. The targeted subclavian neuromodulation aims to influence cardiac autonomic tone. | R | "electrically stimulating at least one of a dorsal subclavian loop or a ventral subclavian loop; confirming the stimulation… by monitoring cardiac parameters" (Claim 15) "re-stimulating the subclavian loop after providing the ablation energy, and further providing ablation energy to the subclavian loop if the cardiac parameters confirm the stimulation" (Claim 11) "repeating said restimulation and providing further ablation energy until said cardiac parameters no longer confirm stimulation" (Claim 12) "ablation energy comprises radiofrequency ablation energy… cryogenic ablation energy" (Claims 16–17) The first electrode ablation part performs the ablation operation, the second electrode ablation part is located behind the first electrode ablation part and is waiting for backup, the first temperature sensor monitors the temperature of the first electrode ablation part, and the first pressure sensor detects the impact of the first electrode ablation part on the tissue. Pressure value;" | 1 | [9 empty] | 1 |  |  |  | 1 | [14 empty] | 1 |  | ATHER PARTNERS PLC | ATHER PARTNERS PLC | 2023-06-29 | 2023 | 2025-04-15 | 2025 | 2022-07-01 | 2022 | CN20238061603 | AU2023297473A1¦CA3260367A1¦EP4547321A1¦JP2025521829A¦US12484962B2¦US20250160940A1¦WO2024006466A1
383 | WO2025222080A1 | MEDICAL DEVICES CONFIGURED FOR TEMPERATURE-CONTROLLED THERAPEUTIC ELECTROPORATION | A medical device includes an elongate member and one or more expandable elements positioned on the elongate member, the one or more expandable elements being configured to expand radially outward from the elongate member. The one or more expandable elements have inner surfaces facing the elongate member and outer surfaces opposite the inner surfaces. At least a portion of the one or more expandable elements provide a radiofrequency ablation electrode. A set of temperature sensors is attached to the inner surfaces of at least one of the one or more expandable elements, or on a flexible circuit surrounding at least a portion of the outer surfaces of the one or more expandable elements. | 1. CLAIMSWhat is claimed is: 1. A medical device, comprising: an elongate member; one or more expandable elements positioned on the elongate member, the one or more expandable elements being configured to expand radially outward from the elongate member, the one or more expandable elements having inner surfaces facing the elongate member and outer surfaces opposite the inner surfaces, at least a portion of the one or more expandable elements providing a radiofrequency ablation electrode; and a set of temperature sensors attached to the inner surfaces of at least one of the one or more expandable elements. 2. The medical device of claim 1, wherein the elongate member comprises a delivery catheter. 3. The medical device of claim 1, wherein the one or more expandable elements comprise a wire cage. 4. The medical device of claim 3, wherein the wire cage comprises a plurality of struts, the set of temperature sensors being attached to inner surfaces of at least two different ones of the plurality of struts. 5. The medical device of claim 4 wherein the at least two different ones of the plurality of struts on which the set of temperature sensors are attached are arranged such that when the one or more expandable elements are expanded radially outward from the elongate member, the set of temperature sensors are at two or more different positions of a circumference of the wire cage. 6. The medical device of claim 1, wherein the set of temperature sensors comprise at least four temperature sensors. 7. The medical device of claim 1, wherein the set of temperature sensors comprise thermocouples. 8. The medical device of claim 1, further comprising one or more return electrodes, wherein the radiofrequency ablation electrode and at least one of the one or more return electrodes are configured to perform bipolar radiofrequency ablation. 9. The medical device of claim 8, wherein said at least one of the one or more return electrodes is disposed on the elongate member. 10. The medical device of claim 8, wherein said at least one of the one or more return electrodes is disposed on a guide sheath in which the elongate member is inserted. 11. The medical device of claim 8, wherein said at least one of the one or more return electrodes is disposed on a flexible circuit surrounding at least a portion of the one or more expandable elements. 12. The medical device of claim 1, further comprising a pull wire coupled to a distal tip of the elongate member, the pull wire being configured for retracting the distal tip of the elongate member to expand the one or more expandable elements radially outward from the elongate member. 13. The medical device of claim 1, further comprising a controller configured: to controllably expand the one or more expandable elements to bring the radiofrequency ablation electrode into contact with a wall of a lumen; to monitor a temperature of target tissues in a vicinity of the wall of the lumen utilizing the set of temperature sensors; and to control, based at least in part on the monitored temperature of the target tissues in the vicinity of the wall of the lumen, a current applied to the radiofrequency ablation electrode. 14. The medical device of claim 13, wherein the controller is further configured to monitor a tissue impedance of the target tissues in the vicinity of the wall of the lumen, and wherein the current applied to the radiofrequency ablation electrode is further controlled based at least in part on the monitored tissue impedance. 15. The medical device of claim 1, further comprising a flexible circuit surrounding at least a portion of the one or more expandable elements. 16. The medical device of claim 15, wherein the flexible circuit comprises a plurality of struts, each of at least a subset of the plurality of struts comprising one or more sensing electrodes. 17. The medical device of claim 16, further comprising an isolation layer disposed between the plurality of struts and the one or more expandable elements. 18. The medical device of claim 16, further comprising one or more return electrodes disposed on the flexible circuit. 19. The medical device of claim 16, further comprising a controller configured: to controllably expand the one or more expandable elements to bring the radiofrequency ablation electrode into contact with a wall of a lumen; and to utilize the one or more sensing electrodes of the subset of the plurality of struts of the flexible circuit to determine an extent of ablation of different portions of target tissues in a vicinity of the wall of the lumen, the different portions being positioned at different locations of a circumference of the wall of the lumen. 20. The medical device of claim 19, wherein the controller is further configured to control, based at least in part on the determined extent of the ablation of the different portions of the target tissues, a current applied to the radiofrequency ablation electrode. 21. The medical device of claim 19, wherein the one or more sensing electrodes of the subset of the plurality of struts are configured for selective operation in a therapeutic mode, and wherein the controller is further configured to operate at least one of the one or more sensing electrodes of at least one of the plurality of struts in the therapeutic mode to apply current for radiofrequency ablation of at least one location of the circumference of the wall of the lumen based at least in part on the determined extent of the ablation of the different portions of the target tissues. 22. A medical device comprising: an elongate member; one or more expandable elements positioned on the elongate member, the one or more expandable elements being configured to expand radially outward from the elongate member, the one or more expandable elements having inner surfaces facing the elongate member and outer surfaces opposite the inner surfaces, at least a portion of the one or more expandable elements providing a radiofrequency ablation electrode; and a flexible circuit surrounding at least a portion of the outer surfaces of the one or more expandable elements, the flexible circuit comprising a set of temperature sensors. 23. The medical device of claim 22, wherein the elongate member comprises a delivery catheter. 24. The medical device of claim 22, wherein the one or more expandable elements comprise a wire cage. 25. The medical device of claim 22, wherein the set of temperature sensors comprise at least four temperature sensors. 26. The medical device of claim 22, wherein the set of temperature sensors comprise thermocouples. 27. The medical device of claim 22, further comprising one or more return electrodes, wherein the radiofrequency ablation electrode and at least one of the one or more return electrodes are configured to perform bipolar radiofrequency ablation. 28. The medical device of claim 27, wherein said at least one of the one or more return electrodes is disposed on the elongate member. 29. The medical device of claim 27, wherein said at least one of the one or more return electrodes is disposed on a guide sheath in which the elongate member is inserted. 30. The medical device of claim 27, wherein said at least one of the one or more return electrodes is disposed on the flexible circuit. 31. The medical device of claim 22, further comprising a pull wire coupled to a distal tip of the elongate member, the pull wire being configured for retracting the distal tip of the elongate member to expand the one or more expandable elements radially outward from the elongate member. 32. The medical device of claim 22, further comprising a controller configured: to controllably expand the one or more expandable elements to bring the radiofrequency ablation electrode into contact with a wall of a lumen; to monitor a temperature of target tissues in a vicinity of the wall of the lumen utilizing the set of temperature sensors; and to control, based at least in part on the monitored temperature of the target tissues in the vicinity of the wall of the lumen, a current applied to the radiofrequency ablation electrode. 33. The medical device of claim 32, wherein the controller is further configured to monitor a tissue impedance of the target tissues in the vicinity of the wall of the lumen, and wherein the current applied to the radiofrequency ablation electrode is further controlled based at least in part on the monitored tissue impedance. 34. The medical device of claim 22, wherein the flexible circuit comprises a plurality of struts, each of at least a subset of the plurality of struts comprising one or more sensing electrodes. 35. The medical device of claim 34, further comprising an isolation layer disposed between the plurality of struts and the one or more expandable elements. -SO- 36. The medical device of claim 35 wherein the set of temperature sensors are attached to at least two different ones of the plurality of struts of the flexible circuit such that when the one or more expandable elements are expanded radially outward from the elongate member, the set of temperature sensors are at two or more different positions of a circumference of a wall of a lumen in which the medical device is deployed. 37. The medical device of claim 34, further comprising a controller configured: to controllably expand the one or more expandable elements to bring the radiofrequency ablation electrode into contact with a wall of a lumen; and to utilize the one or more sensing electrodes of the subset of the plurality of struts of the flexible circuit to determine an extent of ablation of different portions of target tissues in a vicinity of the wall of the lumen, the different portions being positioned at different locations of a circumference of the wall of the lumen. 38. The medical device of claim 37, wherein the controller is further configured to control, based at least in part on the determined extent of the ablation of the different portions of the target tissues, a current applied to the radiofrequency ablation electrode. 39. The medical device of claim 37, wherein the one or more sensing electrodes of the subset of the plurality of struts are configured for selective operation in a therapeutic mode, and wherein the controller is further configured to operate at least one of the one or more sensing electrodes of at least one of the plurality of struts in the therapeutic mode to apply current for radiofrequency ablation of at least one location of the circumference of the wall of the lumen based at least in part on the determined extent of the ablation of the different portions of the target tissues. 40. A method, comprising: delivering an elongate member of a medical device to a target site within a lumen; expanding one or more expandable elements positioned on the elongate member radially outward from the elongate member such that outer surfaces of the one or more expandable elements contact a wall of the lumen, at least a portion of the one or more expandable elements providing a radiofrequency ablation electrode; monitoring a temperature of target tissues in a vicinity of the wall of the lumen utilizing a set of temperature sensors attached to inner surfaces of at least one of the one or more expandable elements; and controlling a current applied to the radiofrequency ablation electrode based at least in part on the monitored temperature of the target tissues in the vicinity of the wall of the lumen. 41. The method of claim 40, wherein the one or more expandable elements comprise a wire cage, wherein the wire cage comprises a plurality of struts, and wherein the set of temperature sensors being attached to inner surfaces of at least two different ones of the plurality of struts. 42. The method of claim 41, wherein the at least two different ones of the plurality of struts on which the set of temperature sensors are attached are arranged such that when the one or more expandable elements are expanded radially outward from the elongate member, the set of temperature sensors are at two or more different positions of a circumference of the wire cage. 43. The method of claim 40, wherein expanding the one or more expandable elements radially outward from the elongate member comprises retracting a pull wire coupled to a distal tip of the elongate member. 44. The method of claim 40, further comprising determining an extent of ablation of different potions of target tissues in a vicinity of the wall of the lumen utilizing sensing electrodes disposed on at least a subset of a plurality of struts of a flexible circuit surrounding at least a portion of the one or more expandable elements, the different portions being positioned at different locations of a circumference of the wall of the lumen. 45. The method of claim 44, wherein controlling the current applied to the radiofrequency ablation electrode is further based at least in part on the determined extent of the ablation of the different portions of the target tissues. 46. The method of claim 44, wherein the one or more sensing electrodes of the subset of the plurality of struts are configured for selective operation in a therapeutic mode, and further comprising operating at least one of the one or more sensing electrodes of at least one of the plurality of struts in the therapeutic mode to apply current for radiofrequency ablation of at least one location of the circumference of the wall of the lumen based at least in part on the determined extent of the ablation of the different portions of the target tissues. 47. The method of claim 40, further comprising monitoring a tissue impedance of the target tissues in the vicinity of the wall of the lumen, and wherein the current applied to the radiofrequency ablation electrode is further controlled based at least in part on the monitored tissue impedance. 48. A method comprising: delivering an elongate member of a medical device to a target site within a lumen; expanding one or more expandable elements positioned on the elongate member radially outward from the elongate member such that outer surfaces of the one or more expandable elements contact a wall of the lumen, at least a portion of the one or more expandable elements providing a radiofrequency ablation electrode; monitoring a temperature of target tissues in a vicinity of the wall of the lumen utilizing a set of temperature sensors attached to a flexible circuit surrounding at least a portion of the outer surfaces of the one or more expandable elements; and controlling a current applied to the radiofrequency ablation electrode based at least in part on the monitored temperature of the target tissues in the vicinity of the wall of the lumen. 49. The method of claim 48, wherein the one or more expandable elements comprise a wire cage. 50. The method of claim 48, wherein the flexible circuit comprises a plurality of struts, and wherein the set of temperature sensors are attached to at least two different ones of the plurality of struts of the flexible circuit such that when the one or more expandable elements are expanded radially outward from the elongate member, the set of temperature sensors are at two or more different positions of a circumference of a wall of a lumen in which the medical device is deployed. 51. The method of claim 48, further comprising determining an extent of ablation of different potions of target tissues in a vicinity of the wall of the lumen utilizing sensing electrodes disposed on at least a subset of a plurality of struts of the flexible circuit, the different portions being positioned at different locations of a circumference of the wall of the lumen. 52. The method of claim 51, wherein controlling the current applied to the radiofrequency ablation electrode is further based at least in part on the determined extent of the ablation of the different portions of the target tissues. 53. The method of claim 51, wherein the one or more sensing electrodes of the subset of the plurality of struts are configured for selective operation in a therapeutic mode, and further comprising operating at least one of the one or more sensing electrodes of at least one of the plurality of struts in the therapeutic mode to apply current for radiofrequency ablation of at least one location of the circumference of the wall of the lumen based at least in part on the determined extent of the ablation of the different portions of the target tissues. 54. The method of claim 48, further comprising monitoring a tissue impedance of the target tissues in the vicinity of the wall of the lumen, and wherein the current applied to the radiofrequency ablation electrode is further controlled based at least in part on the monitored tissue impedance. | A medical electroporation device with expandable elements that radially deploy and provide an RF ablation electrode surface, plus temperature sensors. Temperature data lets the device deliver electroporation while controlling unwanted heating. | R | "at least a portion of the one or more expandable elements providing a radiofrequency ablation electrode… set of temperature sensors attached to the inner surfaces" (Claim 1) "utilize the one or more sensing electrodes of the subset of the plurality of struts of the flexible circuit to determine an extent of ablation of different portions of target tissues" (Claim 19) "monitor a tissue impedance of the target tissues in the vicinity of the wall of the lumen, and wherein the current applied to the radiofrequency ablation electrode is further controlled based at least in part on the monitored tissue impedance" (Claim 14) "operate at least one of the one or more sensing electrodes… in the therapeutic mode to apply current for radiofrequency ablation of at least one location of the circumference of the wall of the lumen based at least in part on the determined extent of the ablation" (Claim 21) | 1 |  |  |  | 1 | [5 empty] | 1 | [13 empty] | 1 | [6 empty] | AUTONOMIX MEDICAL INC | AUTONOMIX MEDICAL INC | 2025-04-18 | 2025 | 2025-10-23 | 2025 | 2024-04-19 | 2024 | WO2025US25299 | [NA]
384 | CN118177977A | Intelligent minimally invasive diagnosis and treatment integrated device and method for common-scanning-path wide-field imaging | According to the intelligent minimally invasive diagnosis and treatment integrated device and method for common-scanning-path wide-field imaging, a mechanical arm controls an optical interference tomography module to obtain a real-time optical tomography image in a biological tissue operation, pathological boundary and volume capacity information of a lesion is recognized, and structural information of the lesion tissue is obtained; calculating tissue optical characteristics and diagnosis information; the multi-dimensional controllable intelligent optical minimally invasive diagnosis and treatment module obtains diagnosis and treatment parameters in minimally invasive diagnosis and treatment and constructs a laser tissue ablation biological effect model; the integrated common scanning path diagnosis and treatment mapping module constructs an intelligent diagnosis and treatment mapping relation according to the position information of the lesion tissue to control the corresponding light beams to perform multi-dimensional scanning; the intelligent optical detection and probe control module controls the mechanical arm to control the minimally invasive diagnosis and treatment time sequence of the optical interference tomography module and the photo-thermal diagnosis and treatment and conduct real-time diagnosis and treatment guidance. The device is simple in structure, accurate in precision, easy to operate, high in imaging speed, small in size, portable and high in intelligent degree. | 1· An intelligent minimally invasive diagnosis and treatment integrated device with common scanning path wide-field imaging, characterized in that it includes: a robotic arm controlled optical interference tomography module, used to obtain real-time optical tomography images of biological tissues during surgery, and identify the pathological boundaries and volume capacity information of the lesions from the real-time optical tomography images, obtain the structural information of the lesion tissue, and calculate the optical properties and diagnostic information of the tissue; a multi-dimensional controllable intelligent optical minimally invasive diagnosis and treatment module, used to obtain the diagnosis and treatment parameters in minimally invasive diagnosis and treatment, and pre-construct a laser tissue ablation biological effect model with the structural information of the lesion tissue and laser parameters; an integrated common scanning path diagnosis and treatment mapping module, used to control the corresponding light beam to perform multi-dimensional scanning according to the constructed intelligent minimally invasive diagnosis and treatment mapping relationship based on the position information of the lesion tissue; and an intelligent optical detection and probe control module, used to control the robotic arm to control the optical interference tomography module and the minimally invasive diagnosis and treatment timing of photothermal diagnosis and treatment, and perform real-time guidance of diagnosis and treatment. 2.The device according to claim 1 is characterized in that the robotic arm controls the optical interference tomography module to obtain high-resolution optical diagnostic images for the physical structure of the tissue under a wide-field imaging structure, and the high-resolution optical diagnostic images include depth structure images and processed tissue elasticity images. 3.The device according to claim 1 or 2 is characterized in that the robotic arm-controlled optical interference tomography module includes a robotic arm-controlled two-dimensional galvanometer or MEMS device with a controllable scanning range. 4.The device according to claim 3 is characterized in that the multi-dimensional controllable intelligent optical minimally invasive diagnosis and treatment module is specifically used to control the scanning laser beam through the oblique incident MEMS so that the scanning area coincides with the scanning range of the two-dimensional galvanometer of the optical interference tomography. 5. The device according to claim 1 is characterized in that the calculation formula of the coordinate mapping relationship of the intelligent minimally invasive diagnosis and treatment wide-field imaging is: endu * camera μ * object μ - object μ base11 * endn * camera*1 - base11 end u , oct μ . object μ - object μ base^ ' md ' octn - basen Among them, b is the coordinate transformation matrix from the base of the manipulator to the end of the manipulator, which is obtained by the 5-point calibration of the manipulator, b is the coordinate transformation matrix from the end of the manipulator to the camera, h is the transformation relationship between the calibration plate and the camera, which is obtained by Zhang's calibration, mh is the coordinate transformation matrix from the base of the manipulator to the calibration plate, m is the coordinate transformation matrix from the end of the manipulator to the OCT probe, and a2 is the coordinate transformation matrix from the OCT probe to the calibration plate; The coordinate system relationship between the two systems before and after changing the posture of the manipulator can be written as: endlu * cameralu * object μ _ end2u * earner a2 μ * object only base*1 * endln * cameraln - base*1 * end2n * camera2n endlu * octlu * object μ - end2u * oct2u * object μ basen endln octln - basen end2n oct2n Among them, H is the coordinate conversion matrix from the base of the manipulator to the end of the manipulator before changing the posture of the manipulator, θ is the coordinate conversion matrix from the end of the manipulator to the binocular camera before changing the posture of the manipulator, θ is the coordinate conversion matrix from the calibration plate to the camera before changing the posture of the manipulator, θ is the coordinate conversion matrix from the base of the manipulator to the end of the manipulator after changing the posture of the manipulator, θ is the coordinate conversion matrix from the end of the manipulator to the binocular camera after changing the posture of the manipulator, and θ is the coordinate conversion matrix from the calibration plate to the camera after changing the posture of the manipulator. is the coordinate transformation matrix from the end of the manipulator to the OCT probe before the manipulator posture is changed, is the coordinate transformation matrix from the OCT probe to the calibration plate before the manipulator posture is changed, is the coordinate transformation matrix from the end of the manipulator to the OCT probe after the manipulator posture is changed, is the coordinate transformation matrix from the OCT probe to the calibration plate after the manipulator posture is changed; Transform the equations to obtain the following equations: end2u-l . endlu . cameralu - camera2μ * object object base*1 ' base*1 ' endln - end2n * camera2n cameraln end2 zj-1 * color strong 1 ear * octi u oct2 u * object μ * object μτ-± basen base11 endln - end2n od2n octln The equation simplifies to: A·X=X·B Μ·Y=Y·N Solve two equations, the unknown x=cam or NH, Y=mengH is the hand-eye relationship between the camera and the OCT probe and the end of the robotic arm. To plan the path and control the scanning area of the OCT probe through the information collected by the camera, it is necessary to find the rotation displacement relationship between the camera and the OCT probe, which can be obtained through X and y: camera_ χ . yT 6·An intelligent minimally invasive diagnosis and treatment integrated method with common scanning path wide-field imaging, characterized in that it includes the following steps: obtaining real-time optical tomography images of biological tissues during surgery, and identifying the pathological boundaries and volume capacity information of the lesions from the real-time optical tomography images, obtaining the structural information of the lesion tissues, and calculating the optical properties and diagnostic information of the tissues; obtaining the diagnosis and treatment parameters in minimally invasive diagnosis and treatment, and pre-constructing a laser tissue ablation biological effect model with the structural information of the lesion tissues and laser parameters; controlling the corresponding light beams to perform multi-dimensional scanning based on the constructed intelligent minimally invasive diagnosis and treatment mapping relationship based on the position information of the lesion tissues; and controlling the robotic arm to control the minimally invasive diagnosis and treatment timing of optical interference tomography and photothermal diagnosis and treatment, and performing real-time guidance of diagnosis and treatment. 7.The method according to claim 6 is characterized in that the acquisition of real-time optical tomographic images of biological tissues during surgery includes: acquiring high-resolution optical diagnostic images for the physical structure of the tissue under a wide-field imaging structure, wherein the high-resolution optical diagnostic images include depth structure images and processed tissue elasticity images. 8.The method according to claim 6 or 7 is characterized in that a two-dimensional galvanometer or MEMS device is controlled by a robotic arm with a controllable scanning range to obtain real-time optical tomographic images of biological tissue during surgery. 9.The method according to claim 8 is characterized in that it also includes: controlling the scanning laser beam through an obliquely incident MEMS so that the scanning area coincides with the scanning range of the two-dimensional galvanometer of the optical interference tomography. 10. The method according to claim 6 is characterized in that the calculation formula of the coordinate mapping relationship of the intelligent minimally invasive diagnosis and treatment wide-field imaging is: end μ , camera μ . object Dan_ object basen endn ' cameran ~ base11 end ft,oct ft,obj&ct if - handsome% base ' endn 'octn - basen Among them, a is the coordinate transformation matrix from the base of the manipulator to the end of the manipulator, which is obtained through the 5-point calibration of the manipulator; m is the coordinate transformation matrix from the end of the manipulator to the camera, which is obtained through Zhang's calibration; a is the coordinate transformation matrix from the base of the manipulator to the calibration plate; h is the coordinate transformation matrix from the end of the manipulator to the OCT probe, which is obtained through Zhang's calibration; h is the coordinate transformation matrix from the end of the manipulator to the OCT probe; the coordinate system relationship between the two systems before and after changing the posture of the manipulator can be written as: endlu . earner alii . object μ _ end2 u . camera2 u . object μ basen ' ' cameraln - base11 *2nd2 port*camera2n endl jj octi it object Tj - end2 rj oct2 jj object Tj basen * endln 'octi-base11 * end2n *oct2n Among them, H is the coordinate transformation matrix from the base of the manipulator to the end of the manipulator before the manipulator posture is changed, h is the coordinate transformation matrix from the end of the manipulator to the binocular camera before the manipulator posture is changed, h is the coordinate transformation matrix from the calibration plate to the camera before the manipulator posture is changed, s is the coordinate transformation matrix from the base of the manipulator to the end of the manipulator after the manipulator posture is changed, s is the coordinate transformation matrix from the end of the manipulator to the binocular camera after the manipulator posture is changed, s is the coordinate transformation matrix from the end of the manipulator to the OCT probe before the manipulator posture is changed, s is the coordinate transformation matrix from the OCT probe to the calibration plate before the manipulator posture is changed, s is the coordinate transformation matrix from the end of the manipulator to the OCT probe after the manipulator posture is changed, s is the coordinate transformation matrix from the OCT probe to the calibration plate after the manipulator posture is changed; Transform the equations to obtain the following equations: end2u-i * endl u * camerai u = camera! u * objectobject basen base11 endln endzn camera2n cameraln end2 u-1 * endl u * octi u - oct2 u object μ object rr-± basen basen endln - end2n oct2n octln The equation is simplified to A·X=X·B Μ·Y=Y·N. Solving the two equations, the unknowns x= aminorthH and y=mengH are the hand-eye relationship between the camera and the OCT probe and the end of the robotic arm. To plan the path and control the scanning area of the OCT probe through the information collected by the camera, it is necessary to find the rotational displacement relationship between the camera and the OCT probe. Through X and y, we can get: camerai^ _ χ . yT | An intelligent minimally invasive diagnosis-and-treatment device using a robotic arm controlling an OCT module that scans tissue across a wide field with a common scanning path. OCT identifies pathological boundaries in real time during minimally invasive procedures. | R |  |  | 1 | [11 empty] | 1 | [5 empty] | 1 | 1 | [9 empty] | 1 | BEIJING INSTITUTE TECH | BEIJING INSTITUTE TECH | 2024-02-05 | 2024 | 2024-06-14 | 2024 | 2024-02-05 | 2024 | CN202410166876 | [NA]
385 | CN118177950A | Tumor electrical impedance change analysis and electric pulse parameter optimization method | The invention discloses a tumor electrical impedance change analysis and electric pulse parameter optimization method, which comprises the following steps: S100, applying irreversible electroporation pulse stimulation to a to-be-detected tumor tissue to form nanoscale micropores on the cell surface of the to-be-detected tumor tissue; s200, collecting electrical impedance information of the tumor tissue to be detected; s300, reconstructing an electrical impedance image for the electrical impedance information by using a time domain difference or frequency domain difference algorithm, and quantitatively describing the tumor tissue to be detected; s400, selecting parameters to be optimized from the pulse parameters and the electrode parameters, performing statistical model modeling by adopting a polynomial regression method, and establishing a target function; s500, selecting a target function needing to be optimized according to the optimization target, and establishing an optimization function; and S600, solving the optimization function to obtain an optimal pulse parameter and an optimal electrode parameter. According to the invention, the tumor can be eliminated in the maximum area, and the surrounding normal tissues can be minimally damaged at the same time. | 1A method for analyzing tumor electrical impedance changes and optimizing electrical pulse parameters, comprising: S100, applying irreversible electroporation pulse stimulation to the tumor tissue to be tested, so that nanoscale micropores are formed on the cell surface of the tumor tissue to be tested; S200, collecting electrical impedance information of the tumor tissue to be tested; S300, using a time domain difference or frequency domain difference algorithm on the electrical impedance information to reconstruct an electrical impedance image and quantitatively describe the tumor tissue to be detected; S400, selecting parameters to be optimized from the pulse parameters and the electrode parameters, using a polynomial regression method to perform statistical model building, and establishing an objective function; S500, according to the optimization target, select the objective function to be optimized and establish the optimization function; S600, solving the optimization function to obtain the optimal pulse parameters and electrode parameters. 2.The method for analyzing tumor electrical impedance changes and optimizing electric pulse parameters guided by nanoscale micropores according to claim 1 is characterized in that, in S100, the irreversible electroporation pulse stimulation includes at least one of single pulse stimulation, continuous stimulation, and specific frequency stimulation. 3The method for analyzing electrical impedance changes of tumors and optimizing electrical pulse parameters guided by nanoscale micropores according to claim 1 is characterized in that, in S200, collecting electrical impedance information of the tumor tissue to be tested comprises: S201, prepare the electrical impedance measurement equipment and ensure that the equipment is in normal working condition; S202, setting initial values of pulse parameters according to the size, shape and conductivity distribution of the tumor tissue to be measured; S203, applying irreversible electroporation pulse stimulation to the tumor tissue to be measured by the electrical impedance measuring device, so that current passes through the tumor tissue to be measured, and recording electrical impedance information. 4The method for analyzing tumor electrical impedance changes and optimizing electrical pulse parameters guided by nanoscale micropores according to claim 1 is characterized in that, in S300, the electrical impedance information is subjected to a time domain difference algorithm to reconstruct an electrical impedance image, and the quantitative description of the tumor tissue to be measured includes: S3011, converting the electrical impedance information into an electrical impedance digital signal; S3012, using the collected electrical impedance digital signal as a time series signal V (t); S3013, sampling the sampling points of the time series signal V(t), the number of sampling points is n, and obtaining a discrete time point electrical impedance sampling value sequence V(nΤ), where T is a sampling interval; S3014, calculating a leading differential sequence AV(n) =V((n+1)Τ) -V(nt) for the electrical impedance sampling value sequence V(nT); S3015, performing weighted summation on the leading differential sequence AV (n) to obtain a weighted summation value S (n) = a·Δv(n) + (1-a)·S(nl), where a is a smoothing coefficient used for weighting, and has a value range of 0 to 1; S3016, repeatedly executing the calculation of the leading difference sequence and performing weighted summation on the leading difference sequence, and performing iterative summation until the last sampling point is processed; S3017, reconstruct according to the S (η) values of all sampling points to obtain the electrical impedance image reconstruction result. 5The method for analyzing tumor electrical impedance changes and optimizing electric pulse parameters guided by nanoscale micropores according to claim 1 is characterized in that, in S300, the electrical impedance information is subjected to a frequency domain difference algorithm to reconstruct an electrical impedance image, and the quantitative description of the tumor tissue to be measured includes: S3021, converting the electrical impedance information into an electrical impedance digital signal; S3022, convert the collected electrical impedance digital signal V (n) from the time domain to the frequency domain, perform discrete Fourier transform, and obtain the frequency domain signal V(n) = Σ2/n, where n is the number of sampling points, and the value is repeatedly taken with n1. The value corresponds to each discrete sampling point in the time domain signal, k is the number of frequency units, which are repeated from 0 to N-1, j is the imaginary unit, and e is the Napier constant; S3023, taking the modulus of the frequency domain signal V(k) to obtain the amplitude spectrum A(k) = ¦V(k) I of each frequency unit; S3024, performing a differential operation on the amplitude spectrum to obtain a differential value of the frequency domain signal between adjacent frequency points AA (k) = A(k+1) -A(k); S3025, repeatedly perform the differential operation until deep differential analysis of the frequency domain signal is achieved, and obtain the frequency domain differential result S (k) = aX AA(k) + (la) XS (kl), where a is a smoothing parameter; S30 26, perform inverse Fourier transform on the difference result S (k) to obtain the reconstructed time domain signal S (n) = (1 / M L: F (k) m), double, this group of sampling points is the electrical impedance image reconstruction result. 6.The method for analyzing tumor electrical impedance changes and optimizing electric pulse parameters guided by nanoscale micropores according to claim 3 is characterized in that, in S400, the pulse parameters include at least one of pulse voltage, pulse width, number of pulses and pulse frequency. 7.The method for analyzing tumor electrical impedance changes and optimizing electric pulse parameters guided by nanoscale micropores according to claim 6 is characterized in that, in S400, the electrode parameters include at least one of the electrode needle spacing, the number of electrode needles, the electrode needle length and the electrode needle diameter. 8The method for analyzing the change of tumor electrical impedance and optimizing electric pulse parameters guided by nanoscale micropores according to claim 7 is characterized in that, in S400, the parameters to be optimized are selected from the pulse parameters and the electrode parameters, and the statistical model is built by the polynomial regression method, and the objective function is established, including: S401, selecting a parameter to be optimized from pulse parameters and electrode parameters, wherein the parameter to be optimized includes at least one of pulse voltage U/V, number of electrode needles n, electrode needle spacing D/mm, and electrode needle length L/mm; S402, using the polynomial regression method, according to the form of the multivariate quadratic regression equation, establish the objective function y = ko+k]X1+··· +3[+Jinnai2+·.·+1</;+1<6;+benevolence2kj2+...+One tile high one high, where y is the response value, Xj is the variable of the parameter to be optimized, i is the number of parameters to be optimized, and k is the regression coefficient. 9.The method for analyzing tumor electrical impedance changes and optimizing electric pulse parameters guided by nanoscale micropores according to claim 8 is characterized in that, in S500, according to the optimization target, the objective function to be optimized is selected and the optimization function is established; S501, determining an optimization goal of achieving maximum ablation area with minimum damage to surrounding normal tissues; S502, establish the objective function of the ablation area f(x) =&0+&ιυ+&2N+&3Ο+&4ί+&5υ2+&6Ν2+&7Ο2+&8ί2+&9υϋ+&10α+&1ιυΝ+&12ΝΏ+&13 NL, where a. to ai4 are regression coefficients; S503, establish the objective function of damage degreef2(X)=b0+b1U+b2N+b3D+b4L+b5U2+b6N2+b7D2+b8L2+b9UD+b10UL+b11UN+b12ND+b13 NL, where b0 to s are regression coefficients; S504, establish optimization function Max f 1(X), Min f2 (x). 10. The method for analyzing tumor electrical impedance changes and optimizing electric pulse parameters guided by nanoscale micropores according to claim 9, characterized in that, 5600φ, solving the optimization function to obtain the optimal pulse parameters and electrode parameters includes: S601, limiting the pulse voltage U/V to a range of 1000WUW2500, and the number of electrode needles N to a range of 1WN W5, the electrode needle spacing D/mm ranges from 4WDW13, and the electrode needle length L/mm ranges from 4WLW13; S602, solving the optimization function according to the limited range to obtain the optimal pulse parameters and electrode parameters. | A method that applies IRE pulses to a tumor, measures impedance, and reconstructs an electrical impedance image to analyze the tumor's response. Optimization of pulse parameters is then driven by these tumor impedance changes. | R | "applying irreversible electroporation pulse stimulation to the tumor tissue to be tested, so that nanoscale micropores are formed on the cell surface" (Claim 1) "using a time domain difference or frequency domain difference algorithm on the electrical impedance information to reconstruct an electrical impedance image and quantitatively describe the tumor tissue to be detected" (Claim 1) "establish the objective function of the ablation area f(x)…; establish the objective function of damage degree f2(x)…; Max f1(X), Min f2(x)" (Claim 9) "pulse parameters include at least one of pulse voltage, pulse width, number of pulses and pulse frequency… electrode parameters include at least one of the electrode needle spacing, the number of electrode needles, the electrode needle length and the electrode needle diameter" (Claims 6–7) | 1 | [10 empty] | 1 | [6 empty] | 1 |  | 1 | [7 empty] | 1 |  |  | BEIJING ZHONGLIAN GUOCHENG TECH CO LTD | BEIJING ZHONGLIAN GUOCHENG TECH CO LTD | 2024-03-14 | 2024 | 2024-06-14 | 2024 | 2024-03-14 | 2024 | CN202410295411 | [NA]
386 | EP4321091A1 | VISUALIZING AND CLUSTERING MULTIPLE ELECTRODES OF A HIGH-DEFINITION CATHETER PROJECTED ON TISSUE SURFACE | A system includes a display and a processor. The processor is configured to: (i) estimate, based on signals, multiple regions of tissue on a surface of an organ, which are affected by multiple respective electrodes positioned in the organ, (ii) calculate, for at least a subset of the electrodes, a unified region of the tissue affected by the subset, and (iii) display a mark indicative of the unified region on the display. | 1. A system (20), comprising: a display (35); and a catheter end effector (40) having a plurality of electrodes disposed thereon to sense signals; a processor (33) electrically connected to the display and the plurality of electrodes, which is configured to estimate, based on the signals, multiple regions (55a, 55b, 55c) of tissue on a surface of an organ (26), which are affected by multiple respective electrodes (50) positioned in the organ (26); calculate, for at least a subset of the electrodes (50), a unified region (99) of the tissue affected by the subset; and display a mark (77, 88) indicative of the unified region (99) on the display (35). 2. The system according to claim 1, wherein the signals comprise position signals indicative of respective positions of the multiple respective electrodes, and wherein the processor is configured to calculate one or more distances between a surface of the tissue and one or more of the electrodes of the subset, respectively. 3. The system according to claim 2, wherein the processor is configured to display the mark, which is indicative of the distance between the unified region and the subset of the electrodes. 4. The system according to claim 3, wherein the position signals comprise first and second position signals indicative of first and second positions of first and second subsets of the electrodes, respectively, and wherein, based on the first and second position signals, the processor is configured to calculate: (i) a first distance between the first subset and a first surface of a first unified region, and (ii) a second distance between the second subset and a second surface of a second unified region, different from the first surface of the first unified region. 5. The system according to claim 4, wherein the first distance is different from the second distance, and wherein the processor is configured to display: (i) over a first map of the first unified region, a first mark, which is indicative of the first distance, and (ii) over a second map of the second unified region, a second mark, which is different from the first mark and is indicative of the second distance. 6. The system according to claim 5, wherein the processor is configured to display a polygon surrounding at least the first map and the second map, and wherein an outline of the polygon comprises the first and second marks. 7. A method, comprising: estimating, based on signals from a plurality of electrodes, multiple regions (55a, 55b, 55c) of tissue on a surface of an organ (26), which are affected by multiple respective electrodes (50) positioned in the organ (26); calculating, for at least a subset of the plurality of electrodes (50), a unified region (99) of the tissue affected by the subset; and displaying a mark (77, 88) indicative of the unified region (99). 8. The method according to claim 7, wherein the signals comprise position signals indicative of respective positions of the multiple respective electrodes, and wherein calculating the unified region comprises calculating one or more distances between a surface of the tissue and one or more of the electrodes of the subset, respectively. 9. The method according to claim 8, wherein the mark is indicative of the distance between the unified region and the subset of the electrodes. 10. The method according to claim 9, wherein the position signals comprise first and second position signals indicative of first and second positions of first and second subsets of the electrodes, respectively, and wherein, based on the first and second position signals, calculating the unified region comprises calculating: (i) a first distance between the first subset and a first surface of a first unified region, and (ii) a second distance between the second subset and a second surface of a second unified region, different from the first surface of the first unified region. 11. The method according to claim 10, wherein the first distance is different from the second distance, and wherein displaying the mark comprises displaying: (i) over a first map of the first unified region, a first mark, which is indicative of the first distance, and (ii) over a second map of the second unified region, a second mark, which is different from the first mark and is indicative of the second distance. 12. The method according to claim 11, wherein displaying the first and second marks comprises displaying a polygon surrounding at least the first map and the second map, and wherein an outline of the polygon comprises the first and second marks. 13. The system according to claim 6 or the method according to claim 12, wherein the first mark comprises a first line, which is a first section of the outline of the polygon, and the second mark comprises a second line, which is a second section of the outline of the polygon. 14. The system or the method according to claim 13, wherein the first line has a first color, and the second line has a second color, different from the first color. 15. The system or the method according to claim 13, wherein the first line comprises a solid line, and the second line comprises a broken line. 16. The method according to any of claims 7 to 15, wherein the signals are indicative of one or more parameters of an ablation of the tissue, and wherein: (i) calculating the unified region comprises calculating, based on the signals, a size of a lesion intended to be formed by the subset in the unified region, and (ii) displaying the mark comprises displaying, over the uniform region, the mark, which is indicative of the size of the lesion. 17. The system according to any of claims 1 to 6 and 13 to 15, wherein the signals are indicative of one or more parameters of an ablation of the tissue, and wherein the processor is configured to: (i) calculate, based on the signals, a size of a lesion intended to be formed by the subset in the unified region, and (ii) display, over the uniform region, the mark, which is indicative of the size of the lesion. | A system that estimates which tissue regions each electrode of a high-density catheter is affecting, then computes a unified region for a chosen subset of electrodes and draws it on the display. Operators see at a glance which combined area is actually being treated. | R | "estimate, based on the signals, multiple regions of tissue on a surface of an organ, which are affected by multiple respective electrodes" (Claim 1) "calculate, for at least a subset of the electrodes, a unified region of the tissue affected by the subset" (Claim 1) "signals are indicative of one or more parameters of an ablation of the tissue… calculate, based on the signals, a size of a lesion intended to be formed by the subset in the unified region" (Claims 16–17) "calculate one or more distances between a surface of the tissue and one or more of the electrodes of the subset" (Claim 2) | 1 | [6 empty] | 1 | [11 empty] | 1 |  |  |  |  | 1 | [6 empty] | BIOSENSE WEBSTER ISRAEL LTD | BIOSENSE WEBSTER ISRAEL LTD | 2023-08-09 | 2023 | 2024-02-14 | 2024 | 2022-08-10 | 2022 | EP20230190529 | CN117582282A¦IL304864A¦JP2024025729A¦US20240050017A1
387 | CN119131797A | Cardiac ultrasonic intelligent recognition system and method for myocardial ablation | The invention discloses a cardiac ultrasound intelligent recognition system for myocardial ablation, which relates to the technical field of ultrasound medicine, and comprises a cardiac echocardiogram acquisition module, a neural network construction module, a neural network training module, a neural network recognition module and an image processing module, and is used for assisting in setting ablation parameters in an ablation process. Comprise output power, time, electrode position and the like. Wherein the cardiac echocardiogram acquisition module comprises an ultrasonic diagnosis device and an ablation instrument, compared with a traditional separated device, the cardiac echocardiogram intelligent identification system is higher in integration degree and more convenient to use, in addition, the invention further discloses a cardiac ultrasound intelligent identification method for myocardial ablation, an echocardiogram two-dimensional dynamic image of a detected object is acquired, and the detection accuracy is improved. The two-dimensional dynamic image of the echocardiogram of the detected object is input into the convolutional neural network model, the features of each section of the ultrasound image are automatically recognized and positioned, and the problem that the score deviation is large due to too strong subjective factors can be avoided by adopting the method. | 1.A cardiac ultrasound intelligent recognition system for myocardial ablation, characterized by: comprising a cardiac ultrasound acquisition module, for collecting and acquiring a two-dimensional dynamic image of the ultrasound cardiogram of the object to be detected; a neural network construction module, for constructing a convolutional neural network model; a neural network training module, the neural network training module is used to train the convolutional neural network model; a neural network recognition module, using the trained convolutional neural network model to process and recognize the two-dimensional dynamic image of the ultrasound cardiogram, and automatically screen out the end-diastolic frame image of each standard section; an image processing module, for performing image processing on each recognized frame image, automatically identifying, locating, and marking the mitral valve, tricuspid valve, aortic valve, and left and right atria and chambers on the image, and measuring the inner diameter of the left and right atria and chambers and the thickness of each chamber wall on the standard section; an image quality evaluation module, for evaluating and analyzing the input two-dimensional dynamic image of the ultrasound cardiogram, and performing quantitative scoring according to whether it contains the end-diastolic frame image of the standard section required for cardiac ultrasound examination, the clarity of the image, and the fit of the measured data; The cardiac ultrasound cardiogram acquisition module includes an ultrasonic diagnostic device and an ablation device. The lower end of the ultrasonic diagnostic device is fixedly connected to a plurality of vertical poles, and the upper ends of the vertical poles are fixedly connected to a circular track. The ablation device is arranged on a mounting mechanism; the mounting mechanism includes a placement plate for mounting the ablation device, a first drive component for driving the placement plate to rotate, and a second drive component for driving the first drive component to slide on the circular track. 2.According to claim 1, a cardiac ultrasound intelligent recognition system for myocardial ablation is characterized in that: the neural network training module specifically includes the following units: a training sample acquisition unit, used to acquire training samples; a neural network training unit, used to input the training set in the training sample into the initial convolutional neural network to obtain an initial convolutional neural network model; an output result acquisition unit, used to input the verification set in the training sample into the initial convolutional neural network model for processing to obtain an output result; a result comparison unit, used to compare the output result with the verification set image label to obtain a comparison result; a parameter optimization unit, used to optimize the parameters of the initial convolutional neural network model using the obtained comparison result to obtain a trained convolutional neural network model. 3.According to claim 1, a cardiac ultrasound intelligent recognition system for myocardial ablation is characterized in that: the neural network construction module specifically includes the following units: a convolution layer construction unit, which applies a preset convolution kernel according to the received echocardiographic data, performs convolution calculation according to the convolution step size, and obtains a convolution result; a pooling layer construction unit, which performs aggregation and statistical processing on the received convolution result to obtain a pooling result; a fully connected layer construction unit, which integrates the pooling results of the received local information to obtain global information. 4.According to claim 1, a cardiac ultrasound intelligent recognition system for myocardial ablation is characterized in that: the image processing module specifically includes the following units: an image preprocessing unit, used to preprocess the ultrasound cardiogram obtained from the object to be detected to obtain a preprocessed image; an image recognition unit, used to identify and extract corresponding cardiac standard section frames from the preprocessed image according to the image features of each standard section of the heart; an image structure positioning unit, used to determine the positions of the mitral valve, tricuspid valve, aorta, left and right atrioventricular cavities of various cardiac structures in the preprocessed image according to the image features of each standard section of the heart, and calibrate the corresponding structures; The image structure measurement unit is used to measure the inner diameters of the left and right atria and ventricles and the thickness of the ventricular septum of each structure of the heart in the preprocessed image according to the image features of each standard section of the heart. 5.A cardiac ultrasound intelligent recognition system for myocardial ablation according to claim 1 is characterized in that: the second driving component includes an arc-shaped slider, a mounting plate, a plug-in plate plugged into the side of the arc-shaped slider, and a second driving module for driving the arc-shaped slider to slide, the lower end of the arc-shaped slider is provided with a slot that can be plugged into the circular track, the side of the circular track is provided with a slide groove for plugging the plug-in plate, the upper half of the arc-shaped slider is provided with a through hole, the mounting plate is slidably arranged in the through hole along the up and down directions, and the driving module is arranged at the lower end of the mounting plate located in the circular track. 6.According to claim 5, a cardiac ultrasound intelligent recognition system for myocardial ablation is characterized in that: a fixing rod is fixedly connected to the lower end of the mounting plate outside the annular track, a through groove for plugging the fixing rod is provided at the end of the plug plate away from the slide groove, the lower end of the fixing rod is an inverted right-angled trapezoid, the upper end of the mounting plate is fixedly connected to a tension spring, and the upper end of the tension spring is fixedly connected to the upper end of the arc-shaped slider. 7.A cardiac ultrasound intelligent recognition system for myocardial ablation according to claim 6, characterized in that: the second driving module includes a first motor fixedly connected to the lower end of the mounting plate, a driving gear fixedly connected to the output shaft of the first motor, and an inner gear ring fixedly connected to the inner side of the annular track, and the inner gear ring and the driving gear are meshed and connected; when the fixing rod is temporarily fixed to the mounting plate, the inner gear ring and the driving gear are meshed and connected, and when not temporarily fixed, the inner gear ring and the driving gear are not meshed. 8.A cardiac ultrasound intelligent recognition system for myocardial ablation according to claim 7, characterized in that: the first driving assembly includes a circular plate, a rectangular groove provided on the circular plate, a telescopic rod fixed in the rectangular groove, a second motor fixed in the arc-shaped slider, an inner polygonal sleeve rotatably connected to the middle of the circular plate, and a fan-shaped gear fixedly connected to the outer side of the inner polygonal sleeve; the placement plate is provided with a circular hole for matching the installation of the circular plate, and the placement plate at the upper end of the circular hole is provided with an annular tooth groove for engaging the fan-shaped gear; the upper end of the mounting plate is fixedly connected with an insertion rod that can be plugged into the rectangular groove, and the insertion rod is provided with a limiting groove for the insertion of the telescopic rod; the output shaft of the second motor is fixedly connected with a polygonal shaft for matching the installation of the inner polygonal sleeve; the upper end of the arc-shaped slider is provided with a mounting groove for mounting the second motor; the placement plate at the lower end of the circular hole is fixedly connected with a lower ring plate, and the upper end of the outer surface of the circular plate is fixedly connected with an upper ring plate, and the upper ring plate is used to limit the lower ring plate on the arc-shaped slider; the placement plate is fixed with a bandage for mounting the ablation device. 9. A cardiac ultrasound intelligent recognition method for myocardial ablation, comprising an intelligent recognition system for cardiac ultrasound examination according to any one of claims 1 to 8, characterized in that it comprises the following steps: Step S1, move the ultrasonic diagnostic device to the treatment position, use the ultrasonic diagnostic device to determine the position of the lesion, and then control the first motor and the second motor to adjust the position of the ablation device, and plan the path of the ablation needle or catheter; Step S2, disinfect and locally anesthetize the puncture site; Step S3, under ultrasonic guidance, accurately insert the ablation needle or catheter into the lesion area, start the ablation device, set the ablation parameters according to the treatment plan, start the ablation process, and continuously use the ultrasonic diagnostic device for monitoring during the entire ablation process to ensure that the position of the ablation needle or catheter is correct, while monitoring changes in tissue temperature; Step S4, obtain a two-dimensional dynamic image of the ultrasonic cardiogram of the detection object; Step S5, input the two-dimensional dynamic image of the ultrasonic cardiogram of the detection object into a convolutional neural network model, and automatically recognize The ultrasonic image features of each segment are located separately, the end-diastolic image frames of each standard section are captured, the mitral valve, tricuspid valve, aortic valve, left and right atrioventricular cavities are marked, and the thickness of each atrioventricular cavity and ventricular wall is measured and automatically marked. | A cardiac-ultrasound intelligent recognition system for myocardial ablation, with modules to acquire echocardiograms, build/train/run a neural network, and process images. AI-recognized cardiac structures help set ablation parameters more accurately. | R | "uses the trained convolutional neural network model to process and recognize the two-dimensional dynamic image of the ultrasound cardiogram, and automatically screen out the end-diastolic frame image of each standard section" (Claim 1) "performing image processing on each recognized frame image, automatically identifying, locating, and marking the mitral valve, tricuspid valve, aortic valve, and left and right atria and chambers… measuring the inner diameter of the left and right atria and chambers and the thickness of each chamber wall" (Claim 1) "start the ablation process, and continuously use the ultrasonic diagnostic device for monitoring during the entire ablation process to ensure that the position of the ablation needle or catheter is correct, while monitoring changes in tissue temperature" (Claim 9, Step S3) |  |  | 1 |  | 1 | [12 empty] | 1 |  |  |  | 1 |  |  | 1 | [6 empty] | FIRST AFFILIATED HOSPITAL ZHENGZHOU UNIV | FIRST AFFILIATED HOSPITAL ZHENGZHOU UNIV | 2024-09-03 | 2024 | 2024-12-13 | 2024 | 2024-09-03 | 2024 | CN202411226495 | [NA]
388 | CN121129419A | Radio frequency ablation electrode control method | The invention discloses a radio frequency ablation electrode control method which comprises the following steps: S1, initializing an ablation temperature target value, a safety limit value and a power adjusting range, and loading a pre-stored anal fistula ablation three-stage feature library and impedance-temperature correlation degree thresholds of all stages; s2, collecting the electrode temperature, the electrode-tissue impedance and the ablation duration in real time, and judging the current ablation stage according to the matching relationship between the ablation duration and the three-stage feature library; s3, formulating a state judgment rule according to the current stage; s4, a fuzzy PID algorithm is adopted to calculate a power basic adjusting amount, a power dynamic compensation amount is calculated, and final output power is determined; and S5, repeating the steps S2-S4 to realize closed-loop control, triggering first-level early warning when the parameters are close to the boundary, triggering second-level alarm when the parameters exceed the safety limit value, and recording abnormal data. The stage adaptability is high, the problem of temperature adjustment lag can be solved, the adjustment precision is accurate enough, and the ablation effect can be improved. | 1 A method for controlling a radiofrequency ablation electrode, characterized by comprising the following steps: S1. Initialize the target ablation temperature value, safety limit value, and power adjustment range; load the pre-stored anal ablation three-stage feature library and impedance-temperature correlation thresholds for each stage. S2. Real-time acquisition of electrode temperature, electrode-tissue impedance and ablation time; determination of the current ablation stage based on the matching relationship between ablation time and the three-stage feature library. S3. Develop state determination rules according to the current stage: determine the impedance range and temperature response matching during the initial contact period, determine the impedance-temperature correlation and parameter fluctuation during the stable period, predict the impedance rise rate and temperature rise risk during the drying period, and determine the energy output state. S4. Use the fuzzy PID algorithm to calculate the basic power adjustment amount and the dynamic power compensation amount to determine the final output power. S5. Repeat S2-S4 to achieve closed-loop control. When the parameter approaches the boundary, a first-level warning is triggered. When the parameter exceeds the safety limit, a second-level alarm is triggered, and abnormal data is recorded. 2 According to claim 1, a radiofrequency ablation electrode control method is characterized in that: in S1, the three-stage feature library for anal ablation includes: initial contact period: duration 0-5s, impedance decrease rate > 30 Ω/s, temperature rise rate v > 5°C/s; stable ablation period: duration 5s to 10s before the end of the procedure, impedance fluctuation amplitude v > 20 Ω/s, temperature fluctuation amplitude v > 0.5°C/s; tissue drying period: duration greater than 10s before the end of the procedure, impedance rise rate > 10 Ω/s, temperature rise rate > 0.2°C/s; the impedance-temperature correlation threshold for the stable ablation period is set to 0.6 and 0.8, and the impedance-temperature correlation threshold for the tissue drying period is set to 0.4-0.6, and the correlation is calculated using the Pearson correlation coefficient. 3 According to claim 2, a radiofrequency ablation electrode control method is characterized in that: in S2, the electrode temperature data is acquired in real time as temperature T by a thermocouple temperature sensor attached to the tip of the electrode; Electrode-tissue impedance data is acquired in real-time using a 0.01Ω high-precision sampling resistor connected in series with a 24-bit ADC chip. The real-time impedance Z is calculated as: U<sub>sample</sub> = U<sub>sample</sub> × R<sub>sample</sub>, where U<sub>sample</sub> is the voltage across the sampling resistor acquired by the ADC chip, R<sub>sample</sub> = 0.01Ω, which is the high-precision sampling resistor of the RF circuit, and Ur is the total voltage of the RF circuit. Duration data is recorded by a built-in timer in the sampling system to record the ablation duration. 4 According to claim 3, the radiofrequency ablation electrode control method is characterized in that: in S3, the judgment logic for each stage is as follows: Initial contact period: The core judgment is whether the impedance change trend matches the temperature response. If the impedance has dropped to the range of 100Ω~2000Ω but the temperature is still >64°C, it is judged as "insufficient energy"; if the impedance has not dropped to the target range but the temperature is >66°C, it is judged as "local overheating"; if both are within the initial adaptation range, with impedance of 100Ω~2000Ω and temperature of 64°C~66°C, it is judged as "normal"; Stable ablation period: The focus is on monitoring the impedance-temperature correlation and parameter fluctuation. If the correlation is in the range of 0.6~0.8 and the real-time temperature T is in the range of 64.5°C~65.5°C and the real-time impedance Z is in the range of 100Ω~1000Ω. If the correlation coefficient v is greater than 0.6 and the real-time temperature Tv is 64.5°C, it is judged as "normal"; if the correlation coefficient v is greater than 0.8 and the real-time temperature Tv is 65.5°C, it is judged as "energy excess"; During the tissue drying period: Prioritize the risk of a sudden temperature rise; if the real-time impedance rise rate is > 10 Q/s or the real-time temperature is above 7... If the rise rate is >0.5°C/s, it is judged as "predicted overheating"; if the real-time impedance Z rises but the real-time temperature TV is 64°C, it is judged as "insufficient energy"; if the real-time temperature T is between 64.5°C and 65.5°C and the real-time impedance Z rise rate is W80Q/s, it is judged as "normal". 5 According to claim 3, the radiofrequency ablation electrode control method is characterized in that: in S4, the fuzzy PID algorithm is used to calculate the power base adjustment specifically as follows: Input variable quantization: the actual temperature deviation AT = TT target and the impedance deviation AZ = ZZ stage target are converted into fuzzy universes of discourse, and the quantization formulas are: X = AT/ATmax * n, y = AZ/AZmax xn, where ATmax=5°C, Zmax = 100Q /=3, the calculation result is rounded to the nearest integer; Fuzzy set partitioning: AZ and the power base adjustment quantity API are both divided into 7 fuzzy sets: "negative large, negative medium, negative small, zero, positive small, positive medium, and positive large". The membership function of each fuzzy set adopts a triangular distribution; Fuzzy inference: The smaller value method is used to calculate the output membership degree of each rule. The rule base contains 49 core rules. The formula for calculating the output membership degree of a single rule is: Rule i(APl) = min[gAT(Ai),UAZ(Bi)]; where, i and Bi are the fuzzy sets of AT and AZ, respectively; Defuzzification: The output membership degrees of all rules are combined to calculate the precise value of P1. The formula is:Lri - £Lu1 μruleQPl) X PL d = :ΣChangμruleQP1); where Pli is the center value of the API corresponding to the i-th rule. 6 According to claim 5, a radiofrequency ablation electrode control method is characterized in that: power dynamic compensation is calculated: the compensation amount AP2 is used to offset the effects of electrode wear and tissue thermal conductivity differences, and the calculation formula is: AP2 = Ktxt + KWithx ¦ΔΖ¦; Where Kt=0.02W/s is the time compensation coefficient, Kz=0.01W/Q is the impedance deviation compensation coefficient, and t is the ablation time in seconds. 7 According to claim 6, a radiofrequency ablation electrode control method is characterized in that: determining the final output power specifically involves: Preliminary calculation: P<sub>preliminary</sub> = P<sub>0</sub> + API + AP<sub>2</sub>, where P<sub>0</sub> is the current output power; Single adjustment amplitude constraint: the upper limit of single power adjustment P<sub>upper limit</sub> = P<sub>0</sub> x 20%, and the lower limit p<sub>lower limit</sub> = -P<sub>0</sub> x 20%; if API + Δp<sub>2</sub> > p<sub>upper limit</sub>, then p<sub>constraint</sub> = P<sub>0</sub> + P<sub>upper limit</sub>; if ΔP<sub>1</sub> + ΔP<sub>2</sub> < P<sub>upper limit</sub>, then P<sub>constraint</sub> = P<sub>0</sub> + P<sub>lower limit</sub>; otherwise, P<sub>constraint</sub> = P<sub>preliminary</sub>; Stage power upper limit constraint: the maximum allowable power for each stage is 20W for the initial contact period, 25W for the stabilization period, and 15W for the drying period. If P<sub>constraint</sub> > P<sub>stage max</sub>, then the final output power P = P<sub>stage max</sub>; if P<sub>constraint</sub> < 5W, then the final output power P = 5W; otherwise, the final output power P = P<sub>constraint</sub>. 8 According to claim 5, a radiofrequency ablation electrode control method is characterized in that: in S5, the first-level warning is triggered when the real-time temperature T approaches the stage temperature boundary or the real-time impedance Z approaches the stage impedance boundary, prompting the doctor to pay attention; the second-level alarm is triggered immediately when the real-time temperature T exceeds the safety limit or the real-time impedance Z exceeds the safety limit, and the power is automatically reduced to the safety value, and the real-time temperature T, real-time impedance Z, and final output power P data at the abnormal moment are recorded for postoperative analysis. | An RF ablation electrode controller specialized for anal-fistula ablation: it loads a three-stage feature library and impedance–temperature thresholds, then monitors temperature, impedance, and time live. The system advances through the three stages based on measured signals. | R | "Real-time acquisition of electrode temperature, electrode-tissue impedance and ablation time; determination of the current ablation stage based on the matching relationship between ablation time and the three-stage feature library" (Claim 1, S2) "impedance-temperature correlation threshold for the stable ablation period is set to 0.6 and 0.8… correlation is calculated using the Pearson correlation coefficient" (Claim 2) "Initial contact period: The core judgment is whether the impedance change trend matches the temperature response… 'insufficient energy'… 'local overheating'… 'normal'" (Claim 4) "Use the fuzzy PID algorithm to calculate the basic power adjustment amount and the dynamic power compensation amount to determine the final output power" (Claim 1, S4) | 1 |  |  |  | 1 | [5 empty] | 1 | [9 empty] | 1 | [6 empty] | 1 |  |  |  | FIRST HOSPITAL JILIN UNIV | FIRST HOSPITAL JILIN UNIV | 2025-11-17 | 2025 | 2025-12-16 | 2025 | 2025-11-17 | 2025 | CN202511677524 | [NA]
389 | WO2025046567A1 | A SYSTEM AND METHOD FOR ELECTROPORATION | A system and method are disclosed to position a 2D array active elements on a 2D array of transperineal stylets in a tissue containing a tumor under guidance by imaging technologies, for evaluating, monitoring, and treating tumors by electroporation. | 1. CLAIMSWhat is claimed is: 1. A system for insertion of an array of stylets into tissue of a subject, comprising: a cartridge housing a plurality of stylets each of said stylets including a pointed distal end and a proximal portion connected to the cartridge; and a first template with a plurality of channels through which a distal portion of the stylets pass into the tissue. 2. The system of claim 1, further comprising an actuator configured to move said stylets longitudinally with respect to said first template thereby moving said distal portion of said plurality of stylets longitudinally through said channels of said first template. 3. The system of claim 2, wherein the actuator is further configured to move said stylets longitudinally with respect to said cartridge. 4. The system according to claim 2, further comprising an x-y actuator for positioning a pusher against a selected stylet of said plurality of stylets and wherein said actuator pushes the selected stylet to a desired depth. 5. The system according to claim 4, further wherein at least one of said x-y actuator and a second actuator is configured for detaching the pusher from the selected stylet. 6. The system according to claim 1, further comprising a second template, said first template in contact with the subject and the second template positioned between the first template and the cartridge, wherein the second template is connected to at least a portion of said stylets and wherein said second template is moved longitudinally by an actuator between said first template thereby moving said portion of the stylets longitudinally. 7. The system of claim 6 wherein said second template is configured for retracting said stylets for a predetermined distance. 8. The system according to claim 7, further comprising a processor running an algorithm configured for said retracting and for positioning active electrodes to a next more superficial depth of a procedure. 9. The system according to claim 7, wherein said retracting is for a distance between 5 mm to 30 mm. 10. The system according to claim 1, wherein each stylet is configured to be extended and retracted individually through the first template. 11. The system according to claim 1, comprising an operator interface providing a real time image of the tissue, wherein the operator interface is configured for the operator marking on the real time image a region of a procedure. 12. The system of claim 11, configured to perform a procedure in two disconnected regions simultaneously. 13. The system according to claim 11, wherein the marking is performed at various depths. 14. The system according to claim 13, further comprising a processor running an algorithm to activate to activate only stylets of said plurality of stylets in the region of a procedure at each of said various depths. 15. The system according to claim 11, wherein an electronic grid is superimposed on the real time image and wherein of the grid, coincide with expected positions of the stylets after being pushed into tissue. 16. The system according to claim 11, further comprising a processor running an algorithm to activate a remote actuator that will push each stylet of the plurality of stylets into the tissue at a desired depth and within boundaries of the region of a procedure. 17. The system according to claim 1, further comprising a processor running an algorithm to pass current from a first subset of said stylets to a second subset of said stylets that are intermingled with each other. 18. The system according to claim 17, wherein the first subset of said stylets are dispersed at equal distance from stylets of the second subset. 19. The system according to claim 17, wherein the first subset of said stylets and the second subset of said stylets are distributed and intercalated according to a pattern chosen from parallel rows, concentric pattern and other such equidistant patterns. 20. The system according to claim 19, wherein the pattern is predetermined to facilitate uniformity of current distribution within an area of treatment. 21. The system according to claim 19, wherein the first subset of said stylets and the second subset of said stylets are intercalated along a long axis of a shape to be ablated. 22. The system according to claim 1, further comprising a bioimpedance module that is connected to the stylets fortesting impedance between electrodes on the stylets. 23. The system according to claim 22, wherein the impedance testing is for differentiating between different types of tissue. 24. The system of claim 23, wherein said different types of tissue includes benign and malignant tissue. 25. The system according to claim 22, wherein the impedance testing is for evaluation of a performance and effect of a treatment treated tissue. 26. The system according to claim 1, further comprising an imaging modality and wherein the imaging modality and the system are configured for independent positioning. 27. The system according to claim 1, further comprising an imaging modality and wherein the imaging modality and the system are configured for synchronized positioning. 28. The system according to claim 14, further comprising an imaging modality and wherein an actuator is configured to facilitate rotation of the imaging modality for imaging stylet positioning. 29. The system according to claim 28, wherein an angle of rotation is determined by a predetermined position of the stylet in the cartridge relative to the imaging modality. 30. The system according to claim 29, wherein the imaging modality and the system are configured to be held in pre -determined relative positions. 31. The system of claim 1, wherein the at least one of the cartridge and the stylet includes an individual actuator for longitudinal movement of an individual stylet. 32. The system according to claim 1, wherein each stylet includes a lumen and wherein the lumen is configured to permit insertion of active electrodes. 33. The system according to claim 1, wherein each stylet is at least partially covered by a covering including at least one of a sheath and a coating. 34. The system according to claim 33, wherein the covering is at least one of hydrophobic, hydrophilic, electrically insulating, thermally insulating, and friction reducing. 35. The system according to claim 1, wherein each of said stylets includes an active electrode of /i cm to 3 cm length. 36. The system according to claim 33, wherein the sheath is removable or retractable. 37. The system according to claim 1, wherein each of said stylets is provided with at least two sensing electrodes for bioimpedance measurements. 38. The system according to claim 1, wherein further comprising activatable visual markers on the cartridge. 39. The system according to claim 38, where the activable markers are include one or more of LEDs, projected bright dots, and bright project lines. 40. The system according to claim 39, wherein the activatable markers are produced by a projector situated behind the cartridge. 41. The system according to claim 39, further comprising a processor running an algorithm for activating the activatable markings corresponding to a proximal end of the stylets that were chosen to be inserted into tissue for easy identification. 42. The system of claim 1, wherein the stylets are configured for a procedure selected from the group consisting of: irreversible electroporation (IRE), reversible electroporation (RE), instillation of chemotherapeutics, calcium solution or gene therapy, radiofrequency ablation, microwave ablation, laser ablation, photodynamic therapy, bioimpedance measurement and a combination thereof. 43. The system according to claim 1, wherein the template is configured to provide a variety of densities and patterns of the plurality of stylets. 44. A method for electro-stimulation of tissue of a subject, the method comprising: attaching a cartridge with an array of stylets to a cradle of an imaging modality, wherein the cartridge is configured for parallel insertion of stylets into the tissue; introducing the stylets into the tissue to a desired depth, under monitoring by the imaging modality; selecting a desired area; executing a procedure on the selected area; retracting the stylets to a more superficial transverse plane and repeating the previous steps, as needed. 45. The method according to claim 44, wherein the retracting is accomplished manually, automatically, robotically, remotely or any combination thereof. 46. The method according to claim 45, wherein the retracting is in a longitudinal direction and wherein retracting stops periodically for a procedure on a new layer in the transverse and lateral directions. 47. The method according to claim 44, wherein the area is selected manually or automatically. 48. The method according to claim 44, wherein the area is pre-designated. 49. The method according to claim 44, wherein the procedure is selected from a group including: contact radiofrequency energy, non-contact radiofrequency energy, electroporation, ultrasonic energy, laser energy, gamma radiation, beta radiation, alpha radiation, immunotherapy, or a combination thereof. 50. The method according to claim 44, wherein multiple stylets are configured for acting simultaneously. | A 2D-array electroporation system using transperineal stylets carrying active elements, guided by imaging to position them in tumor-bearing tissue. The 2D array allows mapping, monitoring, and electroporation treatment of solid tumors. | R | "each of said stylets is provided with at least two sensing electrodes for bioimpedance measurements" (Claim 37) "a bioimpedance module that is connected to the stylets for testing impedance between electrodes on the stylets" (Claim 22) "the impedance testing is for differentiating between different types of tissue… different types of tissue includes benign and malignant tissue" (Claims 23–24) "the impedance testing is for evaluation of a performance and effect of a treatment treated tissue" (Claim 25) | 1 | [10 empty] | 1 | [16 empty] | 1 |  |  | FOCAL MEDICAL TECH LTD | FOCAL MEDICAL TECH LTD | 2024-08-21 | 2024 | 2025-03-06 | 2025 | 2023-08-29 | 2023 | WO2024IL50840 | CN121712553A
390 | WO2022093848A1 | SYSTEMS AND METHODS FOR IMAGE GUIDED TISSUE ABLATION | A real-time, intra-uterine operating field image is displayed to an operator. A processor identifies anatomical features and marks anatomical boundaries on the image with the aid of a trained classifier model. Labels for the identified anatomical features and the anatomical boundaries are displayed on the image. The trained classifier model is trained based on a set of prior surgical field images that have anatomical features and anatomical boundaries therein identified and marked, respectively, by previous operators. The trained classifier model is typically a convolutional deep neural network. The operator is allowed to modify the anatomical feature labels and anatomical boundaries, as well as update the set of training images with the modified image. The labeling of the anatomical features and establishment of the anatomical features facilitate performing treatments in the uterus, such as the ablation of uterine fibroids. | 1. CLAIMSWHAT IS CLAIMED IS: 1. A computer-implemented method of planning and implementing a treatment procedure, the method comprising: (i) displaying a real-time operating field image to an operator; (ii) identifying one or more of (a) at least one anatomical feature or (b) at least one anatomical boundary on the real-time operating field image, with the aid of the trained classifier model; (iii) displaying one or more of (a) at least one label for the identified at least one anatomical feature or (b) at least one marking for the identified at least one anatomical boundary on the displayed real-time operating field image; (iv) allowing the operator to modify the real-time operating field image by one or more of (a) re-labelling the identified at least one anatomical feature or (b) re-positioning the at least one marking for the at least one anatomical boundary; and(v) provide an update to the trained classifier model based on the modifications by the operator. 2. The method of claim 1, wherein the trained classifier model is generated by: (vi) displaying an operating field image to a user; (vii) establishing at least one anatomical boundary in the operating field image, the at least one anatomical boundary being marked by the user; (viii) labelling at least one anatomical feature separated by the at least one anatomical boundary in the operating field image, the at least one anatomical feature being labeled by user; and(ix) repeating steps (vi) to (viii) to generate a training set of operating field images each with at least one marked anatomical boundary and at least one labeled anatomical feature; and(x) training a classifier model, based on the training set of operating field images, to identify (a) at least one anatomical feature or (b) at least one anatomical boundary on the realtime operating image. 3. The method of claim 2, wherein the trained classifier model is updated at step (v) by adding the modified real-time operating field image from step (iv) to the training set of operating field images and updating the classifier model based on the set of operating field images with the added modified real-time operating field image. -40-RECTIFIED SHEET (RULE 91) 4. The method of claim 1, further comprising confirming with the operator the one or more of (a) the re-labelled identified at least one anatomical feature or (b) the re-positioned at least one marking for the at least one anatomical boundary before providing the update to the trained classifier model. 5. The method of claim 4, wherein providing the update to the trained classifier model comprises adding the modified real-time operating field image to a training set of operating field images. 6. The method of claim 1, wherein the operating field image at step (i) is a surgical field image. 7. The method of claim 1, further comprising treating an anatomical structure based on one or more of (a) the displayed at least one label for the at least one identified anatomical feature or (b) the displayed at least one marking for the at least one anatomical boundary on the real-time operating field image. 8. The method of claim 7, wherein treating the anatomical structure comprises ablating the anatomical structure. 9. The method of claim 8, wherein the anatomical structure is ablated using one or more of radiofrequency (RF) energy, plasma energy, therapeutic ultrasound energy, microwave energy, heat, or cold. 10. The method of claim 7, wherein the treated anatomical structure is a uterine fibroid. 11. The method of claim 7, further comprising repeating steps (ii) to (iv) in real-time during the treatment procedure. 12. The method of claim 7, further comprising repeating one or more of steps (ii) or (iii) after the treatment procedure to identify one or more changes to the anatomical structure from the treatment procedure. 13. The method of claim 1, wherein the operating field image at step (i) comprises an ultrasound image. 14. The method of claim 13, wherein the ultrasound image is an intra-uterine ultrasound image. 15. The method of claim 1, wherein the operating field image at step (i) comprises an image of a uterus of a patient. -41-RECTIFIED SHEET (RULE 91) 16 The method of claim 15, wherein at least one marking for the at least one anatomical boundary separates a plurality of anatomical features of the uterus, the plurality of anatomical features including one or more of a serosa, a myometrium, a fibroid, a uterine wall, a bladder wall, or a bladder. 17. The method of claim 15, wherein the at least one label for the at least one anatomical feature includes one or more of a serosa, a myometrium, a fibroid, a uterine wall, a bladder wall, or a bladder. 18. The method of claim 1, wherein the trained classifier model comprises a machine learning algorithm. 19. The method of claim 18, wherein the machine learning algorithm is a convolutional deep learning network. 20. The method of claim 1, wherein the trained classifier model is one or more of a binary tree, a random forest, a decision tree, a plurality of decision trees, a plurality of decision trees with controlled variance, a multinomial logistic regression, a naive Bayes classifier, a linear classifier, an ensemble of linear classifiers, a boosting algorithm, a boosting algorithm trained with stochastic gradient descent, a boosting algorithm trained with stochastic gradient descent, a boosting algorithm comprising training data weighting, a boosting algorithm comprising updating training data weighting, or a boosting algorithm comprising updating misclassified training data with higher weights. 21. The method of claim 1, further comprising providing a pre-treatment image to the operator as a guide for identifying the one or more of (a) the at least one anatomical feature or (b) the at least one anatomical boundary on the real-time operating field image. 22. The method of claim 21, wherein the pre-treatment image comprises one or more of an MRI image, an X-ray image, a CT image, or an ultrasound image. 23. An apparatus for aiding in planning and implementing a treatment procedure, the apparatus comprising: a display; an interface; a processor; and a non-transitory computer readable storage medium including instructions configured to cause the processor to: (i) cause the display to display a real-time operating field image to an operator,-42-RECTIFIED SHEET (RULE 91) (ii) identify one or more of (a) at least one anatomical feature or (b) at least one anatomical boundary on the real-time operating field image, with the aid of the trained classifier model,(iii) display one or more of (a) at least one label for the identified at least one anatomical feature or (b) at least one marking for the identified at least one anatomical boundary on the displayed real-time operating field image,(iv) allow the operator to modify the real-time operating field image by one or more of (a) re-labelling the at least one identified anatomical features or (b) re-positioning the at least one marking for the at least one identified anatomical boundary, and(v) provide an update to the trained classifier model based on the modifications by the operator. 24. The apparatus of claim 23, wherein the instructions are further configured to cause the processor to: (vi) display an operating field image to a user; (vii) establish at least one anatomical boundary in the operating field image, the at least one anatomical boundary being marked by the user via the interface; (viii) label at least one anatomical feature separated by the at least one anatomical boundary, the at least one anatomical feature being labeled by the user via the interface; (ix) repeat steps (vi) to (viii) to generate a training set of operating field images each with at least one marked anatomical boundary and at least one labeled anatomical feature; and(x) train a classifier model, based on the training set of operating field images, to at least one of (a) identify at least one anatomical feature or (b) establish at least one anatomical boundary on the real-time operating image. 25. The apparatus of claim 24, wherein the trained classifier model is updated at step (v) by adding the modified real-time operating field image from step (iv) to the training set of operating field images and updating the classifier model based on the set of operating field images with the added modified real-time operating field image. 26. The apparatus of claim 23, wherein the instructions are further configured to cause the processor to confirm with the operator the one or more of (a) the re-labelled at least one anatomical feature or (b) the re-positioned at least one marking for the at least one anatomical boundary before providing the update to the trained classifier model. 27. The apparatus of claim 26, wherein the update is provided to the trained classifier model by adding the modified real-time operating field image to a training set of operating field images. -43-RECTIFIED SHEET (RULE 91) 28. The apparatus of claim 23, wherein the operating field image at step (i) is an operating field image. 29. The apparatus of claim 23, wherein the instructions are further configured to cause the processor to further allow the operator to use the interface to treat an anatomical structure based on one or more of (a) the displayed at least one label for the at least one identified anatomical feature or (b) the displayed at least one marking for the at least one anatomical boundary on the real-time operating field image. 30. The apparatus of claim 29, wherein treating the anatomical structure comprises ablating the anatomical structure with an ablation element, and wherein the apparatus further comprises the ablation element. 31. The apparatus of claim 30, wherein the ablation element is configured to delivery one or more of radiofrequency (RF) energy, plasma energy, therapeutic ultrasound energy, microwave energy, heat, or cold to the anatomical structure. 32. The apparatus of claim 29, wherein the treated anatomical structure is a uterine fibroid. 33. The apparatus of claim 29, wherein the instructions are further configured to cause the processor to repeat steps (ii) to (iv) in real-time during the treatment procedure. 34. The apparatus of claim 29, wherein the instructions are further configured to cause the processor to repeat one or more of steps (ii) or (iii) after the treatment procedure to identify one or more changes to the anatomical structure from the treatment procedure. 35. The apparatus of claim 23, wherein the operating field image at step (i) comprises an ultrasound image. 36. The apparatus of claim 35, wherein the ultrasound image is an intra-uterine ultrasound image. 37. The apparatus of claim 23, wherein the operating field image at step (i) comprises an image of a uterus of a patient. 38. The apparatus of claim 37, wherein the at least one marking for the at least one anatomical boundary separates a plurality of anatomical features of the uterus, the plurality of anatomical features including one or more of a serosa, a myometrium, a fibroid, a uterine wall, a bladder wall, or a bladder. -44-RECTIFIED SHEET (RULE 91) 39. The apparatus of claim 37, wherein the at least one label for the at least one anatomical feature includes one or more of a serosa, a myometrium, a fibroid, a uterine wall, a bladder wall, or a bladder. 40. The apparatus of claim 23, wherein the classifier model comprises a machine learning algorithm. 41. The apparatus of claim 40, wherein the machine learning algorithm is a convolutional deep learning network. 42. The apparatus of claim 23, wherein the classifier model is one or more of a binary tree, a random forest, a decision tree, a plurality of decision trees, a plurality of decision trees with controlled variance, a multinomial logistic regression, a naive Bayes classifier, a linear classifier, an ensemble of linear classifiers, a boosting algorithm, a boosting algorithm trained with stochastic gradient descent, a boosting algorithm trained with stochastic gradient descent, a boosting algorithm comprising training data weighting, a boosting algorithm comprising updating training data weighting, or a boosting algorithm comprising updating misclassified training data with higher weights. 43. The apparatus of claim 23, wherein the instructions are further configured to cause the processor to further provide a pre-treatment image to the operator as a guide for identifying the one or more of (a) the at least one anatomical feature or (b) the at least one anatomical boundary on the real-time operating field image. 44. The apparatus of claim 43, wherein the pre-treatment image comprises one or more of an MRI image, an X-ray image, a CT image, or an ultrasound image. 45. A computer-implemented method of planning and implementing a treatment procedure, the method comprising: (i) displaying a real-time operating field image to an operator; (ii) identifying one or more of (a) at least one anatomical feature or (b) at least one anatomical boundary on the real-time operating field image, with the aid of a trained classifier model; and(iii) displaying one or more of (a) at least one label for the identified at least one anatomical feature or (b) at least one marking for the at least one anatomical boundary on the real-time operating field image, wherein the trained classifier model is trained based on a set of operating field images each with at least one marked anatomical boundary and at least one labeled anatomical feature,-45-RECTIFIED SHEET (RULE 91) wherein the at least one marked anatomical boundary is identified by a previous operator, and wherein the at least one labeled anatomical feature is identified by the previous operator. 46. The method of claim 45, further comprising allowing the operator to modify the real-time operating field image by one or more of (a) re-labelling the identified at least one anatomical feature or (b) re-positioning the identified at least one anatomical boundary. 47. The method of claim 46, further comprising updating the set of operating field images by adding the modified real-time operating field image to the set. 48. The method of claim 47, further comprising updating the classifier model based on the updated set of operating field images. 49. The method of claim 45, further comprising treating an anatomical structure based on one or more of (a) the displayed at least one label for the identified at least one anatomical feature or (b) the displayed at least one marking for the at least one anatomical boundary on the real-time surgical field image. 50. The method of claim 49, wherein treating the anatomical structure comprises ablating the anatomical structure. 51. The method of claim 49, wherein the anatomical structure is ablated using one or more of radiofrequency (RF) energy, plasma energy, therapeutic ultrasound energy, microwave energy, heat, or cold. 52. The method of claim 49, wherein the treated anatomical structure is a uterine fibroid. 53. The method of claim 49, further comprising repeating steps (ii) and (iii) in realtime during the treatment procedure. 54. The method of claim 49, further comprising repeating one or more of steps (ii) or (iii) after the treatment procedure to identify one or more changes to the anatomical structure from the treatment procedure. 55. The method of claim 45, wherein the real-time operating field image comprises an ultrasound image. 56. The method of claim 55, wherein the ultrasound image is an intra-uterine ultrasound image. 57. The method of claim 45, wherein the identified at least one anatomical boundary separates a plurality of anatomical features of a uterus, the plurality of anatomical features-46-RECTIFIED SHEET (RULE 91) including one or more of a serosa, a myometrium, a fibroid, a uterine wall, a bladder wall, or a bladder. 58. The method of claim 45, wherein the at least one labeled anatomical feature includes one or more of a serosa, a myometrium, a fibroid, a uterine wall, a bladder wall, or a bladder. 59. The method of claim 45, wherein the classifier model comprises a machine learning algorithm. 60. The method of claim 59, wherein the machine learning algorithm is a convolutional deep learning network. 61. The method of claim 59, wherein the classifier model is one or more of a binary tree, a random forest, a decision tree, a plurality of decision trees, a plurality of decision trees with controlled variance, a multinomial logistic regression, a naive Bayes classifier, a linear classifier, an ensemble of linear classifiers, a boosting algorithm, a boosting algorithm trained with stochastic gradient descent, a boosting algorithm trained with stochastic gradient descent, a boosting algorithm comprising training data weighting, a boosting algorithm comprising updating training data weighting, or a boosting algorithm comprising updating misclassified training data with higher weights. 62. The method of claim 45, further providing a pre-treatment image to the operator as a guide for identifying the one or more of (a) the at least one anatomical feature or (b) the at least one anatomical boundary on the real-time operating field image. 63. The method of claim 62, wherein the pre-treatment image comprises one or more of an MRI image, an X-ray image, a CT image, or an ultrasound image. 64. An apparatus for aiding in performing implementing a treatment procedure, the apparatus comprising: a display; an interface; a processor; and a non-transitory computer readable storage medium including instructions configured to cause the processor to: (i) display a real-time operating field image to an operator; (ii) identify one or more of (a) at least one anatomical feature or (b) at least one anatomical boundary on the real-time operating field image, with the aid of a trained classifier model; and-47-RECTIFIED SHEET (RULE 91) (iii) display one or more of (a) at least one label for the identified at least one anatomical feature or (b) at least one marking for the identified at least one anatomical boundary on the real-time operating field image, wherein the trained classifier model is trained based on a set of operating field images each with at least one marked anatomical boundary and at least one labeled anatomical feature, wherein the at least one marked anatomical boundary is identified by a previous operator, and wherein the at least one labeled anatomical feature is identified by the previous operator. 65. The apparatus of claim 64, wherein the instructions are further configured to cause the processor to allow the operator to modify the real-time operating field image by one or more of (a) re-labelling the identified at least one anatomical feature or (b) re-positioning the identified at least one anatomical boundary. 66. The apparatus of claim 64, wherein the instructions are further configured to cause the processor to update the set of operating field images by adding the modified real-time operating field image to the set. 67. The apparatus of claim 66, wherein the instructions are further configured to cause the processor to further update the classifier model based on the updated set of operating field images. 68. The apparatus of claim 64, wherein the instructions are further configured to cause the processor to further allow the operator to use the interface to treat an anatomical structure based on one or more of (a) the displayed at least one label for the identified at least one anatomical feature or (b) the displayed at least one marking for the at least one anatomical boundary on the real-time surgical field image. 69. The apparatus of claim 68, wherein treating the anatomical structure comprises ablating the anatomical structure with an ablation element, and wherein the apparatus further comprises the ablation element. 70. The apparatus of claim 69, wherein the ablation element is configured to delivery one or more of radiofrequency (RF) energy, plasma energy, therapeutic ultrasound energy, microwave energy, heat, or cold to the anatomical structure. 71. The apparatus of claim 68, wherein the treated anatomical structure is a uterine fibroid. 72. The apparatus of claim 68, wherein the instructions are further configured to cause the processor to repeat steps (ii) and (iii) during the treatment procedure. -48-RECTIFIED SHEET (RULE 91) 73. The apparatus of claim 68, wherein the instructions are further configured to cause the processor to repeat one or more of steps (ii) or (iii) after the treatment procedure to identify one or more changes to the anatomical structure from the treatment procedure. 74. The apparatus of claim 64, wherein the real-time operating field image comprises an ultrasound image. 75. The apparatus of claim 73, wherein the ultrasound image is an intra-uterine ultrasound image. 76. The apparatus of claim 64, wherein the real-time operating field image comprises an image of a uterus of a patient. 77. The apparatus of claim 64, wherein the identified at least one anatomical boundary separates a plurality of anatomical features of the uterus, the plurality of anatomical features including one or more of a serosa, a myometrium, a fibroid, a uterine wall, a bladder wall, or a bladder. 78. The apparatus of claim 64, wherein the at least one labeled anatomical feature includes one or more of a serosa, a myometrium, a fibroid, a uterine wall, a bladder wall, or a bladder. 79. The apparatus of claim 64, wherein the classifier model comprises a machine learning algorithm. 80. The apparatus of claim 79, wherein the machine learning algorithm is a convolutional deep learning network. 81. The apparatus of claim 79, wherein the classifier model is one or more of a binary tree, a random forest, a decision tree, a plurality of decision trees, a plurality of decision trees with controlled variance, a multinomial logistic regression, a naive Bayes classifier, a linear classifier, an ensemble of linear classifiers, a boosting algorithm, a boosting algorithm trained with stochastic gradient descent, a boosting algorithm trained with stochastic gradient descent, a boosting algorithm comprising training data weighting, a boosting algorithm comprising updating training data weighting, or a boosting algorithm comprising updating misclassified training data with higher weights. 82. The apparatus of claim 64, wherein the instructions are further configured to cause the processor to further provide a pre-treatment image to the operator as a guide for-49-RECTIFIED SHEET (RULE 91) identifying the one or more of (a) the at least one anatomical feature or (b) the at least one anatomical boundary on the real-time operating field image. 83. The apparatus of claim 82, wherein the pre-treatment image comprises one or more of an MRI image, an X-ray image, a CT image, or an ultrasound image. 84. A computer-implemented method of planning and implementing a treatment procedure, the method comprising: (i) displaying a real-time operating field image to an operator; (ii) identifying one or more of (a) at least one anatomical feature or (b) at least one anatomical boundary on the real-time operating field image, with the aid of the trained classifier model; (iii) displaying one or more of (a) at least one label for the identified at least one anatomical feature or (b) at least one marking for the identified at least one anatomical boundary on the displayed real-time operating field image; and(iv) treating an anatomical structure based on one or more of (a) the displayed at least one label for the identified at least one anatomical feature or (b) the displayed at least one marking for the identified at least one anatomical boundary on the real-time operating field image. 85. The method of claim 84, wherein the real-time operating field image is a real-time surgical field image. 86. The method of claim 84, wherein the real-time operating field image is an ultrasound image. 87. The method of claim 86, wherein the ultrasound image is an intra-uterine ultrasound image. 88. The method of claim 84, wherein the operating field image is an image of a uterus of a patient. 89. The method of claim 88, wherein the identified at least one anatomical boundary separate a plurality of anatomical features of the uterus, the plurality of anatomical features including one or more of a serosa, a myometrium, a fibroid, a uterine wall, a bladder wall, or a bladder. 90. The method of claim 88, wherein the identified at least one anatomical feature includes one or more of a serosa, a myometrium, a fibroid, a uterine wall, a bladder wall, or a bladder. -50-RECTIFIED SHEET (RULE 91) 91. The method of claim 84, wherein treating the anatomical structure comprises ablating the anatomical structure. 92. The method of claim 91, wherein the anatomical structure is ablated using one or more of radiofrequency (RF) energy, plasma energy, therapeutic ultrasound energy, microwave energy, heat, or cold. 93. The method of claim 84, wherein the treated anatomical structure is a uterine fibroid. 94. The method of claim 84, wherein the trained classifier model comprises a machine learning algorithm. 95. The method of claim 84, further comprising: allowing the operator to modify the real-time operating field image by one or more of (a) re-labelling the identified at least one anatomical feature or (b) re-positioning the marked at least one anatomical boundary; and providing an update to the trained classifier model based on the modifications by the operator. 96. The method of claim 84, further comprising repeating steps (ii) and (iii) during the treatment procedure. 97. The method of claim 84, further comprising repeating one or more of steps (ii) or (iii) after the treatment procedure to identify one or more changes to the anatomical structure from the treatment procedure. 98. The method of claim 84, further comprising providing a pre-treatment image to the operator as a guide for identifying the one or more of (a) the at least one anatomical feature or (b) the at least one anatomical boundary on the real-time operating field image. 99. The method of claim 98, wherein the pre-treatment image comprises one or more of an MRI image, an X-ray image, a CT image, or an ultrasound image. 100. An apparatus for aiding in implementing a treatment procedure, the apparatus comprising: a display; an interface; a processor; and a non-transitory computer readable storage medium including instructions configured to cause the processor to: -51-RECTIFIED SHEET (RULE 91) (i) display a real-time operating field image to an operator; (ii) identify one or more of (a) at least one anatomical feature or (b) at least one anatomical boundary on the real-time operating field image, with the aid of the trained classifier model; (iii) display one or more of (a) at least one label for the identified at least one anatomical feature or (b) at least one marking for the identified at least one anatomical boundary on the displayed real-time operating field image; and(iv) treat an anatomical structure based on one or more of (a) the displayed at least one label for the identified at least one anatomical feature or (b) the displayed at least one marking for the identified at least one anatomical boundary on the real-time operating field image. 101. The apparatus of claim 100, wherein the real-time operating field image is a realtime surgical field image. 102. The apparatus of claim 100, wherein the real-time operating field image is an ultrasound image. 103. The apparatus of claim 102, wherein the ultrasound image is an intra-uterine ultrasound image. 104. The apparatus of claim 100, wherein the operating field image is an image of a uterus of a patient. 105. The apparatus of claim 104, wherein the identified at least one anatomical boundary separates a plurality of anatomical features of the uterus, the plurality of anatomical features including one or more of a serosa, a myometrium, a fibroid, a uterine wall, a bladder wall, or a bladder. 106. The apparatus of claim 104, wherein the identified at least one anatomical feature includes one or more of a serosa, a myometrium, a fibroid, a uterine wall, a bladder wall, or a bladder. 107. The apparatus of claim 100, wherein treating the anatomical structure comprises ablating the anatomical structure, and wherein the apparatus further comprises an ablation element. 108. The apparatus of claim 107, wherein the ablation element is configured to ablate the anatomical structure using one or more of radiofrequency (RF) energy, plasma energy, therapeutic ultrasound energy, microwave energy, heat, or cold. -52-RECTIFIED SHEET (RULE 91) 109. The apparatus of claim 100, wherein the treated anatomical structure is a uterine fibroid. 110. The apparatus of claim 100, wherein the trained classifier model comprises a machine learning algorithm. 111. The apparatus of claim 100, wherein the instructions are further configured to cause the processor to: allow the operator to modify the real-time operating field image by one or more of (a) relabelling the identified at least one anatomical feature or (b) re-positioning the marked at least one anatomical boundary; and provide an update to the trained classifier model based on the modifications by the operator. 112. The apparatus of claim 100, wherein the instructions are further configured to cause the processor to repeat steps (ii) and (iii) during the treatment procedure. 113. The apparatus of claim 100, wherein the instructions are further configured to cause the processor to repeat one or more of steps (ii) or (iii) after the treatment procedure to identify one or more changes to the anatomical structure from the treatment procedure. 114. The apparatus of claim 100, wherein the instructions are further configured to cause the processor to further provide a pre-treatment image to the operator as a guide for identifying the one or more of (a) the at least one anatomical feature or (b) the at least one anatomical boundary on the real-time operating field image. 115. The apparatus of claim 114, wherein the pre-treatment image comprises one or more of an MRI image, an X-ray image, a CT image, or an ultrasound image. -53-RECTIFIED SHEET (RULE 91) | A real-time intra-uterine image-guided ablation tool that uses a trained classifier to identify and label anatomical features and boundaries on the live image. Operators see automatic anatomical labels overlaid on the operating field. | R | "identifying one or more of (a) at least one anatomical feature or (b) at least one anatomical boundary on the real-time operating field image, with the aid of the trained classifier model" (Claim 1) "treating an anatomical structure based on one or more of (a) the displayed at least one label for the at least one identified anatomical feature or (b) the displayed at least one marking for the at least one anatomical boundary" (Claim 7) "repeating one or more of steps (ii) or (iii) after the treatment procedure to identify one or more changes to the anatomical structure from the treatment procedure" (Claim 12) "anatomical structure is ablated using one or more of radiofrequency (RF) energy, plasma energy, therapeutic ultrasound energy, microwave energy, heat, or cold" (Claim 9) |  |  | 1 | [5 empty] | 1 | [12 empty] | 1 | [8 empty] | 1 (uterine fibroid — gynecologic) | GYNESONICS INC | GYNESONICS INC | 2021-10-26 | 2021 | 2022-05-05 | 2022 | 2020-10-27 | 2020 | WO2021US56692 | CN116669635A¦EP4238106A1¦EP4238106A4¦US20230260121A1
391 | US20250380980A1 | SYSTEMS AND METHODS FOR UTERINE FIBROID ABLATION | Various systems, devices, and methods for tissue penetration are disclosed. The system includes an ablation element that penetrates the tissue and a radiofrequency generator that delivers energy to the ablation element. The radiofrequency generator may have two modes: a cutting or insertion mode and an ablation or coagulation mode. The cutting or insertion mode is used to cut or penetrate the tissue by providing a voltage or power modulation that acts to cut or soften tissue contacted by the ablation element. The ablation or coagulation mode is used to ablate or coagulate the tissue. | 1. A system for penetrating target tissue, the system comprising: an ablation instrument; and a controller configured to control the delivery of energy to the ablation instrument, wherein the controller is configured to monitor a temperature measured by the ablation instrument at an interface with the target tissue and maintain the temperature between about 80° C. and about 115° C. as the ablation instrument penetrates into the target tissue. 2. The system of claim 1, wherein the ablation instrument comprises an introducer and a plurality of electrode ablation needles extendible from a retracted position within the introducer. 3. The system of claim 2, wherein the temperature measured by the ablation instrument is measured by a thermocouple positioned on one of the electrode ablation needles when the plurality of electrode ablation needles is retracted within the introducer. 4. The system of claim 3, wherein the plurality of electrode ablation needles comprises a central electrode extendible from a retracted position within the introducer, and wherein the thermocouple is positioned on the central electrode. 1. ystem of claim 1, wherein the controller is configured to modulate power delivered to the ablation instrument to maintain the temperature between about 80° C. and about 115° C. as the ablation instrument penetrates into the target tissue. 6. The system of claim 1, wherein the controller is configured to modulate power delivered to the ablation instrument as the ablation instrument penetrates into the target tissue to provide an oscillation in the temperature at the interface with the target tissue. 7. The system of claim 1, wherein the controller is configured to provide an alert when the temperature at the interface is between about 80° C. and about 115° C. to provide an indication to begin penetration by the ablation instrument into the target tissue. 8. The system of claim 1, wherein the system further comprises an ultrasonic imaging device configured to provide for visualization of the ablation instrument as the ablation instrument penetrates the target tissue. 9. The system of claim 8, wherein the ultrasonic imaging device is coupled to the ablation instrument. 10. The system of claim 1, wherein the controller is configured to control the delivery of energy to the ablation instrument to ablate the target tissue after the ablation instrument has penetrated into the target tissue. 11. The system of claim 10, wherein the controller is configured to control the delivery of energy to the ablation instrument to maintain a substantially constant temperature at the target tissue to ablate the target tissue. 12. The system of claim 1, further comprising a radiofrequency generator configured to deliver energy to the ablation instrument while the ablation instrument penetrates into the target tissue. 13. A method of penetrating target tissue, the method comprising: delivering an ablation instrument into contact with a target tissue; delivering energy to the ablation instrument to cause a temperature measured by the ablation instrument at an interface with the target tissue to rise to between about 80° C. and about 115° C.; and advancing the ablation instrument to penetrate into the target tissue while modulating the power to maintain the temperature at the interface with the target tissue between about 80° C. and about 115° C. 14. The method of claim 13, wherein the ablation instrument penetrates into the target tissue while the temperature maintained at the interface with the target tissue oscillates. 15. The method of claim 13, further comprising ablating the target tissue by delivering energy to the ablation instrument to maintain the temperature at the interface with the target tissue at a substantially constant temperature. 16. The method of claim 13, wherein the ablation instrument comprises an introducer and a plurality of electrode ablation needles extendible from a retracted position within the introducer, wherein the plurality of electrode ablation needles is at least partially extended from the introducer while maintaining the temperature at the interface with the target tissue between about 80° C. and about 115° C. 17. A system for penetrating target tissue, the system comprising: an ablation element configured to penetrate the target tissue; and a radiofrequency generator configured to deliver energy to the ablation element, wherein the radiofrequency generator comprises: a cutting or insertion mode configured to cut through target tissue, wherein the cutting or insertion mode provides for a power oscillation configured to cut tissue contacted by the ablation element; and an ablation or coagulation mode configured to ablate and/or coagulate the target tissue, wherein the ablation or coagulation mode provides for an increase in power and then a decrease in power after the target tissue reaches a target temperature. 18. The system of claim 17, wherein the cutting or insertion mode is configured to maintain a substantially constant surface temperature of the ablation element. 19. The system of claim 17, wherein the cutting or insertion mode is configured to cause a rapid increase in temperature in tissue in contact with the ablation element. 20. The system of claim 17, wherein the radiofrequency generator is configured to be controlled to provide a limit on temperature during the cutting or insertion mode and during the ablation or coagulation mode, wherein the limit on temperature during the cutting or insertion mode is lower than the limit on temperature during the ablation or coagulation mode. 21. The system of claim 17, further configured to provide a mechanical vibration or cutting force during the cutting or insertion mode. 22. The system of claim 17, wherein the ablation element comprises an introducer. 23. The system of claim 17, wherein the ablation element comprises a plurality of electrode ablation needles. 24. The system of claim 17, further comprising a controller configured to control the delivery of energy to the ablation element, wherein the controller is configured to monitor a temperature measured by the ablation element at an interface with the target tissue and maintain the temperature between about 80° C. and about 115° C. as the ablation element penetrates into the target tissue. 25. The system of claim 17, wherein the target tissue is a uterine fibroid. 26. A method of uterine fibroid ablation, comprising: delivering an ablation element into contact with a uterine fibroid; delivering radiofrequency energy according to a cutting or insertion mode to the ablation element while the ablation element is in contact with the uterine fibroid, wherein in the cutting or insertion mode a voltage or power is modulated based on a temperature measured at an interface between the ablation element and the uterine fibroid to assist the ablation element in penetrating into the uterine fibroid; and after the ablation element penetrates into the uterine fibroid, delivering radiofrequency energy according to a coagulation mode, wherein in the coagulation mode a voltage or power applied to the ablation element is sufficient to ablate the uterine fibroid. 27. The method of claim 26, wherein the voltage or power is oscillated in the cutting or insertion mode. 28. The method of claim 26, wherein the voltage or power is modulated in the cutting or insertion mode to maintain a temperature at the interface between about 80° C. and about 115° C. 29. The method of claim 26, wherein the voltage or power in the coagulation mode increases and then decreases after tissue in contact with the ablation element reaches a target temperature. 30. The method of claim 26, wherein delivering radiofrequency energy according to the cutting or insertion mode preheats the uterine fibroid. 31. The method of claim 26, wherein delivering radiofrequency energy according to the cutting or insertion mode softens the uterine fibroid thereby allowing the ablation element to penetrate the uterine fibroid without substantially deforming the uterine fibroid upon penetration. 32. The method of claim 26, wherein delivering radiofrequency energy according to the cutting or insertion mode maintains a substantially constant surface temperature of the ablation element. 33. The method of claim 32, wherein the substantially constant surface temperature of the ablation element is between about 80° C. and about 115° C. 34. The method of claim 26, wherein delivering radiofrequency energy according to the cutting or insertion mode is controlled to a maximum output not to exceed 70 watts. 35. The method of claim 26, wherein delivering radiofrequency energy according to the cutting or insertion mode is controlled not to exceed 30 seconds. 36. The method of claim 26, further comprising detecting a force on the ablation element when in contact with the uterine fibroid, and wherein delivering radiofrequency energy according to the cutting or insertion mode is controlled based on detection of the force. | A uterine-fibroid ablation device whose RF generator has two modes — cutting/insertion (modulated voltage that softens tissue) and ablation/coagulation. The same probe penetrates the fibroid and then ablates it without switching tools. | R | "controller is configured to monitor a temperature measured by the ablation instrument at an interface with the target tissue and maintain the temperature between about 80° C. and about 115° C. as the ablation instrument penetrates into the target tissue" (Claim 1) "modulate power delivered to the ablation instrument as the ablation instrument penetrates into the target tissue to provide an oscillation in the temperature at the interface with the target tissue" (Claim 6) "ultrasonic imaging device configured to provide for visualization of the ablation instrument as the ablation instrument penetrates the target tissue" (Claim 8) "detecting a force on the ablation element when in contact with the uterine fibroid, and wherein delivering radiofrequency energy according to the cutting or insertion mode is controlled based on detection of the force" (Claim 36) |  |  | 1 |  | 1 | [5 empty] | 1 | [19 empty] | 1 (uterine fibroid — gynecologic) | GYNESONICS INC | GYNESONICS INC | 2023-06-27 | 2023 | 2025-12-18 | 2025 | 2022-06-28 | 2022 | US202318879178 | CN119630358A¦EP4547135A1¦JP2025521812A¦WO2024006789A1
392 | CN116983075A | Temperature monitoring compensation algorithm and system for ablation needle | The invention relates to a temperature monitoring compensation algorithm and system for an ablation needle, and relates to the technical field of ablation. According to the temperature monitoring compensation algorithm of the ablation needle, the temperature of the target area is obtained by calibrating the measured temperature data x and considering the relative temperature change of the cooling medium, and the actual temperature of the target area can be indicated, so that precise surgical treatment of tumors can be realized; and the risk of relapse caused by incomplete ablation inactivation and serious complications caused by excessive ablation can be greatly reduced. | 1. A temperature monitoring and compensation algorithm for an ablation needle, characterized in that it includes the following steps: Obtaining the measured temperature data X of the proximal end of the needle shaft of the ablation needle; Obtaining the temperature data % of the cooling medium in the needle shaft; Calibrate the above-mentioned measured temperature data Calculate the temperature F(Xk) of the target area based on the calibration value f(X) and the relative temperature change value fω of the cooling medium, where, F (ζ, t) =f (ζ) + f (t) ο 2. The temperature monitoring and compensation algorithm of the ablation needle according to claim 1, characterized in that calibrating the measured temperature data X to obtain the calibration value f ω includes the following sub-steps: The standard temperature of the environment is divided into multiple standard temperature intervals; Obtain the starting measurement value and end measurement value corresponding to the starting standard temperature and end standard temperature in any standard temperature interval; According to the starting standard temperature, end standard temperature, The starting measurement value and the end measurement value are used to fit the measured temperature data X to obtain the calibration value f CO. 3. The temperature monitoring and compensation algorithm of the ablation needle according to claim 2, characterized in that the calibration value f (z) satisfies the following relationship: f fx) = - ~ (yi - y.) + y. ; Sun 1 3 Among them, is the starting standard temperature in any standard temperature interval; is the end standard temperature in the standard temperature interval; X. It is the starting measurement value corresponding to the starting standard temperature within the standard temperature range; X1 is the end measurement value corresponding to the end standard temperature within the standard temperature range. 4. The temperature monitoring and compensation algorithm of the ablation needle according to any one of claims 1 to 3, characterized in that when the measured temperature data X is calibrated to obtain the calibration value f ω, the temperature of the same cooling medium is Reference; or when calibrating the measured temperature data X to obtain the calibration value, the temperature of the cooling medium is the initial temperature. 5. The temperature monitoring and compensation algorithm of the ablation needle according to claim 3, characterized in that, obtaining the relative temperature change value f&) of the cooling medium according to the temperature data t of the cooling medium includes the following sub-steps: Obtaining the initial temperature of the cooling medium; Obtaining The first measured value corresponding when the temperature of the cooling medium is the first cooling temperature, and the second measured value corresponding when the temperature of the cooling medium is the second cooling temperature is obtained; For the initial temperature, the first cooling temperature, the second cooling temperature, The first measured value and the second measured value are subjected to a fitting operation to obtain the relative temperature change value of the cooling medium fl). 6. The temperature monitoring and compensation algorithm of the ablation needle according to claim 5, characterized in that the relative temperature change value f of the cooling medium satisfies the following relationship: where is the initial cooling water temperature; J is the first cooling temperature; G is the second cooling temperature; Xi is the first measured value corresponding to the first cooling temperature J; X2 is the second measured value corresponding to the second cooling temperature t2. 7. The temperature monitoring and compensation algorithm of the ablation needle according to claim 5, characterized in that the relative temperature change value f of the cooling medium satisfies the following relationship: f (t) = kx (t - to); wherein 4 is the coefficient. Water=0.2-0.4; Product is the initial cooling water temperature. 8. The temperature monitoring and compensation algorithm of the ablation needle according to any one of claims 1 to 3, characterized in that the needle shaft of the ablation needle is obtained through the first temperature measurement element at the proximal end of the needle shaft of the ablation needle. Proximal measured temperature data The second temperature measuring element in the needle bar obtains the temperature data of the cooling medium in the needle bar, wherein the second temperature measuring element is located on the distal outer surface of the needle bar or in the cooling medium. 9. The temperature monitoring and compensation algorithm of the ablation needle according to claim 6, characterized in that, when the relative temperature change value f(t) of the cooling medium is obtained according to the temperature data t of the cooling medium, the same needle of the ablation needle is used. The standard temperature of the environment near the rod's proximal end is used as a reference. 10. A temperature monitoring and compensation system for an ablation needle, characterized in that it includes an ablation needle and an ablation instrument connected to the ablation needle. The ablation instrument includes: a temperature acquisition module, which is used to obtain the needle temperature of the ablation needle respectively. The measured temperature data t obtains the relative temperature change value f S) of the cooling medium, and calculates the temperature of the target area based on the calibration value f 0O and the relative temperature change value f1) of the cooling medium, where, F(X, t) =f(X)+ f (t) ο 11. The temperature monitoring and compensation system of the ablation needle according to claim 10, wherein the temperature acquisition module includes a first temperature measurement element for acquiring the measured temperature data X of the proximal end of the needle shaft of the ablation needle and A second temperature measuring element used to obtain the temperature data t of the cooling medium in the needle bar. The ablation needle includes a needle bar and a multi-layer isolation structure provided on the needle bar. The multi-layer isolation structure includes an internal Isolation structure and external isolation structure, the internal isolation structure is arranged outside the needle bar, the external isolation structure includes an external isolation film covering the outer wall of the internal isolation structure, wherein the first temperature measurement The element is located on the proximal outer surface of the needle bar or between the internal isolation structure and the outer isolation film; the second temperature measuring element is located on the distal outer surface of the needle bar or is located on the in the cooling medium. | An ablation-needle temperature compensation algorithm that calibrates measured temperature using the relative temperature change of the cooling medium. The compensation gives a more accurate read of the actual target tissue temperature. | R | "Calculate the temperature F(Xk) of the target area based on the calibration value f(X) and the relative temperature change value f(t) of the cooling medium, where, F(x,t) = f(x) + f(t)" (Claim 1) "first temperature measurement element at the proximal end of the needle shaft of the ablation needle… second temperature measuring element… located on the distal outer surface of the needle bar or in the cooling medium" (Claim 8) "precise surgical treatment of tumors can be realized; and the risk of relapse caused by incomplete ablation inactivation and serious complications caused by excessive ablation can be greatly reduced" (Abstract) |  |  |  |  | 1 | [15 empty] | 1 | [7 empty] | 1 |  |  | HYGEA MEDICAL TECH CO LTD | HYGEA MEDICAL TECH CO LTD | 2023-09-28 | 2023 | 2023-11-03 | 2023 | 2023-09-28 | 2023 | CN202311267891 | CN116983075B
393 | CN120616749A | Temperature data acquisition and report generation method for cervical cryotherapy | The invention belongs to the technical field of cervical cryotherapy, and discloses a cervical cryotherapy temperature data acquisition and report generation method, which comprises the following steps: capturing a temperature gradient based on a multi-point thermosensitive probe array, and generating temperature distribution original data; performing thermal diffusion path analysis to generate freezing front migration characteristic data; positioning a thermal diffusion critical point, and generating freezing depth critical value data; analyzing the interface temperature field to generate a three-dimensional distribution characteristic spectrum; representing an ice crystal formation dynamic state, and generating ice crystallization process data; quantifying the heat conduction effect to generate freezing response data; evaluating tissue damage, and generating depth prediction data; constructing a decision model, and generating regulation and control guidance data; the treatment effect is comprehensively evaluated, and a standardized electronic report is automatically generated; according to the invention, temperature monitoring and individualized scheme optimization in the whole process of cervical cryotherapy are realized, and the therapeutic accuracy and safety are improved. | 1A method for temperature data collection and report generation for cervical cryotherapy, characterized in that it includes: step S1: capturing the freezing interface temperature gradient based on a multi-point thermal probe array to generate raw data on cervical tissue temperature distribution; performing heat diffusion path tracing analysis based on the raw data on cervical tissue temperature distribution to generate freezing front migration characteristic data; step S2: locating and identifying the heat diffusion critical point based on the freezing front migration characteristic data to generate freezing depth critical value data; performing interface temperature field analysis based on the freezing depth critical value data to generate a three-dimensional distribution characteristic map of the freezing zone temperature field; step S3: dynamically characterizing tissue ice crystal formation based on the three-dimensional distribution characteristic map of the freezing zone temperature field to generate tissue ice crystallization process data; performing multi-level heat conduction effect quantification based on the tissue ice crystallization process data to generate tissue freezing response channel characteristic data; Step S4: Map the tissue damage assessment under different freezing cycles for the tissue freezing response channel characteristic data to generate treatment effect depth prediction data; construct an individualized freezing plan decision model based on the treatment effect depth prediction data to generate freezing parameter dynamic control guidance data; conduct a comprehensive evaluation of the treatment effect based on the freezing parameter dynamic control guidance data and the three-dimensional distribution characteristic map of the freezing zone temperature field, and automatically generate a standardized electronic report for cervical cryotherapy. 2.The temperature data collection and report generation method for cervical cryotherapy according to claim 1 is characterized in that, step S1 includes: using a ring-distributed thermal probe array to capture the freezing interface temperature gradient to generate the original data of cervical tissue temperature distribution; performing temperature-time curve deconstruction processing based on the original data of cervical tissue temperature distribution to generate non-uniform cooling characteristic analysis data; calculating the tissue thermal diffusion coefficient based on the non-uniform cooling characteristic analysis data to generate patient-specific thermal conduction characteristic data; performing heat diffusion path tracking analysis based on the patient-specific thermal conduction characteristic data to generate freezing front migration characteristic data. 3.The temperature data collection and report generation method for cervical cryotherapy according to claim 2 is characterized in that, step S2 includes: calculating the temperature contour change rate based on the freezing front migration characteristic data to generate freezing rate gradient change data; accurately locating the tissue freezing transition point based on the freezing rate gradient change data to generate freezing depth critical value data; using high-precision thermal imaging technology to perform temperature field scanning on the freezing depth critical value data to generate deep temperature distribution high-density sampling data; accurately depicting and quantifying the freezing zone boundary based on the deep temperature distribution high-density sampling data to generate a freezing range boundary map; performing interface temperature field analysis based on the freezing range boundary map to generate a three-dimensional distribution characteristic map of the freezing zone temperature field. 4The temperature data collection and report generation method for cervical cryotherapy according to claim 3 is characterized in that the method accurately depicts and quantifies the freezing zone boundary based on the deep temperature distribution high-density sampling data to generate a freezing range boundary map, including: calculating the temperature gradient vector based on the deep temperature distribution high-density sampling data to generate a temperature gradient vector field number. According to the temperature gradient vector field data, the temperature gradient vector field data includes freezing peak point distribution data and temperature reduction rate quantification data; based on the freezing peak point distribution data, low-temperature area clustering analysis is performed to generate freezing cluster area data; based on the freezing cluster area data and the temperature reduction rate quantification data, potential unfrozen areas are identified to generate freezing dead corner warning data; based on the freezing dead corner warning data, high-density sampling data of deep temperature distribution are quantitatively graded to generate a freezing range boundary map. 5.The temperature data collection and report generation method for cervical cryotherapy according to claim 4 is characterized in that step S3 includes: dynamically characterizing tissue ice crystal formation based on the three-dimensional distribution characteristic map of the freezing zone temperature field, and generating tissue ice crystal process data; analyzing the growth pattern of ice crystals inside and outside the cells based on the tissue ice crystal process data, and generating microscopic ice crystal morphological characteristic data; quantifying multi-level heat conduction effects based on the microscopic ice crystal morphological characteristic data, and generating tissue freezing response channel characteristic data. 6.The temperature data collection and report generation method for cervical cryotherapy according to claim 5 is characterized in that, the analysis of ice crystal growth patterns inside and outside the cells based on tissue ice crystal process data to generate microscopic ice crystal morphological characteristic data includes: performing microscopic freezing injury mechanism analysis based on tissue ice crystal process data to generate cell freezing mechanism characteristic data; performing tissue dehydration effect analysis based on tissue ice crystal process data to generate cell osmotic pressure change characteristic data; constructing an ice crystal growth pattern classification matrix based on cell freezing mechanism characteristic data and cell osmotic pressure change characteristic data to generate an ice crystal morphological evolution matrix; performing analysis of ice crystal growth patterns inside and outside the cells based on the ice crystal morphological evolution matrix to generate microscopic ice crystal morphological characteristic data. 7.The temperature data collection and report generation method for cervical cryotherapy according to claim 6 is characterized in that, the multi-level heat conduction effect quantification based on the microscopic ice crystal morphological characteristic data to generate tissue freezing response channel characteristic data includes: performing micro-macro freezing effect correlation analysis under the influence of tissue blood supply factors on the microscopic ice crystal morphological characteristic data to generate cross-scale freezing effect linkage characteristic data; performing multi-level heat conduction effect quantification based on the cross-scale freezing effect linkage characteristic data to generate tissue freezing response channel characteristic data. 8.The temperature data collection and report generation method for cervical cryotherapy according to claim 7 is characterized in that, step S4 includes: performing tissue damage assessment mapping under different freezing cycles on the tissue freezing response channel characteristic data to generate treatment effect depth prediction data; extracting the freezing-thawing cycle influencing factors based on the treatment effect depth prediction data to generate freezing cycle impact quantitative data, and establishing tissue necrosis prediction distribution nodes through the freezing cycle impact quantitative data; constructing an individualized freezing scheme decision model for the tissue necrosis prediction distribution nodes and the treatment effect depth prediction data to generate freezing parameter dynamic control guidance data; performing matching analysis on the three-dimensional distribution characteristic map of the freezing zone temperature field and the treatment effect depth prediction data to generate disease Lesion-treatment range fit assessment data; Based on the lesion-treatment range fit assessment data and the dynamic control guidance data of freezing parameters, a comprehensive evaluation of the cryotherapy effect is conducted to generate a treatment integrity evaluation result; Based on the treatment integrity evaluation results, basic patient information, treatment parameter records, temperature collection data and efficacy predictions are integrated to automatically generate a standardized electronic report for cervical cryotherapy. 9. A temperature data collection and report generation device for cervical cryotherapy, characterized in that the wiring harness terminal intelligent assembly device includes: a memory and at least one processor, the memory storing instructions; the at least one processor calls the instructions in the memory to enable the wiring harness terminal intelligent assembly device to execute the temperature data collection and report generation method for cervical cryotherapy as described in any one of claims 1-8. 10. A computer-readable storage medium having instructions stored thereon, wherein when the instructions are executed by a processor, the temperature data collection and report generation method for cervical cryotherapy as described in any one of claim 187 is implemented. | A cervical cryotherapy temperature acquisition and report-generation method using a multi-point thermosensitive probe array to map temperature gradients, analyze heat-diffusion paths, and locate critical points. The output includes freezing front migration data and depth criteria. | R | "capturing the freezing interface temperature gradient based on a multi-point thermal probe array to generate raw data on cervical tissue temperature distribution" (Claim 1, S1) "Map the tissue damage assessment under different freezing cycles for the tissue freezing response channel characteristic data to generate treatment effect depth prediction data; construct an individualized freezing plan decision model" (Claim 1, S4) "performing matching analysis on the three-dimensional distribution characteristic map of the freezing zone temperature field and the treatment effect depth prediction data to generate disease lesion-treatment range fit assessment data" (Claim 8) |  |  |  |  | 1 | [9 empty] | 1 |  |  |  |  | 1 | 1 | [9 empty] | 1 (cervix ) | JIANGSU HAIHUI DATONG MEDICAL TECH CO LTD | JIANGSU HAIHUI DATONG MEDICAL TECH CO LTD | 2025-06-06 | 2025 | 2025-09-12 | 2025 | 2025-06-06 | 2025 | CN202510751552 | [NA]
394 | CN119564320A | Ablation parameter determination method and device for high-frequency pulse electric field ablation | The invention discloses an ablation parameter determination method and device for high-frequency pulsed electric field ablation, and the method comprises the steps: obtaining collected medical image data containing depth information, and determining a corresponding lesion box containing the three-dimensional information of a lesion according to the medical image data; analyzing and segmenting the lesion box to obtain the contour and depth of the lesion and the contour of an electrode for performing ablation on the lesion; taking a plurality of equidistant points on the boundary of the electrode contour as starting points, respectively emitting a preset number of rays to the periphery, and obtaining the number of intersection points of the rays and the focus box to determine whether the electrode is attached to the focus; when the electrode is attached to the focus, corresponding ablation parameters are determined according to the contour and depth of the focus, and the ablation parameters comprise the pulse voltage, the pulse width, the pulse string number and the discharge frequency. The problem that in the prior art, ablation parameters are not accurately determined is solved. | 1.A method for determining ablation parameters of high-frequency pulse electric field ablation, characterized in that the method comprises: acquiring collected medical image data containing depth information, and determining a corresponding lesion box containing three-dimensional information of the lesion based on the medical image data; analyzing and segmenting the lesion box to obtain the contour and depth of the lesion and the contour of the electrode used to ablate the lesion; taking a plurality of equally spaced points on the boundary of the electrode contour as starting points, emitting a preset number of rays to the surroundings respectively, and obtaining the number of intersections between the rays and the lesion box to determine whether the electrode fits the lesion; when the electrode fits the lesion, determining the corresponding ablation parameters based on the contour and depth of the lesion, wherein the ablation parameters include pulse voltage, pulse width, number of pulse trains, and number of discharges. 2.The method for determining ablation parameters of high-frequency pulsed electric field ablation according to claim 1 is characterized in that the step of obtaining the collected medical image data containing depth information and determining the corresponding lesion box containing three-dimensional information of the lesion according to the medical image data includes: inputting the medical image data into a pre-trained lesion localization model to obtain a corresponding lesion box containing three-dimensional information of the lesion; wherein the training process of the lesion localization model is: collecting a preset number of historical medical image data, each of which corresponds to a preset frame to frame the lesion range and mark the lesion type to obtain a training data set; dividing the training data set into a training set, a validation set and a test set according to a preset ratio, and then inputting the training data set into a preset neural network for training until the model parameters tend to be stable to obtain the lesion localization model. 3.The method for determining ablation parameters of high-frequency pulse electric field ablation according to claim 1 is characterized in that the step of analyzing and segmenting the lesion box to obtain the contour and depth of the lesion and the contour of the electrode used to ablate the lesion includes: inputting the lesion box into a pre-trained lesion segmentation model to obtain the corresponding contour and depth of the lesion and the contour of the electrode used to ablate the lesion; wherein the training process of the lesion segmentation model is: collecting a preset number of historical lesion boxes and the corresponding contours and depths of the lesions and the contours of the electrodes used to ablate the lesions to obtain a training data set; dividing the training data set into a training set, a validation set and a test set according to a preset ratio, and then inputting the training data set into a preset neural network for training until the model parameters tend to be stable to obtain the lesion segmentation model. 4.The method for determining ablation parameters of high-frequency pulse electric field ablation according to claim 1 is characterized in that the step of obtaining the number of intersections between the ray and the lesion box to determine whether the electrode and the lesion are in contact with each other includes: when the number of intersections between the ray and the lesion box is an even number, it is determined that the electrode and the lesion are not in contact with each other; when the number of intersections between the ray and the lesion box is an odd number, it is determined that the electrode and the lesion are in contact with each other. 5.The method for determining ablation parameters of high-frequency pulse electric field ablation according to claim 1 is characterized in that the step of determining the corresponding ablation parameters according to the contour and depth of the lesion comprises: inputting the contour and depth of the lesion into a preset first ablation parameter determination model to obtain the corresponding ablation parameters; wherein the training process of the first ablation parameter determination model is: collecting historical ablation parameters and the corresponding ablation depth and ablation length as a training data set, and inputting the training data into a first ablation parameter determination model; The data set is input into the naive Bayesian network for training to obtain the first ablation parameter determination model; the naive Bayesian network is based on the Bayesian probability formula, and the optimal relationship between the ablation depth, ablation length and ablation parameters is adjusted through multiple training; wherein, the mathematical expression of the Bayesian formula is: (4)·Ρ(Α) ~Ρ(^)~ wherein, P (A IB) represents the probability of event 4 occurring under the condition that event 8 occurs, P (BlΑ) represents the probability of event 8 occurring under the condition that event 4 occurs, Ρ (Α) is the prior probability of event 4 occurring independently, Ρ (Β) is the marginal probability of event 8 occurring, event Α= {pulse voltage, pulse width, number of pulse trains, number of discharges}, event 8= {ablation length, ablation depth}. 6.The method for determining ablation parameters of high-frequency pulse electric field ablation according to claim 1 is characterized in that, after the step of determining the corresponding ablation parameters according to the contour and depth of the lesion when the electrode is in contact with the lesion, it also includes: after completing an ablation discharge using the ablation parameters, obtaining the impedance data, contour and depth of the lesion within a preset time period after the end of this ablation discharge; inputting the impedance data and the contour and depth of the lesion after this ablation discharge into a pre-trained second ablation parameter determination model to obtain the corresponding secondary ablation parameters. 7.The ablation parameter determination method of high-frequency pulse electric field ablation according to claim 6 is characterized in that the training process of the second ablation parameter determination model is: collecting historical impedance data, secondary ablation parameters and corresponding ablation depth and ablation length as a training data set, and inputting the training data set into the Markov network for training to obtain the second ablation parameter determination model; the Markov network is based on the Bayesian probability formula, and adjusts the optimal relationship between the ablation depth, ablation length and secondary ablation parameters through multiple training; wherein, the mathematical expression of the Bayesian formula is: wherein /(4/8, represents the probability of event 4, / occurring under the condition that event 8 occurs /6/4,. represents the probability of event I occurring under the condition that event 4 occurs / (4 represents the probability of event I occurring under the condition that event 4 occurs / (B11) represents the probability of event / occurring under the condition that event 8 occurs, event Α= {pulse voltage, pulse width, number of pulse trains, number of discharges}, event 8= {ablation length, ablation depth}, / is impedance data. 8.A device for determining ablation parameters of high-frequency pulse electric field ablation, characterized in that the device comprises: an acquisition module, used to acquire collected medical image data containing depth information, and determine the corresponding lesion box containing three-dimensional information of the lesion according to the medical image data; a segmentation module, used to analyze and segment the lesion box to obtain the contour and depth of the lesion and the contour of the electrode used to ablate the lesion; a collision module, used to take a plurality of equally spaced points on the boundary of the electrode contour as starting points, respectively emit a preset number of rays to the surroundings, and obtain the number of intersections between the rays and the lesion box to determine whether the electrode and the lesion are in contact; a determination module, used to determine the corresponding ablation parameters according to the contour and depth of the lesion when the electrode and the lesion are in contact; Or it is used to determine corresponding ablation parameters according to the contour and depth of the lesion when the electrode is in contact with the lesion, and to obtain the impedance data as well as the contour and depth of the lesion after the corresponding ablation parameters complete one ablation discharge, and to determine corresponding secondary ablation parameters based on the impedance data and the contour and depth of the lesion; wherein the ablation parameters and the secondary ablation parameters both include pulse voltage, pulse width, number of pulse trains and number of discharges. 9.A readable storage medium having a computer program stored thereon, characterized in that when the program is executed by a processor, the steps of the method as claimed in any one of claims 1 to 7 are implemented. 10. An electronic device, characterized in that it includes a memory, a processor, and a computer program stored in the memory and executable on the processor, wherein when the processor executes the program, the steps of the method described in any one of claims 1 to 7 are implemented. | A high-frequency PFA parameter-determination method that pulls 3D depth-bearing medical images, builds a focus box around the lesion, and analyzes lesion outline/depth and electrode outline. Geometry from imaging then drives the choice of ablation parameters. | R | "acquiring collected medical image data containing depth information, and determining a corresponding lesion box containing three-dimensional information of the lesion" (Claim 1) "taking a plurality of equally spaced points on the boundary of the electrode contour as starting points, emitting a preset number of rays to the surroundings… to determine whether the electrode fits the lesion" (Claim 1) "after completing an ablation discharge using the ablation parameters, obtaining the impedance data, contour and depth of the lesion within a preset time period after the end of this ablation discharge; inputting the impedance data and the contour and depth of the lesion after this ablation discharge into a pre-trained second ablation parameter determination model" (Claim 6) | 1 | [7 empty] | 1 |  |  | 1 | [9 empty] | 1 | [9 empty] | JIANGXI SHUNJIE MEDICAL TECH CO LTD | JIANGXI SHUNJIE MEDICAL TECH CO LTD | 2024-12-18 | 2024 | 2025-03-07 | 2025 | 2024-12-18 | 2024 | CN202411869784 | CN119564320B
395 | US12478427B2 | Contact pressure assessment for cryoballoon ablation catheters | A method of ablating tissue includes positioning a treatment device proximate to a target tissue area. The treatment device has an expandable treatment element. The expandable treatment element is inflated with a refrigerant during an inflation phase such that at least a portion of the expandable treatment element is in contact with the target tissue area. A first pressure measurement of the inflated expandable treatment element is recorded and compared to a predetermined pressure threshold. The refrigerant is circulated within the expandable treatment element during an ablation phase to reduce a temperature of the target tissue area to a temperature sufficient to cryoablate the target tissue area. A second pressure measurement of the expandable treatment element is recorded during the ablation phase and compared to the predetermined pressure threshold. | 1. A method of ablating tissue, the method comprising: positioning a treatment device proximate to a target tissue area, the treatment device having an expandable treatment element; inflating the expandable treatment element with a refrigerant during an inflation phase such that at least a portion of the expandable treatment element is in contact with the target tissue area; recording a first pressure measurement within the inflated expandable treatment element during the inflation phase and comparing the first pressure measurement to a predetermined pressure threshold; circulating the refrigerant within the expandable treatment element during an ablation phase to reduce a temperature of the target tissue area to a temperature sufficient to cryoablate the target tissue area; recording a second pressure measurement within the inflated expandable treatment element during the ablation phase and comparing the second pressure measurement to the predetermined pressure threshold; calculating a difference between the first pressure measurement and the predetermined pressure threshold; and assigning a first contact score based on the calculated difference, the assigned first contact score being indicative of a quality of tissue contact by the expandable treatment element during the inflation phase. 2. The method of claim 1, further comprising: calculating a difference between the second pressure measurement and the predetermined pressure threshold; and assigning a second contact score based on the calculated difference between the second pressure measurement and the predetermined pressure threshold, the assigned second contact score being indicative of a quality of tissue contact by the expandable treatment element during the ablation phase. 3. The method of claim 2, further including: determining whether undue force is being applied to the treatment device during at least one of the ablation phase and the inflation phase based, in part, on at least one of the first contact score and the second contact score. 4. The method of claim 3, further comprising: generating an alert if it is determined that undue force is being applied to the treatment device during at least one of the ablation phase and the inflation phase. 5. The method of claim 3, further comprising at least one selected from the group consisting of: adjusting a flow rate of the refrigerant if it is determined that undue force is being applied to the treatment device during at least one of the ablation phase and the inflation phase; and discontinuing the circulation of the refrigerant if it is determined that undue force is being applied to the treatment device during at least one of the ablation phase and the inflation phase. 6. The method of claim 2, further including: determining whether the second contact score is indicative of inadequate tissue contact; generating an alert if the second contact score indicates inadequate tissue contact during the ablation phase; and at least one of adjusting a flow rate of the refrigerant and discontinuing the circulation of the refrigerant if the second contact score indicates inadequate tissue contact during the ablation phase. 7. A system for ablating tissue, comprising: a treatment device including an expandable treatment element and a pressure monitoring tube having a proximal portion and an opposite distal portion, the distal portion being partially disposed within the expandable treatment element; a control unit in communication with the treatment device, the control unit including processing circuitry configured to: initiate and transition between an inflation phase and an ablation phase; record a first pressure measurement of the expandable treatment element during the inflation phase and compare the first pressure measurement to a predetermined pressure threshold; record a second pressure measurement of the expandable treatment element during the ablation phase and compare the second pressure measurement to the predetermined pressure threshold; calculate a difference of the first pressure measurement and the second pressure measurement from the predetermined pressure threshold; assign a first contact score based on a comparison of the calculated difference between the first pressure measurement and the predetermined pressure threshold to a predefined table; and assign a second contact score based on a comparison of the calculated difference between the second pressure measurement and the predetermined pressure threshold to the predefined table. 8. The system of claim 7, wherein the control unit further includes: a fluid supply reservoir in fluid communication with the expandable treatment element; and a pressure sensor configured to monitor a pressure of the expandable treatment element. 9. The system of claim 8, wherein the inflation phase includes inflating the expandable treatment element such that the expandable treatment element is in contact with a target tissue area. 10. The system of claim 9, wherein the ablation phase includes delivering the refrigerant from the fluid supply reservoir to the expandable treatment element to treat the target tissue area. 11. The system of claim 7, wherein the processing circuitry is further configured to: determine whether undue force is being applied to the treatment device during at least one of the ablation phase and the inflation phase based, in part, on at least one of the first contact score and the second contact score. 12. The system of claim 11, wherein the processing circuitry is further configured to generate an alert if it is determined that undue force is being applied to the treatment device during at least one of the ablation phase and the inflation phase. 13. The system of claim 12, wherein the processing circuitry is further configured to perform one operation selected from the group consisting of: adjust a flow rate of the refrigerant if undue force is being applied to the treatment device during the at least one of the ablation phase and the inflation phase; and discontinue the delivery of the refrigerant if undue force is being applied to the treatment device during the at least one of the ablation phase and the inflation phase. 14. The system of claim 11, wherein the processing circuitry is further configured to determine whether at least one of the first contact score and the second contact score is indicative of inadequate tissue contact. 15. The system of claim 14, wherein the processing circuitry is further configured to: generate an alert if the at least one of the first contact score and the second contact score indicates inadequate tissue contact during at least one of the ablation phase and the inflation phase; and at least one of adjust a flow rate of the refrigerant and discontinue the circulation of the refrigerant if the at least one of the first contact score and the second contact score indicates inadequate tissue contact during at least one of the ablation phase and the inflation phase. 16. A system for ablating tissue, comprising: a treatment device including an expandable treatment element and a pressure monitoring tube having a proximal portion and an opposite distal portion, the distal portion being partially disposed within the expandable treatment element; a control unit in communication with the treatment device, the control unit including: a fluid supply reservoir in fluid communication with the expandable treatment element; a pressure sensor configured to monitor a pressure of the expandable treatment element; and processing circuitry configured to: initiate and transition between an inflation phase and an ablation phase, the inflation phase including inflating the expandable treatment element with a refrigerant such that the expandable treatment element is in contact with a target tissue area, the ablation phase including delivering the refrigerant from the fluid supply reservoir to the expandable treatment element to treat the target tissue area; record a first pressure measurement of the expandable treatment element during the inflation phase; calculate a difference between the first pressure measurement and a predetermined pressure threshold; assign a first contact score based on a comparison of the calculated difference between the first pressure measurement and the predetermined pressure threshold to a predefined scale; record a second pressure measurement of the expandable treatment element during the ablation phase; calculate a difference between the second pressure measurement and the predetermined pressure threshold; assign a second contact score based on a comparison of the calculated difference between the second pressure measurement and the predetermined pressure threshold to a predefined table; determine whether undue force is being applied to the treatment device during at least one of the ablation phase and the inflation phase based, in part, on at least one of the first contact score and the second contact score; and generate an alert if it is determined that undue force is being applied to the treatment device during at least one of the ablation phase and the inflation phase. | A cryoballoon ablation method that records pressure of the inflated balloon and compares it to a threshold to assess tissue contact. The pressure-based contact check helps confirm the balloon is firmly seated against the target before ablating. | R | "recording a first pressure measurement within the inflated expandable treatment element during the inflation phase and comparing the first pressure measurement to a predetermined pressure threshold" (Claim 1) "calculating a difference between the first pressure measurement and the predetermined pressure threshold; and assigning a first contact score based on the calculated difference, the assigned first contact score being indicative of a quality of tissue contact" (Claim 1) "determining whether undue force is being applied to the treatment device during at least one of the ablation phase and the inflation phase based, in part, on at least one of the first contact score and the second contact score" (Claim 3) "adjusting a flow rate of the refrigerant if it is determined that undue force is being applied… discontinuing the circulation of the refrigerant" (Claim 5) | [5 empty] | 1 | [8 empty] | 1 | [5 empty] | 1 | [10 empty] | MEDTRONIC CRYOCATH LP | MEDTRONIC CRYOCATH LP | 2022-07-22 | 2022 | 2025-11-25 | 2025 | 2021-08-10 | 2021 | US202217814405 | CA3228706A1¦CN117693316A¦US20230053149A1¦WO2023015377A1¦WO2023015377A8
396 | US20240374933A1 | TISSUE TREATMENT SYSTEMS AND METHODS WITH ACOUSTIC DOSE MANAGEMENT | A tissue treatment system may comprise a catheter including a distal portion on which is located an ultrasound transducer, where at least the distal portion of the catheter is insertable into a segment of a body lumen having a diameter within a range of diameters having a lower subrange and an upper subrange; an excitation source that provides energy to the ultrasound transducer; and a controller that controls the excitation source to cause the ultrasound transducer to emit at least a first amount of acoustic energy and a second amount of acoustic energy greater than the first amount of acoustic energy; where the controller controls the excitation source to cause the ultrasound transducer to emit the first amount of acoustic energy into a body lumen having a diameter within the lower subrange, and the second amount of acoustic energy into a body lumen having a diameter within the upper subrange. | 1. A tissue treatment system, comprising: a catheter including a distal portion on which is located an ultrasound transducer, wherein the catheter is configured such that at least the distal portion of the catheter is insertable into a segment of a body lumen having a diameter within a specified range of diameters, wherein the range of diameters has a lower subrange and an upper subrange; an excitation source configured to selectively provide energy to the ultrasound transducer of the catheter; and a controller communicatively coupled to the excitation source, the controller configured to control the excitation source to cause the ultrasound transducer to emit two different amounts of acoustic energy, which include a first amount of acoustic energy; and a second amount of acoustic energy, which is greater than the first amount of acoustic energy; wherein the controller is configured to control the excitation source to cause the ultrasound transducer to emit the first amount of acoustic energy into a body lumen having a diameter within the lower subrange, and the second amount of acoustic energy into a body lumen having a diameter within the upper subrange. 2. The system of claim 1, wherein the controller is further configured to: automatically estimate a diameter of the segment of the body lumen; control the excitation source to cause the ultrasound transducer to emit the first amount of acoustic energy when the controller estimated a diameter of the segment of the body lumen to be within the lower subrange of the specified range of diameters; and control the excitation source to cause the ultrasound transducer to emit the second amount of acoustic energy when the controller estimated a diameter of the segment of the body lumen to be within the upper subrange of the specified range of diameters. 3. The system of claim 2, wherein the system further comprises a user interface connected to the controller, and the controller is further configured to: receive a selection from the user interface of one of the lower subrange of the specified range of diameters and the upper subrange of the specified range of diameters, and control the excitation source to cause the ultrasound transducer to emit the first amount of acoustic energy when the selection is the lower subrange of the specified range of diameters; and control the excitation source to cause the ultrasound transducer to emit the second amount of acoustic energy when the selection is the upper subrange of the specified range of diameters. 4. The system of claim 1, wherein the distal portion of the catheter further includes a balloon within which is located the ultrasound transducer; and the balloon is configured to generally center the ultrasound transducer within the body lumen and to have a fluid circulated through the balloon to cool at least a portion of tissue adjacent to the body lumen within which the ultrasound transducer is positioned. 5. The system of claim 4, wherein the balloon comprises a compliant balloon, wherein the compliant balloon is configured such that when the compliant balloon is partially inflated, such that its diameter is less than a nominal balloon diameter of the compliant balloon, the compliant balloon includes one or more folds that at least partially attenuate the acoustic energy emitted by the ultrasound transducer. 6. The system of claim 5, wherein a specified intermediate diameter separates the lower subrange from the upper subrange, and wherein the nominal balloon diameter is larger than the specified intermediate diameter by an amount selected from the group consisting of at least 10%, at least 20%, and at least 30%. 7. The system of claim 6, wherein the one or more folds are configured to occur such that when the balloon is inserted into a body lumen segment and partially inflated to less than its nominal balloon diameter, a folded surface of the balloon generates additional acoustic reflections compared to when the balloon is inflated to the point that there are no, or less, or smaller folds. 8. The system of claim 4, further comprising a fluid supply subsystem configured to circulate the fluid through the balloon, wherein the controller is also configured to control the fluid supply subsystem. 9. The system of claim 8, wherein an amount of energy absorbed by tissue surrounding the segment of the body lumen in which the ultrasound transducer is located depends in part on a flowrate of the fluid circulated through the balloon, and wherein the controller is configured to control the flowrate of the fluid circulated through the balloon to be within a flowrate range of about 5 mL/min to about 40 mL/min, in particular about 10 mL/min to about 15 mL/min. 10. The system of claim 8, wherein the controller is configured to control the fluid supply system to circulate the fluid through the balloon, after the first amount of energy has been emitted by the ultrasound transducer, for a predetermined amount of time within a range of 0.5 seconds to 20 seconds. 11. The system of claim 1, wherein: a lower end of the specified range of diameters is selected from the group consisting of: about 2 mm; about 2.5 mm; and about 3.0 mm; and an upper end of the specified range of diameters is selected from the group consisting of: about 7.5 mm; about 8.0 mm; and about 8.5 mm. 12. The system of claim 1, wherein the second amount of acoustic energy is selected from the group consisting of at least 8%, at least 10%, at least 12%, and at least 15% greater than the first amount of acoustic energy. 13. The system of claim 1, wherein the first amount of acoustic energy is based on a first acoustic signal having a first acoustic frequency, a first acoustic power and a first duration; and the second amount of acoustic energy is based on an second acoustic signal having a second acoustic frequency, a second acoustic power and a second duration; wherein the second acoustic power is greater than the first acoustic power. 14. The system of claim 13, wherein the second acoustic power is selected from the group consisting of at least 8%, at least 10%, at least 12% and at least 15% greater than the first acoustic power. 15. The system of claim 13, wherein the first acoustic frequency equals the second acoustic frequency. 16. The system of claim 13, wherein the first duration equals the second duration. 17. The system of claim 13, wherein a total tolerance range for generating the first and second acoustic powers is selected from the group consisting of: below 15%, below 11% and below 8%. 18. The system of claim 1, wherein: the first amount of acoustic energy is obtained by a first acoustic power of about 30.1 or 32.0 watts; and the second amount of acoustic energy is obtained by a second acoustic power of about 36.0 watts. 19. The system of claim 1, wherein the system is configured to produce at least one lesion in tissue surrounding the body lumen in which the ultrasound transducer is positioned, such that a lesion depth of the one or more lesions is within a lesion depth range of about 2 mm to about 10 mm. 20. A method for use with a tissue treatment system having a catheter that includes a distal portion on which is located an ultrasound transducer, the method comprising: inserting the distal portion of the catheter into a segment of a body lumen having a diameter within a specified range of diameters that is at least 4 mm, such that the ultrasound transducer is located within the segment of the body lumen having the diameter within the specified range of diameters that is at least 4 mm; receiving input, at a controller associated with the ultrasound transducer, whether the diameter of the segment of the body lumen, within which the ultrasound transducer is located, is within a lower subrange of the specified range of diameters or an upper subrange of the specified range of diameters; upon receiving input that the diameter of the segment of the body lumen is within the lower subrange of the specified range of diameters, causing the ultrasound transducer to emit a first amount of acoustic energy; and upon receiving input that the diameter of the segment of the body lumen is within the upper subrange of the specified range of diameters, causing the ultrasound transducer to emit a second amount of acoustic energy, which is greater than the first amount of acoustic energy. 21. The method of claim 20, wherein the receiving input comprises: accepting an indication, from a user via a user interface connected to the controller, of which one of a plurality of different types of body lumens the ultrasound transducer is located; and automatically categorizing whether the diameter of the segment of the body lumen, within which the ultrasound transducer is located, is within the lower subrange of the specified range of diameters or the upper subrange of the specified range of diameters, based on the indication accepted via the user interface. 22. The method of claim 20, wherein the receiving input comprises receiving input from the controller, which is configured to automatically determine whether the diameter of the segment of the body lumen, within which the ultrasound transducer is located, is within the lower subrange of the specified range of diameters or the upper subrange of the specified range of diameters. 23. The method of claim 20, wherein the distal portion of the catheter further includes a compliant balloon within which is located the ultrasound transducer, wherein the compliant balloon is configured such that when the compliant balloon is partially inflated, to a diameter less than a nominal balloon diameter of the compliant balloon, the compliant balloon includes one or more folds; the method further comprising: causing inflation of the compliant balloon so that the compliant balloon is in apposition with a body lumen wall of a segment of the body lumen within which is located the ultrasound transducer; utilizing the one or more folds in the compliant balloon to at least partially attenuate some of the acoustic energy emitted by the ultrasound transducer, and thereby reduce an amount of the acoustic energy that passes through the compliant balloon when the diameter of the segment of the body lumen in which the compliant balloon is in apposition is within a smaller diameter subset of the specified range of diameters, compared to when the compliant balloon is inflated such that its diameter is at least the nominal balloon diameter of the compliant balloon and the diameter of the segment of the body lumen in which the compliant balloon is in apposition with is within a larger diameter subset of the specified range of diameters. 24. The method of claim 23, wherein the compliant balloon has a balloon wall having a nominal balloon wall thickness; and wherein the causing inflation comprises causing inflation beyond the nominal balloon diameter; and wherein the causing inflation beyond the nominal balloon diameter causes the balloon wall to stretch and become thinner than the nominal balloon wall thickness, wherein the thinner balloon wall provides less attenuation of the acoustic energy emitted by the ultrasound transducer than the nominal balloon wall thickness. 25. The method of claim 20, wherein the body lumen is one of a vein; a main renal artery, an accessory renal artery, a renal artery branch, a pulmonary artery; a vascular lumen; a celiac artery; a common hepatic artery; a proper hepatic artery; a gastroduodenal artery; a hepatic artery; a splenic artery; a gastric artery; a blood vessel; a nonvascular lumen; an airway; a sinus; an esophagus; a respiratory lumen; a digestive lumen; a stomach; a duodenum; a jejunum; or a cancer tissue. 26. A tissue treatment system, comprising: a catheter including a distal portion configured to deliver neuromodulation energy to a segment of a body lumen; an excitation source configured to selectively provide neuromodulation energy to the catheter; and a controller communicatively coupled to the excitation source, the controller configured to control the excitation source to cause the catheter to emit about a same amount of neuromodulation energy when a diameter of the segment of the body lumen is within a first specified range of diameters that is at least 4 mm. 27. The system of claim 26, wherein the controller is further configured to control the excitation source to cause the catheter to emit a different amount of neuromodulation energy when the diameter of the segment of the body lumen is within a second specified range of diameters that is at least 1 mm. 28. The system of claim 27, wherein: the first specified range of diameters, that is at least 4 mm, includes diameters between about 4.1 mm to about 8.0 mm; and the second specified range of diameters, which is at least 1 mm, includes diameters between about 2.0 mm to about 4.0 mm. 29. The system of claim 26, wherein: the distal portion configured to deliver neuromodulation energy to a segment of a body lumen comprises an ultrasound transducer, wherein the neuromodulation energy is acoustic energy; an amount of acoustic energy emitted by the ultrasound transducer is based on an output power of the excitation source that drives the ultrasound transducer, a frequency of an acoustic signal emitted by the ultrasound transducer, and a duration of the acoustic signal emitted by the ultrasound transducer; and the controller is configured to control the excitation source to cause the ultrasound transducer to emit the same amount of acoustic energy, when the diameter of the segment of the body lumen is within the first specified range of diameters, by controlling the excitation source so that the output power is about the same, the frequency of the acoustic signal emitted by the ultrasound transducer is about the same, and the duration of the acoustic signal emitted by the ultrasound transducer is about the same, when the diameter of the segment of the body lumen is within the first specified range of diameters that is at least 4 mm. 30. The system of claim 29, wherein: the distal portion of the catheter further includes a balloon within which is located the ultrasound transducer; and the balloon is configured to generally center the ultrasound transducer within the body lumen and to receive a fluid circulated through the balloon to cool at least a portion of tissue adjacent to the body lumen within which the ultrasound transducer is positioned. 31. The system of claim 30, further comprising a fluid supply subsystem configured to circulate the fluid through the balloon, and wherein: the controller is also configured to control the fluid supply subsystem; an amount of energy absorbed by tissue surrounding the segment of the body lumen in which the ultrasound transducer is located depends in part on a flowrate of the fluid circulated through the balloon; and the controller is also configured to control the fluid supply subsystem so that the flowrate of the fluid circulated through the balloon is about the same when the diameter of the segment of the body lumen is within the specified range of diameters that is at least 4 mm. 32. The system of claim 30, wherein the balloon is a compliant balloon having a Shore D durometer between 50 and 60. | A tissue-treatment system with an ultrasound transducer at the catheter's distal end whose excitation source is varied to deliver controlled doses, sized for body lumens within a defined diameter range. Acoustic dose management keeps energy delivery safe across vessel sizes. | R | "the controller is configured to control the excitation source to cause the ultrasound transducer to emit the first amount of acoustic energy into a body lumen having a diameter within the lower subrange, and the second amount of acoustic energy into a body lumen having a diameter within the upper subrange" (Claim 1) "automatically estimate a diameter of the segment of the body lumen" (Claim 2) "produce at least one lesion in tissue surrounding the body lumen in which the ultrasound transducer is positioned, such that a lesion depth of the one or more lesions is within a lesion depth range of about 2 mm to about 10 mm" (Claim 19) "the body lumen is one of a vein; a main renal artery, an accessory renal artery, a renal artery branch, a pulmonary artery…" (Claim 25) |  |  | 1 | [12 empty] | 1 |  |  |  |  | 1 |  |  |  |  | 1 | [5 empty] | OTSUKA MEDICAL DEVICES CO LTD | OTSUKA MEDICAL DEVICES CO., LTD. | 2024-05-13 | 2024 | 2024-11-14 | 2024 | 2023-05-12 | 2023 | US202418662757 | EP4709304A1¦TW202508669A¦TWI902250B¦WO2024236449A1
397 | CN113994380A | METHOD FOR DETERMINING AN ABLATION REGION BASED ON DEEP LEARNING | The invention relates to a method for evaluating in post-treatment an ablation of a portion of an anatomy of interest (130) of an individual (110), the anatomy of interest comprising at least one lesion (165). The evaluation method comprises in particular a step of automatically determining a contour of the ablation region by means of an automatic learning method, such as a neural network, analyzing the post-treatment image of the anatomy of interest of the individual, said automatic learning method being preloaded (290) during a so-called training phase using a database comprising a plurality of post-operative medical images of an anatomy of identical interest of a set of patients, each medical image of the database being associated with an ablation region of the anatomy of interest of said patient. The invention also relates to an electronic device (150) comprising a processor and a computer memory storing instructions of such an evaluation method. | 1A post-treatment assessment method (200) for ablation of a portion of a target anatomy (130) of an individual (110), the target anatomy comprising at least one lesion (165, 510, 600), a portion of the target anatomy The ablation is defined by an ablation zone (160), characterized in that the evaluation method comprises the steps of: acquiring (210) a post-operative medical image of the target anatomy of the individual, including all or part of the ablation zone; The machine learning method automatically determines the contour of the ablation area, analyzes the individual's post-treatment images of the target anatomy, pre-trained in the training phase using a database of multiple post-operative medical images of the same target anatomy for a set of patients (290) In the machine learning method, each medical image in the database is associated with an ablation region of a target anatomy of the patient. 2. The post-treatment evaluation method of claim 1, wherein the training phase includes a preceding step of training using medical images of the same target anatomy containing unablated lesions. 3. The post-treatment assessment method according to any one of the preceding claims, further comprising the step (220) of registering a post-operative image of a target anatomy of the individual with a medical image (170, 500), the medical Images obtained before surgical treatment are called preoperative medical images. 4. The post-treatment assessment method of any preceding claim, further comprising assessing between the ablation area and the lesion, between the ablation area and the target anatomy, or between the lesion and the lesion Steps to assess recurrence risk by relative characteristics between target anatomical structures. 5. A post-treatment assessment method according to any preceding claim, further comprising, when a risk of recurrence is demonstrated, according to between the ablation area and the lesion, the ablation area and the target anatomy The step of determining the location of recurrence based on relative characteristics between the lesion or between the lesion and the target anatomy. 6. The post-treatment evaluation method according to any one of claims 4 and 5, wherein the relative feature is the ablation margin between the ablation area and the lesion. 7. The post-treatment evaluation method according to any one of claims 4 to 6, wherein the relative features are the centroid of the lesion and the centroid of the ablation area. 8. The post-treatment evaluation method of any one of claims 4 to 7, wherein the relative characteristics are the uniformity and sharpness of the edges of the ablation zone relative to surrounding healthy tissue. 9. The post-treatment evaluation method according to any one of claims 4 to 8, wherein the relative feature is the ratio of the volume of the lesion to the volume of the ablation area. 10. The post-treatment assessment method of any one of claims 4 to 9, wherein the relative feature is the location of the lesion relative to the center of the target anatomy. 11. The post-treatment assessment method of any one of claims 3 to 10, further comprising the step of segmenting the lesion in a preoperative medical image of the subject's target anatomy. 12. The post-treatment assessment method of any one of claims 3 to 11, further comprising the step of detecting the lesion in a preoperative medical image of the individual target anatomy. 13. The post-treatment evaluation method according to any one of claims 3 to 12, wherein all or part of the training post-operative medical images in the database are cropped around an ablation region comprising at least one lesion, the cropping of the image being cropped This is performed using a common frame of predetermined dimensions, within which the collection of ablation region centers in the cropped post-operative medical image forms clusters of distinct points. 14. The post-treatment assessment method according to any one of claims 3 to 13, wherein, for the set of medical images in the database, the parts of an individual's body included in the images are divided into a plurality of single-sized basic unit, The number of base cells is divided into two nearly equal parts between the part of the human body bounded by the ablation area and the rest of the individual's body included in the image. 15. The post-treatment evaluation method according to any one of claims 3 to 14, wherein the post-operative medical image database comprises at least one pre-operative medical image containing at least one unablated lesion. 16. The post-treatment assessment method of any one of claims 4 to 15, further comprising the step of determining a supplemental ablation mask when recurrence risk is demonstrated. 17. A post-treatment assessment method according to any preceding claim, further comprising the step of proposing a path to be followed by a medical instrument to a target point of the supplemental ablation mask. 18. A post-treatment assessment method according to any preceding claim, wherein the medical image is a three-dimensional image. 19. A post-treatment assessment method according to any one of claims 3 to 18, wherein each post-operative image is acquired using the same image acquisition technique. 20. An electronic device (150) comprising a processor and a computer memory storing instructions for a post-treatment assessment method according to any preceding claim. 21. twenty one. An electronic device according to the preceding claim, wherein the device may be a control device, a navigation system, a robotic device or an augmented reality device. | A deep-learning method that automatically determines ablation-region contours from post-treatment images using a trained neural network. By comparing pre- and post-treatment images, it produces a quantitative post-ablation assessment. | R | The present invention is particularly useful for evaluating minimally invasive medical interventions and predicting the risk of recurrence of lesions, such as tumor or metastases. This minimally invasive medical intervention corresponds, for example, to percutaneous ablation of lesions, eg tumors in the liver, lungs, kidneys or any other organ. Percutaneous ablation typically involves image-guided insertion of one or more needles through the skin to reach and destroy the lesion. | [8 empty] | 1 | [10 empty] | 1 |  | 1 | [6 empty] | 1 |  | 1 | QUANTUM SURGICAL | QUANTUM SURGICAL | 2021-05-20 | 2021 | 2022-01-28 | 2022 | 2020-05-20 | 2020 | CN20218003924 | CA3176333A1¦CN113994380B¦EP4154263A1¦FR3110763A1¦FR3110763B1¦IL298313A¦JP2023526909A¦JP7834030B2¦KR20230013041A¦US20230177681A1¦WO2021234305A1
398 | US20240001487A1 | MEDICAL TREATMENT APPARATUS AND TREATMENT PROBE THEREOF | Disclosed is a medical treatment apparatus. The medical treatment apparatus includes a magnetic resonance imaging (MRI) device, configured for imaging of a specific region including a target object and generating a magnetic resonance image; a laser interstitial thermal therapy (LITT) device, including an LITT probe, the LITT probe being positioned close to the target object based on the magnetic resonance image, and being configured to treat the target object by emitting laser; and a temperature measurement element, the temperature measurement element and the LITT probe being integrated as an integrated probe, and the temperature measurement element being configured to obtain a temperature of the target object. | 1. A laser interstitial thermal therapy (LITT) lateral ablation probe, which is used in an LITT device, wherein the LITT lateral ablation probe comprises: a probe main body; and a coating layer covered on an end surface of the probe main body, wherein the end surface of the probe main body and an axis of the probe main body form a first angle, the probe main body comprises a probe core and a hard cladding structure located at an outer peripheral of the probe core, the probe core is made of pure silicon oxide material, the hard cladding structure is made of technology enhanced clad silica (TECS) material, and the coating layer comprises noble metal. 2. The LITT lateral ablation probe of claim 1, wherein the coating layer has a dual-layer structure, a layer of the dual-layer structure which is close to the end surface of the probe main body comprises pure silver, and the other layer of the dual-layer structure which is away from the end surface of the probe main body comprises silicon monoxide. 3. The LITT lateral ablation probe of claim 1, wherein the LITT lateral ablation probe is obtained by brightening and cleaning the end surface of the probe main body, and then further covering the coating layer on the end surface through magnetron sputtering. 4. A laser interstitial thermal therapy (LITT) lateral ablation probe, which is used in an LITT device, wherein the LITT lateral ablation probe comprises: a probe main body; and a lens connected to an end surface of the probe main body, wherein the end surface of the probe main body is perpendicular to an axis of the probe main body, the lens has a first vertex angle, the probe main body comprises a probe core and a hard cladding structure located at an outer peripheral of the probe core, and the lens is a sapphire lens. 5. The LITT lateral ablation probe of claim 4, wherein the lens is a chamfered cylindrical lens, or a chamfered hemi-spherical or hemi-ellipsoidal lens. 6. The LITT lateral ablation probe of claim 4, wherein the LITT lateral ablation probe is obtained by brightening and cleaning the end surface of the probe main body, connecting the probe main body and the lens through fusion welding, and processing a welded surface between the probe main body and the lens using fire polishing. 7. A laser interstitial thermal therapy (LITT) circumferential ablation probe, which is used in an LITT device, wherein the LITT circumferential ablation probe comprises: a probe main body, the main body of the probe comprising a cone-shaped surface, one or more grooves being set on the cone-shaped surface, the one or more grooves being evenly distributed on the cone-shaped surface in an arabesquitic pattern, wherein the LITT circumferential ablation probe emits laser out of the one or more grooves. 8. The LITT circumferential ablation probe of claim 7, wherein a diameter of the cone-shaped surface decreases from an initial diameter to a preset diameter, the initial diameter is 630 μm, and the preset diameter is 100 μm. 9. The LITT circumferential ablation probe of claim 7, wherein the arabesquitic pattern comprises a single thread pattern, a cross-thread pattern, or a rhombic grid pattern, or a combination thereof. 10. The LITT circumferential ablation probe of claim 7, wherein the cone-shaped surface is formed by performing operations comprising: fixing two ends of a first workpiece with cone-shaped surface to be carved on a sliding rail and a fixing device, respectively; controlling, by a laser device controller, a laser device, to emit laser, wherein a power of the emitted laser is 30 W, a wavelength of the emitted laser is 10600 nm, the emitted laser is attenuated to 6.6-11.2 W by a laser power attenuator, the attenuated laser is split into two laser beams by a diffraction splitting lens unit, and the two laser beams have equal power and are focused onto a surface of the first workpiece by a focusing lens unit; and driving, by a motion driver, the first workpiece to move, the motion driver being controlled by a motion driver controller. 11. The LITT circumferential ablation probe of claim 7, wherein the one or more grooves that are evenly distributed on the cone-shaped surface in the arabesquitic pattern are formed by performing operations comprising: fixing two ends of a second workpiece with arabesquitic pattern to be carved on a sliding rail and a fixing device, respectively; controlling, by a laser device controller, a laser device, to emit laser, wherein a power of the emitted laser is 30 W, a wavelength of the emitted laser is 10600 nm, the emitted laser is focused on a surface of the second workpiece by a lens set, the lens set comprises a first concave lens, a first convex lens, and a second concave lens, the first concave lens and the first convex lens are used to expand a diameter of the emitted laser into a first diameter, the second concave lens is used to focus the laser beam of the first diameter to form a spot on the surface of the second workpiece, and the spot has a diameter of 45 μm; and driving, by a motion driver, the second workpiece to move, the motion driver being controlled by a motion driver controller. 12. The LITT circumferential ablation probe of claim 7, wherein a spatial intensity distribution of the laser emitted by the LITT circumferential ablation probe is obtained through operations comprising: obtaining signals output by a laser measurement sensor, and generating a polar intensity of a specific orientation of the laser emitted by the LITT circumferential ablation probe, wherein the LITT circumferential ablation probe is connected to a helium-neon laser device that generates laser with a wavelength of 632.8 nm, a slit diaphragm is set between the laser measurement sensor and the LITT circumferential ablation probe, and a slit size of the slit diaphragm is 0.3 mm; and determining the spatial intensity distribution of the laser emitted by the LITT circumferential ablation probe by driving, by a motion driver, the slit diaphragm and the laser measurement sensor to move along a circumferential direction and an axial direction of the cone-shape surface of the LITT circumferential ablation probe, the motion driver being controlled by a motion driver controller, wherein the slit diaphragm and the laser measurement sensor are fixedly connected to the motion driver. 13. A medical treatment apparatus, comprising: a magnetic resonance imaging (MRI) device, configured for imaging of a specific region comprising a target object and generating a magnetic resonance image; a laser interstitial thermal therapy (LITT) device, comprising: an LITT probe, the LITT probe being positioned close to the target object based on the magnetic resonance image, and configured to treat the target object by emitting laser; and a temperature measurement element, the temperature measurement element and the LITT probe being integrated as an integrated probe, and the temperature measurement element being configured to obtain a temperature of the target object, wherein the LITT probe comprises an LITT lateral ablation probe or an LITT circumferential ablation probe. 14. The medical treatment apparatus of claim 13, wherein the LITT lateral ablation probe comprises: a probe main body; and a coating layer covered on an end surface of the probe main body, wherein the end surface of the probe main body and an axis of the probe main body form a first angle, the probe main body comprises a probe core and a hard cladding structure located at an outer peripheral of the probe core, and the coating layer comprises noble metal. 15. The medical treatment apparatus of claim 14, wherein the coating layer has a dual-layer structure, a layer of the dual-layer structure which is close to the end surface of the probe main body comprises pure silver, and the other layer of the dual-layer structure which is away from the end surface of the probe main body comprises silicon monoxide. 16. The medical treatment apparatus of claim 14, wherein the LITT lateral ablation probe is obtained by brightening and cleaning the end surface of the probe main body, and then further covering the coating layer on the end surface through magnetron sputtering. 17. The medical treatment apparatus of claim 13, wherein the LITT lateral ablation probe comprises: a probe main body; and a lens connected to an end surface of the probe main body, wherein the end surface of the probe main body is perpendicular to an axis of the probe main body, the lens has a first vertex angle, the probe main body comprises a probe core and a hard cladding structure located at an outer peripheral of the probe core, and the lens is a sapphire lens. 18. The medical treatment apparatus of claim 17, wherein the lens is a chamfered cylindrical lens, or a chamfered hemi-spherical or hemi-ellipsoidal lens. 19. The medical treatment apparatus of claim 17, wherein the LITT lateral ablation probe is obtained by brightening and cleaning the end surface of the probe main body, connecting the probe main body and the lens through fusion welding, and processing a welded surface between the probe main body and the lens using fire polishing. 20. The medical treatment apparatus of claim 13, wherein the LITT circumferential ablation probe comprises: a probe main body, the main body of the probe comprising a cone-shaped surface, one or more grooves being set on the cone-shaped surface, the one or more grooves being evenly distributed on the cone-shaped surface in an arabesquitic pattern, wherein the LITT circumferential ablation probe emits laser out of the one or more grooves. 21. The medical treatment apparatus of claim 20, wherein a diameter of the cone-shaped surface decreases from an initial diameter to a preset diameter, the initial diameter is 630 μm, and the preset diameter is 100 μm. 22. The medical treatment apparatus of claim 20, wherein the arabesquitic pattern comprises a single thread pattern, a cross-thread pattern, or a rhombic grid pattern, or a combination thereof. 23. The medical treatment apparatus of claim 20, wherein the cone-shaped surface is formed by performing operations comprising: fixing two ends of a first workpiece with cone-shaped surface to be carved on a sliding rail and a fixing device, respectively; controlling, by a laser device controller, a laser device, to emit laser, wherein a power of the emitted laser is 30 W, a wavelength of the emitted laser is 10600 nm, the emitted laser is attenuated to 6.6-11.2 W by a laser power attenuator, the attenuated laser is split into two laser beams by a diffraction splitting lens unit, and the two laser beams have equal power and are focused onto a surface of the first workpiece by a focusing lens unit; and driving, by a motion driver, the first workpiece to move, the motion driver being controlled by a motion driver controller. 24. The medical treatment apparatus of claim 20, wherein the one or more grooves that are evenly distributed on the cone-shaped surface in the arabesquitic pattern are formed by performing operations comprising: fixing two ends of a second workpiece with arabesquitic pattern to be carved on a sliding rail and a fixing device, respectively; controlling, by a laser device controller, a laser device, to emit laser, wherein a power of the emitted laser is 30 W, a wavelength of the emitted laser is 10600 nm, the emitted laser is focused on a surface of the second workpiece by a lens set, the lens set comprises a first concave lens, a first convex lens, and a second concave lens, the first concave lens and the first convex lens are used to expand a diameter of the emitted laser into a first diameter, the second concave lens is used to focus the laser beam of the first diameter to form a spot on the surface of the second workpiece, and the spot has a diameter of 45 μm; and driving, by a motion driver, the second workpiece to move, the motion driver being controlled by a motion driver controller. 25. The medical treatment apparatus of claim 20, wherein a spatial intensity distribution of the laser emitted by the LITT circumferential ablation probe is obtained through operations comprising: obtaining signals output by a laser measurement sensor, and generating a polar intensity of a specific orientation of the laser emitted by the LITT circumferential ablation probe, wherein the LITT circumferential ablation probe is connected to a helium-neon laser device that generates laser with a wavelength of 632.8 nm, a slit diaphragm is set between the laser measurement sensor and the LITT circumferential ablation probe, and a slit size of the slit diaphragm is 0.3 mm; and determining the spatial intensity distribution of the laser emitted by the LITT circumferential ablation probe by driving, by a motion driver, the slit diaphragm and the laser measurement sensor to move along a circumferential direction and an axial direction of the cone-shape surface of the LITT circumferential ablation probe, the motion driver being controlled by a motion driver controller, wherein the slit diaphragm and the laser measurement sensor are fixedly connected to the motion driver. | A laser interstitial thermal therapy (LITT) apparatus combining an MRI scanner, a LITT probe positioned by MR imaging, and temperature monitoring. MR imaging guides probe placement while temperature feedback controls the laser dose. | R | "a magnetic resonance imaging (MRI) device, configured for imaging of a specific region comprising a target object and generating a magnetic resonance image; a laser interstitial thermal therapy (LITT) device, comprising: an LITT probe, the LITT probe being positioned close to the target object based on the magnetic resonance image" (Claim 13) "a temperature measurement element, the temperature measurement element and the LITT probe being integrated as an integrated probe, and the temperature measurement element being configured to obtain a temperature of the target object" (Claim 13) |  |  |  |  | 1 |  |  |  | 1 |  |  |  |  | 1 | [14 empty] | 1 |  |  | SHANGHAI HROCT LITT MEDICAL TECH CO LTD | SHANGHAI HROCT LITT MEDICAL TECHNOLOGY CO., LTD. | 2022-11-03 | 2022 | 2024-01-04 | 2024 | 2022-07-02 | 2022 | US202217980444 | CN117338407A¦CN117530770A¦CN117770947A¦US20240000507A1¦WO2024060314A1
399 | CN117883175A | Radio frequency ablation system and computer readable medium | The invention relates to a radio frequency ablation system and a computer readable medium, the system comprises a radio frequency host, an electrode needle device and an electrode slice module, the radio frequency host comprises a main control module, a liquid cooling module, a temperature detection module, a man-machine interaction module, an electrode needle device connecting end and an electrode slice module connecting end; the temperature detection module is used for collecting the ablation temperature of the working end of the needle device and sending the ablation temperature to the main control module; the liquid cooling module adjusts the flow of liquid conveyed to the electrode needle device according to the ablation temperature and the control signal; the man-machine interaction module is used for sending a control signal to the main control module, the man-machine interaction module is further used for displaying system information, and the main control module is further connected with the electrode needle device connecting end and the electrode slice module connecting end. The computer readable medium is used for realizing the control method when the system carries out ablation. The radiofrequency ablation system and the computer readable medium have the advantages of being accurate, convenient, easy to operate and high in safety. | 1. A radiofrequency ablation system, characterized in that it includes a radiofrequency host, an electrode needle device and an electrode sheet module; the radiofrequency host includes a main control module, a liquid cooling module, a temperature detection module, a human-computer interaction module, and an electrode needle device connection terminal and the connection end of the electrode sheet module; the temperature detection module is connected to the main control module and the connection end of the electrode needle device respectively, and the temperature detection module is used to collect the temperature connected to the connection end of the electrode needle device. The ablation temperature of the working end of the electrode needle device is sent to the main control module; the liquid cooling module is connected to the electrode needle device and the main control module respectively. The liquid cooling module is connected to the control module, and the liquid cooling module adjusts the liquid flow delivered to the corresponding electrode needle device according to the ablation temperature and control signal; the human-computer interaction module is connected to the main control module, and the human-computer interaction module is connected to the main control module. The computer interaction module is used to receive the control signal and send the control signal to the main control module, and the human-computer interaction module is also used to display system information, where the system information includes the ablation temperature ; The main control module is also connected to the connecting end of the electrode needle device and the connecting end of the electrode sheet module respectively. 2. The radiofrequency ablation system according to claim 1, wherein the radiofrequency host further includes an electrical signal acquisition module, and the electrical signal acquisition module is connected to the main control module, the electrode needle device connection end and the electrode sheet module respectively. The electrical signal acquisition module is used to collect electrical signals from the electrode needle device connection end and the electrode sheet module connection end, and generate corresponding electrical signal information according to the electrical signal, and the The electrical signal acquisition module sends the electrical signal to the main control module. 3. The radiofrequency ablation system according to claim 2, characterized in that the main control module determines the electrode needle device connected to the connection end of the electrode needle device and the electrode sheet module based on the electrical signal. The connection state of the electrode sheet module at the connection end, and the connection state of the electrode needle device connected to the connection end of the electrode needle device and/or the electrode sheet module connected to the connection end of the electrode sheet module When an abnormality occurs, a corresponding first warning message is generated and sent to the human-computer interaction module. 4. The radiofrequency , wherein the main control module determines whether the tissue is coagulated and/or carbonized based on the electrical signal, and generates a signal when the tissue is coagulated and/or carbonized. The corresponding second warning prompt information is sent to the human-computer interaction module. 5. The radiofrequency ablation system according to claim 2, wherein the radiofrequency host further includes a radiofrequency generation module, and the radiofrequency generation module is respectively connected to the main control module, the electrode needle device connection end and the electrode sheet module. The electrical signal acquisition module is used to collect the real-time power value output by the radio frequency generating module, and send the real-time power value to the main control module, and the main control module determines the real-time power value according to the real-time power value. The power deviation value between the preset power value and the power value determines whether the radio frequency generating module is working normally, and when the radio frequency generating module works abnormally, a corresponding third warning message is generated and sent to the human-computer interaction module. 6. The radiofrequency ablation system according to claim 5, characterized in that the radiofrequency host further includes a power module, and the power module is respectively connected with the electrode needle device, electrode sheet module, main control module, liquid cooling module, The temperature detection module, the human-computer interaction module, the electrical signal acquisition module and the radio frequency generation module are connected; after the radio frequency ablation system is powered on, the electrical signal acquisition module collects the connection ends of the electrode needle device and the electrode sheet module respectively. The power-on status of the connection end, the main control module, the liquid cooling module, the temperature detection module, the human-computer interaction module and the radio frequency generation module is generated, and the corresponding power-on detection signal is generated and transmitted to the main control module. When the electrode needle device connection When any abnormality occurs in any of the terminals, electrode module connection terminals, main control module, liquid cooling module, temperature detection module, human-computer interaction module, electrical signal acquisition module and radio frequency generation module, the corresponding third message will be generated. Four early warning prompts are sent to the human-computer interaction module. 7. The radio frequency ablation system according to claim 6, wherein after the radio frequency ablation system is powered on, the main control module receives heartbeat data regularly sent by each monitored node in the radio frequency ablation system. When the heartbeat data of the monitored node is refreshed abnormally, it is determined that the monitored node is abnormal, and a corresponding fifth warning message is generated and sent to the human-computer interaction module; wherein the monitored node includes a computer connected to The electrode needle device at the connection end of the electrode needle device, the electrode sheet module connected to the electrode sheet module connection end, the liquid cooling module, the temperature detection module, the human-computer interaction module, the At least one component of the electrical signal acquisition module and the radio frequency generation module. 8. The radiofrequency ablation system according to claim 2, wherein the liquid cooling module includes a water pump, and the electrical signal acquisition module obtains the real-time rotational speed pulse value of the water pump in real time, and sends the real-time rotational speed pulse value. To the main control module, the main control module determines whether the water pump is working normally based on the pulse difference between the real-time speed pulse value and the preset speed pulse value corresponding to the preset flow rate of the water pump, and When the water pump works abnormally, a corresponding sixth warning message is generated and sent to the human-computer interaction module. 9. The radiofrequency ablation system according to claim 1, wherein the human-computer interaction module includes a touch screen, a control switch unit and a speaker; the touch screen and the speaker are respectively connected to the main control module; The control switch unit is respectively connected to the main control module and the electrode needle device. 10. The radiofrequency ablation system according to claim 1, characterized in that: the radiofrequency host further includes an electrode needle multi-channel detection control module; the electrode needle device connection end includes several electrode needle interfaces; the electrode needle device One or more electrode needle mechanisms can be selectively connected to one of the electrode needle interfaces; the temperature detection module and the main control module are connected to each of the electrodes through the electrode needle multi-channel detection control module The electrode needle multi-channel detection control module is used to send the connection status of the electrode needle mechanism detected at each of the electrode needle interfaces to the main control module, and connect it to the corresponding electrode needle The ablation temperature signal of the working end of the electrode needle mechanism at the interface is transmitted to the temperature detection module, so that the temperature detection module obtains the corresponding ablation temperature according to each of the ablation temperature signals; the main control module can according to The connection status of the electrode needle mechanism detected from each of the electrode needle interfaces, the ablation temperature of the working end of the electrode needle mechanism connected to the corresponding electrode needle interface, and the control signal are adjusted and sent to the corresponding The liquid flow rate of the electrode needle mechanism is controlled, and the working status of each electrode needle mechanism connected to the corresponding electrode needle interface is controlled. 11. The radiofrequency ablation system according to claim 10, wherein when two or more electrode needle mechanisms are connected to corresponding electrode needle interfaces, each of the pairs is connected to the corresponding electrode needle interface. Controlling the working status of the electrode needle mechanism at the electrode needle interface includes: controlling the electrode needle mechanism connected to the electrode needle interface to work in turn according to the system's preset power, and in the radiofrequency ablation system When any component is abnormal, the electrode needle mechanism is controlled to stop working. 12. A computer-readable medium, characterized in that a computer program is stored on the computer-readable medium, and when the computer program is executed by a processor, the radiofrequency ablation system according to any one of claims 1 to 11 is implemented. Control method, the control method includes the following steps: Real-time monitoring of the ablation temperature of the working end of the electrode needle device connected to the connection end of the electrode needle device, and sending the ablation temperature to the main control module and the A human-computer interaction module, the human-computer interaction module is used to display the ablation temperature; the liquid cooling module adjusts the liquid flow rate delivered to the corresponding electrode needle device according to the ablation temperature and the control signal. 13. The computer-readable medium according to claim 12, wherein the control method further includes the following steps: collecting electrical signals at the connecting end of the electrode needle device and the connecting end of the electrode sheet module in real time, and Corresponding electrical signal information is generated according to the electrical signal, and the electrical signal information includes loop status information of the connecting end of the electrode needle device and the connecting end of the electrode sheet module; determining the electrode needle device based on the loop status information and the connection status of the electrode sheet module and the radio frequency host, wherein when the loop status information shows that the electrode needle device connection end and/or the electrode sheet module connection end does not form a loop, then It is determined that the connecting end of the electrode needle device and/or the connecting end of the electrode sheet module is not connected to the corresponding electrode needle device and/or the electrode sheet module; when the circuit status information shows that the connecting end of the electrode needle device and/or when the connecting end of the electrode sheet module forms a loop, it is determined that the connecting end of the electrode needle device and/or the connecting end of the electrode sheet module is connected to the corresponding electrode needle device and/or electrode sheet module . 14. The computer-readable medium according to claim 13, wherein the control method further includes the following steps: when the radiofrequency ablation system performs an ablation operation, the circuit status information displays the connection end of the electrode needle device and/or the connecting end of the electrode sheet module does not form a loop, then determine the electrode needle device connected to the connecting end of the electrode needle device and/or the electrode sheet module connected to the connecting end of the electrode sheet module If the connection status of the group is abnormal, the corresponding first warning message will be generated to issue an alarm. 15. The computer-readable medium according to claim 13, wherein the electrical signal information further includes an impedance value formed by the connection end of the electrode needle device and the connection end of the electrode sheet module, and the control method further It includes the following steps: Determining the ablation status of the tissue according to the impedance value. When detecting that the impedance value exceeds a system preset threshold, it is determined that the tissue is coagulated and/or carbonized, and a corresponding second warning prompt is generated. information and make an alarm. 16. The computer-readable medium according to claim 13, wherein the radio frequency host further includes a radio frequency generation module, and the control method further includes the following steps: collecting the real-time power value output by the radio frequency generation module, and based on the The power deviation value between the real-time power value and the preset power value is used to determine whether the radio frequency generating module is working normally, and when the radio frequency generating module is working abnormally, a corresponding third warning message is generated to issue an alarm. 17. The computer-readable medium according to claim 16, characterized in that the control method further includes the following steps: after the radiofrequency ablation system is powered on, separately collecting the connecting end of the electrode needle device and the electrode sheet module The power-on status of the connection end, main control module, liquid cooling module, temperature detection module, human-computer interaction module and radio frequency generation module, and generates a pair of According to the power-on detection signal, when the electrode needle device connection end, the electrode sheet module connection end, the main control module, the liquid cooling module, the temperature detection module, the human-computer interaction module, the electrical signal acquisition module and the radio frequency generation module When an abnormality occurs in any component of If any component among the module, the human-computer interaction module and the radio frequency generation module cannot feed back the corresponding power-on detection signal, it is determined that the component is abnormal. 18. The computer-readable medium according to claim 16, wherein the control method further includes the following steps: regularly receiving heartbeat data regularly sent by each monitored node in the radio frequency ablation system, when any of the monitored nodes When there is a refresh abnormality in the heartbeat data of the monitoring node, it is determined that the monitored node is abnormal, and a corresponding fifth warning message is generated to issue an alarm; wherein the monitored node includes all the nodes connected to the connection end of the electrode needle device. The electrode needle device, the electrode sheet module connected to the electrode sheet module connection end, the liquid cooling module, the temperature detection module, the human-computer interaction module, the electrical signal acquisition module and the At least one component in the radio frequency generation module. 19. The computer-readable medium according to claim 13, wherein the liquid cooling module includes a water pump, and the control method further includes the following steps: obtaining the real-time speed pulse value of the water pump in real time, and based on the real-time speed The pulse difference between the pulse value and the preset speed pulse value corresponding to the preset flow rate of the water pump is used to determine whether the water pump is working normally, and when the water pump is working abnormally, the corresponding sixth warning message is generated. Call the police. 20. The computer-readable medium according to claim 13, wherein the electrode needle device includes two or more electrode needle mechanisms, and the control method further includes the following steps: when two or more electrode needle mechanisms When the electrode needle mechanism is connected to the connection end of the electrode needle device, the ablation temperature and the control signal collected from the working end of each electrode needle mechanism are adjusted and transmitted to the corresponding electrode needle mechanism. liquid flow, and control the working status of each electrode needle mechanism so that each electrode needle mechanism works at a preset operating temperature and continues to work for a preset working time. 21. twenty one. The computer-readable medium according to claim 13, wherein the control method further includes the following steps: when the radiofrequency ablation system is in manual working mode, the radiofrequency host receives the information according to the human-computer interaction module. The real-time control signal received controls the switch of the electrode needle device and the adjustment of the working power; when the radiofrequency ablation system is in the automatic working mode, the radiofrequency host controls the switch according to the preset control signal received by the human-computer interaction module. , control the preset working time of the electrode needle device, monitor the ablation temperature of the electrode needle device in real time, and adjust the liquid flow rate delivered by the liquid cooling module to the corresponding electrode needle device according to the detected ablation temperature. , so that the electrode needle device works within the preset ablation temperature range. | An RF ablation system with a host integrating main control, liquid cooling, temperature detection, HMI, and connectors for the electrode needle and electrode sheet. The host coordinates ablation, cooling, and live temperature reads in one device. | R | determining the ablation status of the tissue according to the impedance value. When detecting that the impedance value exceeds a system preset threshold, it is determined that the tissue is coagulated and/or carbonized, and a corresponding second warning prompt is generated | 1 |  |  |  | 1 | [5 empty] | 1 | [9 empty] | 1 | [9 empty] | 1 | SHANGHAI MEIWEIDA MEDICAL TECH CO LTD | SHANGHAI MEIWEIDA MEDICAL TECH CO LTD | 2022-10-08 | 2022 | 2024-04-16 | 2024 | 2022-10-08 | 2022 | CN202211222360 | [NA]
400 | EP4378393A1 | METHOD AND SYSTEM FOR DETERMINING A TEMPERATURE AND/OR TEMPERATURE DISTRIBUTION IN AN ABLATION REGION | The invention relates to a method and system for determining a temperature and/or temperature distribution (TA) in an ablation area (5), wherein the temperature and/or temperature distribution (TA) based on first and second image data sets (6) and temperatures provided for a first sensor position (S) and a second sensor position (S'). | 1. Method for determining a temperature and/or a temperature distribution (TA) in an ablation area (5), comprising the steps: - providing (100) a first image data set (6) of an object region (8), wherein the first image data set (6) comprises medical image data (7) recorded with a different spectral distribution of an X-ray radiation, wherein the object region (8) comprises the ablation region (5), - determining (200) a first sensor position (S) in the object area (8) based on the first image data set (6), - determining (300) a first density, a first density distribution, a first attenuation value, a first attenuation value distribution and/or a first tissue distribution in a surrounding area of ​​the first sensor position (S) based on the first image data set (6), - providing (400) a first temperature for the first sensor position (S), - providing (500) a second image data set (6) of the object area (8), wherein the second image data set (6) comprises medical image data (7) recorded with a different spectral distribution of an X-ray radiation, - determining (600) a second sensor position (S') in the object area (8) based on the second image data set (6), - determining (700) a second density, a second density distribution, a second attenuation value, a second attenuation value distribution and/or a second tissue distribution in a surrounding area of ​​the second sensor position (S') based on the second image data set (6), - providing (700) a second temperature for the second sensor position (S'), - Determining (1000) a temperature and/or a temperature distribution (TA)in the ablation region (5) based on the first temperature, the first sensor position (S), the second temperature, the second sensor position (S'), the first density, the first density distribution, the first attenuation value, the first attenuation value distribution, the first tissue distribution, the second density, the second density distribution, the second attenuation value, the second attenuation value distribution and/or the second tissue distribution, - Providing (1100) the determined temperature and/or temperature distribution (TA). 2. Method according to claim 1,characterized in that in the step of determining (700) a second density distribution, a second attenuation value distribution and/or a second tissue distribution, the second density distribution, the second attenuation value distribution and/or the second tissue distribution in the object region (8) is determined. 3. Method according to claim 1 or 2,characterized in that in the step of determining (300) a first density, a first density distribution, a first attenuation value, a first attenuation value distribution and/or a first tissue distribution, the first density, the first density distribution, the first attenuation value, the first attenuation value distribution and/or the first tissue distribution in the object region (8) is determined. 4. Method according to one of the preceding claims,marked by: - determining a density change, an attenuation value change and/or a tissue distribution change based on the first density, the first density distribution, the first attenuation value, the first attenuation value distribution, the first tissue distribution, the second density, the second density distribution, the second attenuation value, the second attenuation value distribution and/or the second tissue distribution, wherein the temperature and/or temperature distribution (TA) in the ablation region (5) is determined based on the first temperature, the second temperature, the density change, the attenuation value change and/or the tissue distribution change. 5. Method according to one of the preceding claims,marked by: - determining a calibration function based on the first temperature, the first sensor position (S), the second temperature, the second sensor position (S'), the first density, the first density distribution, the first attenuation value, the first attenuation value distribution, the first tissue distribution, the second density, the second density distribution, the second attenuation value, the second attenuation value distribution and/or the second tissue distribution,where the calibration function assigns a temperature to a density, an attenuation value and/or a tissue. 6. Method according to one of the preceding claims,characterized in that the provision (100, 500) of a first and/or a second image data set (7) of an object region (8) comprises an artifact removal, wherein the artifact removal removes and/or reduces image artifacts and/or metal artifacts in the medical image data (7). 7. Method according to one of the preceding claims,characterized in that the first temperature and/or the second temperature is provided by a sensor device, wherein the sensor device comprises a first sensor element (11) and a second sensor element (11), wherein the first sensor element (11) and the second sensor element (11) are spaced apart at a sensor distance, wherein the first sensor position (S) is a first sensor element position (SE1) for the first sensor element (11) and a second sensor element position (SE2) for the second sensor element (11), wherein the first temperature and/or the second temperature comprises a first temperature value for the first sensor element (11) and a second temperature value for the second sensor element (11). 8. Method according to one of the preceding claims,characterized in that the first temperature and/or the second temperature is provided by a sensor device, wherein the sensor device is comprised by an ablation device (2), wherein the ablation device comprises an ablation section (10) for thermal ablation. 9. Method according to one of the preceding claims,marked by: - providing a third image data set of the object area, wherein the third image data set (6) comprises medical image data (7) recorded with a different spectral distribution of an X-ray radiation, - determining a third temperature and/or a third temperature distribution in the ablation region (5) based on the first temperature, the first sensor position (S), the second temperature, the second sensor position (S'), the first density, the first density distribution, the first attenuation value, the first attenuation value distribution, the first tissue distribution, the second density, the second density distribution, the second attenuation value,the second attenuation value distribution and/or the second tissue distribution, - Determining a planned temperature profile in the ablation area (5) based on the first, second and third temperature and/or temperature distribution (TA), - Providing the planned temperature profile. 10. Method according to claim 9,characterized in that Ablation operating data and/or tissue parameters are provided, whereby the planned temperature profile is determined based on the ablation operating data and/or tissue parameters. 11. Method according to one of claims 9 or 10,characterized in that Based on the planned temperature profile, a planned residual ablation time is determined, whereby the planned residual ablation time is provided. 12. Method according to one of the preceding claims,characterized in that based on the first temperature, the second temperature, the third temperature, the first temperature distribution, the second temperature distribution and/or the third temperature distribution, heat loss areas in the ablation area (5) are determined, wherein the determined heat loss areas are provided. 13. System (1) for determining a temperature and/or a temperature distribution (TA) in an ablation area (5), comprising an image data provision module (20), an image data evaluation module (21), a sensor data provision module (22), a position determination module (23), a temperature determination module (24) and a provision module (25), wherein: - The image data provision module (20) is designed to provide a first image data set (6) of an object region (8), wherein the first image data set (6) comprises medical image data (7) recorded with a different spectral distribution of an X-ray radiation, wherein the object region (8) comprises the ablation region (5), - The image data evaluation module (21) is designed to determine a first density, a first density distribution, a first attenuation value, a first attenuation value distribution and/or a first tissue distribution in at least one section of the object region based on the first image data set (6), - The position determination module (23) is designed to determine a first sensor position (S) in the object area (8) based on the first image data set (6), - The sensor data provision module (22) is designed to provide a first temperature for the first sensor position (S), - The image data provision module (20) is designed to provide a second image data set (6) of the object area (8), wherein the second image data set (6) comprises medical image data (7) recorded with a different spectral distribution of an X-ray radiation, - The position determination module (23) is designed to determine a second sensor position (S') in the object area (8) based on the second image data set (6), - The sensor data provision module (22) is designed to provide a second temperature for the second sensor position (S'), - The image data evaluation module (21) is designed to determine a second density, a second density distribution, a second attenuation value, a second attenuation value distribution and/or a second tissue distribution in a surrounding area of ​​the second sensor position (S') based on the second image data set (6), - The temperature determination module (24) is designed to determine a temperature and/or a temperature distribution (TA) in the ablation area (5) based on the first temperature, the first sensor position (S), the second temperature, the second sensor position (S'), the first density, the first density distribution, the first attenuation value, the first attenuation value distribution, the first tissue distribution, the second density, the second density distribution, the second attenuation value, the second attenuation value distribution and/or the second tissue distribution, - The provision module (25) is designed to determine the specific temperature and/or temperature distribution (TA). 14. Ablation arrangement (1) comprising an ablation device (2) and a system according to claim 13, wherein the ablation device (2) has a sensor device for providing the first, second and/or third temperature and/or temperature distribution. 15. Computer program product comprising program code means for causing a system according to claim 13, an ablation arrangement (1) according to claim 14 and/or a computer to carry out the method for determining a temperature and/or temperature distribution (TA) in an ablation area (5). | A method that determines temperature or temperature distribution in an ablation region using two image datasets and temperature readings from two sensor positions. Combining imaging with two-point sensor data produces a richer temperature map. | R | "Determining a temperature and/or a temperature distribution (TA) in the ablation region (5) based on the first temperature, the first sensor position (S), the second temperature, the second sensor position (S'), the first density, the first density distribution, the first attenuation value… the second tissue distribution" (Claim 1) "determining a calibration function based on the first temperature, the first sensor position (S), the second temperature… where the calibration function assigns a temperature to a density, an attenuation value and/or a tissue" (Claim 5) "Determining a planned temperature profile in the ablation area (5) based on the first, second and third temperature and/or temperature distribution (TA)" (Claim 9) "based on the first temperature, the second temperature, the third temperature… heat loss areas in the ablation area (5) are determined" (Claim 12) |  |  |  |  | 1 |  |  |  | 1 | [13 empty] | 1 | [8 empty] | SIEMENS HEALTHINEERS AG | SIEMENS HEALTHINEERS AG | 2022-11-30 | 2022 | 2024-06-05 | 2024 | 2022-11-30 | 2022 | EP20220210596 | [NA]
401 | CN120241030A | Laser ablation method based on magnetic resonance image temperature calculation and magnetic resonance guidance | The embodiment of the invention provides a magnetic resonance image-based temperature calculation and magnetic resonance guided laser ablation method, and the method comprises the steps: periodically scanning a target region containing a focus through a magnetic resonance fast imaging sequence, and obtaining an initial magnetic resonance image sequence of a target period, the target period is any period of periodically scanning the target area; the initial magnetic resonance image sequence of the target period is corrected according to a reference period, a corrected magnetic resonance image sequence of the target period is obtained, and the reference period is a period in which a reference area containing a focus is scanned by magnetic resonance in advance to obtain a space standard image; the corrected magnetic resonance image sequence of the target period and the corrected magnetic resonance image sequence of the historical period are used to generate the temperature map of the target period, the historical period is any period before the target period, distortion generated in the rapid scanning process is eliminated, and the accuracy and convenience of temperature monitoring are improved. | 1. A method for calculating temperature based on magnetic resonance images, characterized in that the method comprises: periodically scanning a target area containing a lesion through a magnetic resonance rapid imaging sequence to obtain an initial magnetic resonance image sequence of the target cycle, wherein the target cycle is any cycle of periodic scanning of the target area; correcting the initial magnetic resonance image sequence of the target cycle according to a reference cycle to obtain a corrected magnetic resonance image sequence of the target cycle, wherein the reference cycle is a cycle of obtaining a spatial standard image in advance by using magnetic resonance to scan a reference area containing the lesion; using the corrected magnetic resonance image sequence of the target cycle and the corrected magnetic resonance image sequence of a historical cycle to generate a temperature map of the target cycle, wherein the historical cycle is any cycle before the target cycle. 2.The method according to claim 1 is characterized in that the initial magnetic resonance image sequence of the target period is corrected according to the reference period to obtain the corrected magnetic resonance image sequence of the target period, including: acquiring a spatial standard image obtained by pre-scanning; correcting the initial magnetic resonance image sequence of the target period according to the spatial standard image to obtain the corrected magnetic resonance image sequence of the target period. 3.The method according to claim 1 is characterized in that the correcting the initial magnetic resonance image sequence of the target cycle according to the reference cycle to obtain the corrected magnetic resonance image sequence of the target cycle includes: obtaining the distortion correction information of the previous cycle according to the reference cycle and the previous cycle of the target cycle; correcting the initial magnetic resonance image sequence of the target cycle using the distortion correction information to obtain the corrected magnetic resonance image sequence of the target cycle; determining and storing the distortion correction information of the target cycle based on the initial magnetic resonance image sequence of the target cycle and the spatial standard image, wherein the distortion correction information of the target cycle is used to correct the initial magnetic resonance image sequence of the next cycle of the target cycle. 4.The method according to claim 3 is characterized in that the distortion correction information of the target period is determined and stored based on the initial magnetic resonance image sequence of the target period and the spatial standard image, including: determining the first tissue position information of the target tissue in the initial magnetic resonance image sequence of the target period; determining the second tissue position information of the target tissue in the spatial standard image; calculating the spatial transformation matrix based on the first tissue position information and the second tissue position information, using the spatial transformation matrix as the distortion correction information, and storing it. 5.The method according to claim 1 is characterized in that the correcting the initial magnetic resonance image sequence of the target period according to the reference period to obtain the corrected magnetic resonance image sequence of the target period includes: selecting any period as a correction reference period; obtaining constant distortion correction information according to the reference period and the correction reference period; correcting the initial magnetic resonance image sequence of the target period using the constant distortion correction information to obtain the corrected magnetic resonance image sequence of the target period. 6.The method according to claim 5 is characterized in that the constant distortion correction information is obtained according to the reference period and the correction base period, comprising: when the target period is earlier than the correction base period, not correcting the initial magnetic resonance image sequence of the target period, continuing to scan the next period until the target period is the correction base period, and scanning to obtain the initial magnetic resonance image sequence of the correction base period; acquiring the spatial standard image of the reference period, and obtaining the constant distortion correction information according to the spatial standard image and the correction base period. An initial magnetic resonance image sequence is used to obtain constant distortion correction information of the correction reference period. 7.The method according to claim 1 is characterized in that the use of the corrected magnetic resonance image sequence of the target cycle and the corrected magnetic resonance image sequence of the historical cycle to generate the temperature map of the target cycle includes: establishing a temperature difference map using the phase difference values of the corrected magnetic resonance image sequence of the target cycle and the corrected magnetic resonance image sequence of the historical cycle; generating the temperature map of the target cycle based on the temperature difference map and a baseline temperature map, wherein the baseline temperature map is the temperature map of the historical cycle. 8.The method according to any one of claims 1 to 7 is characterized in that the spatial standard image is obtained by scanning the reference area based on a spin echo sequence; the initial magnetic resonance image sequence of the target period is obtained by scanning the target area based on gradient echo imaging. 9.The method according to any one of claims 1 to 7 is characterized in that the periodic scanning of the target area containing the lesion by means of a rapid magnetic resonance imaging sequence to obtain an initial magnetic resonance image sequence of the target period includes: determining the target layer and/or the target sampling column within the layer to be sampled in the target period according to a set downsampling rule; scanning the target layer and/or the target sampling column within the layer by means of the rapid magnetic resonance imaging sequence to obtain a downsampling sequence of the target period; and completing and reconstructing the downsampling sequence by means of a reconstruction algorithm to obtain an initial magnetic resonance image sequence of the target period. 10. A method for calculating temperature based on magnetic resonance images, characterized in that the method includes: by setting a scanning mode, adopting a downsampling strategy to scan a target area containing a lesion, and obtaining a downsampling sequence of a target cycle, wherein the target cycle is any cycle of periodic scanning of the target area, the downsampling strategy is determined based on the set scanning mode, and the downsampling sequence of the target cycle contains image information of at least one optical fiber; reconstructing a complete magnetic resonance image sequence according to the downsampling sequence of the target cycle to obtain a reconstructed magnetic resonance image sequence of the target cycle; acquiring a reconstructed magnetic resonance image sequence of a historical cycle, and generating a temperature map of the target cycle based on the reconstructed magnetic resonance image sequence of the target cycle and the reconstructed magnetic resonance image sequence of the historical cycle, wherein the historical cycle is any cycle before the target cycle, and the temperature map is used to provide information support for the laser ablation process of the at least one optical fiber. 11.The method according to claim 10 is characterized in that the setting of the scanning mode is layer scanning, and the downsampling strategy is inter-layer downsampling; the setting of the scanning mode and the adoption of the downsampling strategy to scan the target area containing the lesion to obtain the downsampling sequence of the target period include: determining the target layer to be sampled in the target period according to the set inter-layer downsampling rule; scanning the target layer by the set scanning mode to obtain the downsampling sequence of the target period. 12.The method according to claim 11 is characterized in that the set inter-layer downsampling rule is to sample odd layers and even layers respectively at intervals; determining the target layer to be sampled in the target cycle according to the set inter-layer downsampling rule comprises: determining the sampling layer scanned in the previous cycle; when the sampling layer is an odd layer, determining the target layer to be sampled in the target cycle to be an even layer; when the sampling layer is an even layer, determining the target layer to be sampled in the target cycle to be an odd layer. 13. The method according to claim 11 is characterized in that the set inter-layer downsampling rule is the layer to be sampled When the number of layers to be sampled exceeds the layer number threshold, the first strategy is adopted, and when the number of layers to be sampled does not exceed the layer number threshold, the second strategy is adopted; the target layer to be sampled in the target cycle is determined according to the set inter-layer downsampling rule, including: determining the number of layers to be sampled in the target area; when the number of layers to be sampled exceeds the layer number threshold, determining the target layer to be sampled in the target cycle according to the first strategy; When the number of layers to be sampled does not exceed the layer number threshold, the target layers to be sampled in the target cycle are determined according to the second strategy. 14.The method according to claim 10 is characterized in that the set scanning mode is column scanning, and the downsampling strategy is inter-layer downsampling and/or intra-layer downsampling; the downsampling strategy is used to scan the target area containing the lesion by setting the scanning mode to obtain a downsampling sequence of the target period, including: determining the target layer and/or the target sampling column within the layer to be sampled in the target period according to the set inter-layer downsampling rules and/or intra-layer downsampling rules; scanning the target layer and/or the target sampling column within the layer by the set scanning mode to obtain the downsampling sequence of the target period. 15.A magnetic resonance-guided laser ablation method, characterized in that the method comprises: obtaining a temperature map of a target period, wherein the temperature map of the target period is obtained by the method for calculating temperature based on a magnetic resonance image as described in any one of claims 1-9 or 1014; generating an ablation map according to the temperature map of the target period, wherein the ablation map is used to provide information support for the laser ablation process of at least one optical fiber. 16.A magnetic resonance-guided laser ablation system, characterized in that the system includes a host, and the host is provided with an application for implementing the method for calculating temperature based on magnetic resonance images as described in any one of claims 1-9 or 10-14, or the magnetic resonance-guided laser ablation method as described in claim 15. 17. A computer-readable storage medium, characterized in that the computer-readable storage medium is arranged in a host of a magnetic resonance-guided laser ablation system, and stores computer-executable instructions, which, when executed by a processor, implement the method for calculating temperature based on magnetic resonance images as described in any one of claims 1-9 or 10-14, or the magnetic resonance-guided laser ablation method as described in claim 15. | An MR-guided laser ablation method that uses a fast MR imaging sequence to scan the lesion repeatedly and computes temperature from successive MR images. MR-derived thermometry then guides laser ablation in real time. | R | "periodically scanning a target area containing a lesion through a magnetic resonance rapid imaging sequence to obtain an initial magnetic resonance image sequence of the target cycle" (Claim 1) "using the corrected magnetic resonance image sequence of the target cycle and the corrected magnetic resonance image sequence of a historical cycle to generate a temperature map of the target cycle" (Claim 1) "establishing a temperature difference map using the phase difference values of the corrected magnetic resonance image sequence of the target cycle and the corrected magnetic resonance image sequence of the historical cycle" (Claim 7) "generating an ablation map according to the temperature map of the target period, wherein the ablation map is used to provide information support for the laser ablation process of at least one optical fiber" (Claim 15) |  |  |  |  | 1 |  |  |  | 1 |  |  |  |  | 1 | [5 empty] | 1 | [11 empty] | SINOVATION BEIJING MEDICAL EQUIPMENT CO LTD | SINOVATION BEIJING MEDICAL EQUIPMENT CO LTD | 2023-12-27 | 2023 | 2025-07-04 | 2025 | 2023-12-27 | 2023 | CN202311828361 | [NA]
402 | US20260083490A1 | METHOD AND DEVICE FOR ABLATING ANATOMICAL TARGETS FROM WITHIN A BLOOD VESSEL WITH IN-PROCEDURE FEEDBACK | A cryo-ablation system treats conditions associated with sympathetic nervous system hyperactivity using a catheter with a configurable heat transfer element that transitions from a flat delivery profile to deployed configurations including helical shapes. The system includes deployable temperature probes that directly measure tissue temperature at target locations inside or outside blood vessels. A computational control system processes temperature data from multiple sensors and deployable probes to calculate real-time lethal isotherm progression using heat transfer modeling algorithms. The system provides vessel occlusion during treatment and real-time feedback for treatment optimization. Methods include introducing the catheter into a renal artery, deploying the heat transfer element, occluding the vessel, deploying temperature probes to target locations within or outside the vessel, delivering cryogenic temperatures while monitoring probe temperatures, and terminating treatment based on computational analysis of temperature data to achieve target ablation depths. | 1. A cryo-ablation system for treating resistant hypertension, comprising: a cryo-ablation balloon configured to be placed within a blood vessel; a deployable temperature probe outside the cryo-ablation catheter, to perform direct temperature measurement of tissue, at a known distance from the balloon. 2. The cryo-ablation catheter of claim 1, including an occlusive cryo-ablation balloon configured to be placed within a blood vessel; the occlusive cryo-ablation balloon being selectively transitionable from a substantially flat configuration to a cylindrical, conical, or helical shape for creating a continuous, circumferential, or helical elongated cryogenic treatment region of at least 5 mm in length; a temperature monitoring system configured to measure temperatures at the temperature probe; and a pressure sensing system to detect arterial occlusion distal to the cryo-ablation balloon. 3. The cryo-ablation system of claim 1, wherein the catheter is an over-the-wire, rapid exchange, deflectable, or shapable configuration. 4. The cryo-ablation system of claim 1, further comprising a shapeable structure including a metallic core wire encapsulated in an atraumatic polymer body, allowing manual shaping of the catheter at a desired angle or curvature. 5. The cryo-ablation system of claim 1, further including a compliant occlusion balloon located on the catheter shaft proximal to the cryo-ablation balloon, wherein the occlusion balloon is configured to be manually inflated by the operator using a saline or contrast solution. 6. The cryo-ablation system of claim 1, wherein the cryo-ablation balloon is configured to transition to a helical shape. 7. The cryo-ablation system of claim 6, further comprising means to control the pitch, outer diameter, and length of the cryo-ablation balloon when deployed in the helical shape. 8. The cryo-ablation system of claim 1, further comprising an ultrasound imaging system to visualize progression of tissue freezing, measure a diameter of frozen tissue volume, and calculate a position of a lethal isotherm to provide real-time feedback during a cryo-ablation procedure. 9. The cryo-ablation system of claim 2, wherein the temperature monitoring system includes one or more temperature probes, at least one of which is situated outside the cryo-ablation balloon at a known distance to compute the distance of a lethal isotherm from the balloon surface. 10. The cryo-ablation system of claim 9, wherein the temperature monitoring system is configured to display the distance of the lethal isotherm as a guide during the cryo-ablation procedure. 11. The cryo-ablation system of claim 2, wherein the temperature monitoring system includes one or more deployable temperature probes that are advanceable through a vessel wall 1-10 mm. 12. The cryo-ablation system of claim 1, wherein the cryo-ablation balloon has a deployed diameter sufficient to completely occlude the blood vessel. 13. A method for treating resistant hypertension, comprising: introducing a cryo-ablation catheter with an occlusive cryo-ablation balloon into a renal artery; transitioning the cryo-ablation balloon from a flat configuration to a cylindrical, conical, or helical configuration within the renal artery to create a continuous cryogenic treatment region; occluding the renal artery using the cryo-ablation balloon to reduce blood flow and increase the efficacy of the cryo-ablation; achieving cryogenic temperatures to ablate renal sympathetic nerves adjacent to the renal artery; measuring the temperature at a known distances outside the cryo-ablation balloon to determine tissue freezing; and monitoring progression of tissue freezing at the known distances and adjusting cryo-ablation parameters based on real-time feedback from the temperature measurements. 14. The method of claim 13, further comprising using ultrasound to monitor progression of tissue freezing and using a pressure sensing system to confirm arterial occlusion distal to the cryo-ablation balloon. 15. The method of claim 14, further comprising manually shaping the cryo-ablation catheter to navigate and engage the distal branches of the renal artery. 16. The method of claim 14, further comprising the step of inflating a compliant occlusion balloon proximal to the cryo-ablation balloon to achieve vessel occlusion if the cryo-ablation balloon does not fully occlude the renal artery. 17. The method of claim 14, further comprising the step of computing a distance of a lethal isotherm from the cryo-ablation balloon and displaying this information to an operator for guiding duration and extent of cryo-ablation. 18. The method of claim 13, wherein cryo-ablation is guided by an ultrasound imaging externally or within the renal artery to visualize the progression of tissue freezing and provide real-time feedback for the procedure. 19. The method of claim 13, wherein the cryo-ablation is performed to treat medical conditions selected from the group consisting of resistant hypertension, heart failure, central sympathetic hyperactivity, metabolic syndrome, and obesity by cryo-ablating targets proximate to the renal artery. | A cryoablation catheter for sympathetic-nervous-system disorders, with a heat-transfer element that goes from a flat delivery profile to deployed shapes (including helical), plus deployable temperature probes. The probes measure temperature inside or outside the vessel for in-procedure feedback. | R | "a deployable temperature probe outside the cryo-ablation catheter, to perform direct temperature measurement of tissue, at a known distance from the balloon" (Claim 1) "the temperature monitoring system is configured to display the distance of the lethal isotherm as a guide during the cryo-ablation procedure" (Claim 10) "an ultrasound imaging system to visualize progression of tissue freezing, measure a diameter of frozen tissue volume, and calculate a position of a lethal isotherm to provide real-time feedback during a cryo-ablation procedure" (Claim 8) "monitoring progression of tissue freezing at the known distances and adjusting cryo-ablation parameters based on real-time feedback from the temperature measurements" (Claim 13) |  |  | 1 |  | 1 | [9 empty] | 1 | [10 empty] | 1 | [5 empty] | STEER MEDICAL INC | STEER MEDICAL INC | 2025-09-23 | 2025 | 2026-03-26 | 2026 | 2024-09-24 | 2024 | US202519337355 | [NA]
403 | US20260069350A1 | Laparoscopic Applicator, and System for Laparoscopic Ablation | The present invention describes a device for use in laparoscopic surgery, and a system for laparoscopic ablation. A method for perivascular and/or perineural neuromodulation is also disclosed. | 1. Device for use in laparoscopic surgery, having a shaft, which has a distal end and a proximal end, a holding unit, which is arranged at the distal end of the shaft and is adapted to hold and/or to release a tissue, at least two electrodes integrated in the holding unit, a force transmission unit, which is arranged on the shaft, is connected to the holding unit and is adapted to be moved relative to the shaft, an electric cable, which is arranged in the shaft, is connected to the at least two electrodes and is adapted to transmit an electrical signal that is to be received via the proximal end of the shaft to the at least two electrodes. 2. Device according to claim 1, having an articulated arrangement, which is arranged between the holding unit and the distal end of the shaft and is adapted to connect the distal end of the shaft, in particular movably, to the holding unit, wherein a movement of the force transmission unit leads to a movement of the holding unit and/or to the tissue being held and/or released by means of the holding unit. 3. Device according to claim 1, wherein the holding unit is configured as a finger gripper, in particular a three-fingered gripper, and the force transmission unit is configured as a pull- and push-rod, wherein a movement of the pull- and push-rod leads to the opening or closing of the finger gripper in order to grip or release the tissue. 4. Device according to claim 3, wherein fingers of the finger gripper each have an inner face oriented, in particular when the finger gripper is in a closed state, towards the tissue, in which inner face there is integrated at least one of the at least two electrodes, in particular one of a plurality of electrodes. 5. Device according to claim 1, wherein the holding unit is configured as a suction unit, wherein a movement of the force transmission unit leads to an angling of the suction unit relative to the shaft, and/or the suction unit is adapted to suction a tissue or to detach itself from the tissue. 6. Device according to claim 5, wherein the suction unit has: a suction cup, at least one suction hole, which is arranged in the suction cup, a suction channel, which is connected to the at least one suction hole, and/or a suction connection provided at the end of the suction channel, via which an external vacuum pump can be connected in order to generate a negative pressure in the suction cup and suction the tissue against the suction cup. 7. Device according to claim 6, wherein the suction cup is configured to be oval, in particular stadium-shaped or circular, along a vertical axis, and the at least two electrodes are arranged in the suction cup along the vertical axis and/or transverse to the vertical axis. 8. Device according to claim 6, wherein the suction cup is configured to be oval, in particular stadium-shaped or circular, along a vertical axis, and at least three electrodes are arranged in the suction cup along the vertical axis and/or transverse to the vertical axis, in particular spaced apart from the at least two electrodes. 9. Device according to claim 1, wherein the holding unit is configured as a clamshell gripper, in particular a double half-clamshell gripper, wherein a movement of the force transmission unit leads to an angling of the clamshell gripper relative to the shaft, and/or to the opening and/or to the closing of the clamshell gripper. 10. Device according to claim 9, wherein the clamshell gripper extends along a transverse axis and is adapted to open or to close along a vertical axis, wherein the at least two electrodes are integrated in an inside of at least one clamshell, in particular along a longitudinal axis. 11. Device according to claim 9, wherein the clamshell gripper extends along a transverse axis and is adapted to open or to close along a vertical axis, wherein at least three electrodes are integrated in an inside of at least one clamshell, in particular along and/or transverse to a longitudinal axis, spaced apart from the at least two electrodes. 12. Device according to claim 9, wherein the force transmission unit is configured as a first pull- and push-rod, which is arranged in the shaft and is connected to the articulated arrangement, wherein a movement of the first pull- and push-rod leads to the opening or closing of the clamshell gripper, and/or the force transmission unit is configured as a second pull- and push-rod, which is arranged in the shaft and is connected, in particular solely, to the clamshell gripper, wherein a movement of the second pull- and push-rod leads to an angling of the clamshell gripper relative to the shaft. 13. Device according to claim 1, wherein the force transmission unit and the holding unit are configured as a contiguous shape memory material, in particular shape memory alloy, and the holding unit forms a trap structure, in particular a trap basket, wherein a movement of the force transmission unit leads to the trap structure, in particular the trap basket, moving back into the shaft and folding, or leads to the trap structure, in particular the trap basket, moving out of the shaft and unfolding. 14. Device according to claim 13, wherein the trap structure has two individual and three contiguous splines, wherein the at least two electrodes and/or at least three electrodes are integrated in the splines, in particular in the contiguous splines. 15. System for laparoscopic ablation, having: a device according to claim 1; a signal generator arrangement which is connected or can be connected to the device; and a control and evaluation unit which is connected or can be connected to the device and/or to the signal generator arrangement. 16. System according to claim 15, wherein the signal generator arrangement has a first signal generator for generating a radiofrequency, RF, signal and a second signal generator for generating a signal for pulsed field ablation. 17. System according to claim 15, wherein the control and evaluation unit is configured to activate the first signal generator and/or the second signal generator to deliver a signal. 18. System according to claim 17, wherein the control and evaluation unit is configured to activate the first signal generator and/or the second signal generator to deliver a signal in dependence on at least one electrode temperature. 19. System according to claim 18, wherein the control and evaluation unit is configured to activate the first signal generator to deliver a signal if the electrode temperature falls below a temperature limit value and to activate the second signal generator to deliver a signal if the electrode temperature assumes or exceeds the temperature limit value. 20. System according to claim 15, wherein the control and evaluation unit is adapted to apply an electric current to the at least two electrodes and/or to measure a voltage via the at least three electrodes, wherein a nerve activity of the tissue and/or the local tissue impedance can be determined or is determined from the electric current and the electric voltage. 21. Method for perivascular and/or perineural neuromodulation, comprising: inserting an applicator into a patient; positioning a distal end of the applicator at a vessel or tissue or nerve of the patient such that at least two electrodes of the applicator are in contact perivascularly and/or perineurally with a perimeter of the vessel or tissue or nerve or lie against a perimeter of the vessel or tissue or nerve; carrying out a denervation by delivering energy via the at least two electrodes, wherein the delivery of energy is effected on the basis of a protocol which has at least one pulse of a pulsed field ablation (PFA); withdrawing the applicator. 22. Method according to claim 21, wherein an irreversible electroporation is carried out. 23. Method according to claim 21, wherein a reversible electroporation is carried out. 24. Method according to claim 21, wherein the applicator is inserted and/or withdrawn by a laparoscopic procedure. 25. Method according to claim 22, wherein the laparoscopic applicator is inserted into a patient's torso by means of a trocar. 26. Method according to claim 21, wherein the applicator is inserted and/or withdrawn by means of an open surgical procedure. 27. Method according to claim 21, wherein, after the applicator has been inserted, geometric unfolding of a distal end of the applicator is initiated and a form of the distal end that conforms to a vessel or tissue or nerve is thus achieved. 28. Method according to claim 21, wherein the electrodes of the applicator are brought into contact with a target vessel or target tissue or target nerve. 29. Method according to claim 21, further comprising: checking a positioning of the applicator by means of a local impedance measurement. 30. Method according to claim 21, wherein the protocol comprises only PFA pulses. 31. Method according to claim 21, wherein the protocol comprises a combination of PFA pulses and radiofrequency (RF) pulses. 32. Method according to claim 21, further comprising: choosing whether the delivery of energy for the denervation is to be carried out on the basis of a protocol by means of PFA pulses or on the basis of a protocol by means of a combination of PFA pulses and RF pulses. 33. Method according to claim 21, wherein a delivery of energy is effected once or multiple times at the same or different locations of the vessel or tissue or nerve. 34. Method according to claim 21, wherein perivascular and/or perineural neuromodulation occurs by carrying out the denervation. 35. Method according to claim 21, further comprising: before the denervation is carried out, quantifying a conductivity of the vessel or tissue or nerve by means of at least two stimulation and measuring electrodes. 36. Method according to claim 21, further comprising: after the denervation has been carried out, characterising the denervation by means of at least two stimulation and measuring electrodes. 37. Method according to claim 21, further comprising: measuring a temperature at a chosen ablation site. | A laparoscopic applicator and system for laparoscopic ablation, plus a method for perivascular and/or perineural neuromodulation. The device targets nerves around vessels via a laparoscopic approach. | R | "the control and evaluation unit is adapted to apply an electric current to the at least two electrodes and/or to measure a voltage via the at least three electrodes, wherein a nerve activity of the tissue and/or the local tissue impedance can be determined or is determined" (Claim 20) "to activate the first signal generator to deliver a signal if the electrode temperature falls below a temperature limit value and to activate the second signal generator to deliver a signal if the electrode temperature assumes or exceeds the temperature limit value" (Claim 19) "before the denervation is carried out, quantifying a conductivity of the vessel or tissue or nerve by means of at least two stimulation and measuring electrodes" (Claim 35) "after the denervation has been carried out, characterising the denervation by means of at least two stimulation and measuring electrodes" (Claim 36) | 1 |  |  |  | 1 | [6 empty] | 1 | [17 empty] | 1 |  | STOCKERT GMBH | STOCKERT GMBH | 2024-11-12 | 2024 | 2026-03-12 | 2026 | 2024-09-11 | 2024 | US202418944629 | CN121647795A¦DE102024126105A1¦EP4710881A1
404 | CN115802962A | Ultrasonic surgical aspirator for probing and ablating tissue | Systems and methods for probing a patient's problem using an ultrasound surgical aspirator that includes a tip for treating tissue of a patient and a console coupled to a handpiece. The console provides an AC drive signal to the ultrasound handpiece that vibrates the tip against the patient tissue. The AC drive signal includes a first component having a resonant frequency of the ultrasonic handpiece and a second component having a probing frequency less than the resonant frequency. The console measures a voltage and a current of the AC drive signal, calculates a resistance associated with the ultrasound handpiece based on the measured voltage and current, and then provides at least one of an audible, visual, or tactile indication based on the calculated resistance. | 1. An ultrasonic tool system for probing patient tissue, the system comprising: an ultrasonic handpiece comprising a tip for treating a distal region of patient tissue and at least one driver, the tip being coupled to the driver and an AC drive signal is applied to the driver to vibrate the tip, the ultrasonic handpiece defines a first passage for providing suction at a distal region of the tip and a first passage for supplying fluid to the distal region of the tip. a second pathway; and a console coupled to the ultrasonic handpiece and configured to generate said AC drive signal applied to said at least one driver of the ultrasonic handpiece to vibrate the tip of the ultrasonic handpiece, said The console includes: a first sensor for measuring a voltage of the AC drive signal; a second sensor for measuring a current of the AC drive signal; and a processor coupled to the first and second sensors and configured to: providing said AC drive signal to said at least one driver of an ultrasonic handpiece, said AC drive signal comprising a first component having a resonant frequency of the ultrasonic handpiece and a second component having a probe frequency less than said resonant frequency; Measure voltage and current of the AC drive signal using the first and second sensors; calculate a resistance associated with the ultrasonic handpiece based on the measured voltage and measured current; and provide an audible, visual, or tactile sense based on the calculated resistance at least one of the instructions. 2. The system of claim 1 , wherein the processor is configured to provide at least one of an auditory, visual, or tactile indication based on the calculated resistance by being configured to: identifying a property of the patient tissue; and providing at least one of an audible, visual, or tactile indication of the identified property. 3. The system of claim 1 or 2, wherein the AC drive signal provided to the ultrasound handpiece is defined by a base signal having a resonant frequency amplitude modulated according to the probe frequency. 4. The system of any one of claims 1-3, wherein the resonant frequency is about 25 kHz and the probing frequency is about 4 Hz. 5. The system of any one of claims 1-4, wherein the AC drive signal provided to the ultrasonic handpiece is configured to induce vibration of the tip insufficient to ablate patient tissue. 6. The system of claim 5 , wherein the AC drive signal provided to the ultrasonic handpiece is controlled by being configured to induce vibrations with a peak-to-peak displacement of less than or equal to 100 microns at the distal region of the tip. Configured to induce vibration of the tip insufficient to ablate patient tissue. 7. The system of any one of claims 1-6, wherein the AC drive signal is defined as a first AC drive signal, and the processor is configured to: determine that the ultrasonic tool system is set whether to operate in an exploratory mode or an ablation mode; in response to determining that the ultrasonic tool system is configured to operate in an exploratory mode, providing a first AC drive signal to an ultrasonic handpiece; and in response to determining that the ultrasonic tool system is configured to Operating in the ablation mode, a second AC drive signal is provided to the ultrasonic handpiece, the second AC drive signal being configured to induce vibrations of the tip sufficient to ablate tissue of the patient. 8. The system of claim 7 , wherein the second AC drive signal provided to the ultrasonic handpiece is configured to induce a peak-to-peak displacement greater than 100 microns and less than or equal to 300 microns at the distal region of the tip. Micron vibrations configured to induce vibrations of the tip sufficient to ablate tissue of the patient. 9. The system of claim 7 or 8, further comprising a switch communicatively coupled to the processor, the switch having a first setting and a second setting, wherein the processor is configured to: determine that the ultrasound tool system is set to operate in a probe mode in response to the switch being set to the first setting; and in response to the switch being set To the second setting, it is determined that the ultrasound tool system is set to operate in an ablation mode. 10. The system of any one of claims 7-9, wherein the processor is configured to, in response to determining that the ultrasound tool system is configured to operate in an ablation mode: through a first pathway defined by the ultrasound handpiece Suction is provided at the distal region of the tip; and fluid is supplied to the distal region of the tip through a second channel defined by the ultrasonic handpiece. 11. The system according to any one of claims 1-9, wherein the processor, by being configured to perform Based on the measured voltage and measured current, the equivalent current through the mechanical part of the ultrasonic handpiece is calculated; and based on the calculated equivalent current through the mechanical part of the ultrasonic handpiece, the calculation and ultrasonic The resistance associated with the handpiece. 12. The system of claim 11 , wherein the processor, by being configured to perform the following, is configured to calculate an ultrasonic handpiece based on the calculated equivalent current through a mechanical component of the ultrasonic handpiece resistance associated with the piece: Calculate the first amplitude of the measured voltage at the probe frequency, the second amplitude of the calculated equivalent current through the mechanical part of the ultrasonic handpiece at the probe frequency, and the second amplitude at the probe frequency the phase difference between the measured voltage and the calculated equivalent current through the mechanical components of the ultrasonic handpiece; and based on the calculated first amplitude, the calculated second amplitude, and the calculated phase difference, calculate the ultrasonic The real part of the impedance of the handpiece. 13. The system according to any one of claims 1-12, wherein the processor is configured to provide an audible, visual or tactile indication based on the calculated resistance by being configured to perform At least one of: calculating a difference between the calculated resistance and the no-load resistance of the ultrasonic handpiece; and providing at least one of an audible, visual, or tactile indication based on the calculated difference. 14. The system of claim 13 , wherein the processor is configured to provide at least one of an auditory, visual, or tactile indication based on the calculated difference by being configured to: The identified difference identifies a property of the patient tissue; and provides at least one of an audible, visual, or tactile indication of the identified property. 15. The system of any one of claims 1-14, wherein the console further comprises a memory coupled to the processor that stores tissue property data indicative of potential tissue properties, and for potential tissue One or more values specific to the potential tissue property for each of the potential tissue properties, and providing at least one of an auditory, visual or tactile indication based on the calculated resistance comprises: the one or more values of the property and the calculated resistance, identifying one of the potential tissue properties indicated by the tissue property data as a property of the patient tissue; and providing an audible sense of the identified property of the patient tissue , at least one of visual or tactile indications. 16. The system of claim 15 , wherein the processor is configured based on the one or more values specific to the underlying tissue property and the calculated resistance by being configured to perform One of the potential tissue properties indicated by the tissue property data is identified: calculating a difference between the calculated resistance and the no-load resistance of the ultrasonic handpiece; and identifying a potential tissue property indicated by the tissue property data based on the one or more values specific to the potential tissue property and the calculated difference The one underlying tissue property in . 17. The system of claim 13 , 14 , or 16 , wherein the processor is configured to determine the absence of the ultrasound handpiece by being configured to perform the following in response to the ultrasound handpiece being connected to the console. Load resistance: with the ultrasonic handpiece in an unloaded state, an AC drive signal is provided to the ultrasonic handpiece; with the ultrasonic handpiece in an unloaded state, measured using the first and second sensors a second voltage and a second current of the AC drive signal; and calculating an unloaded resistance of the ultrasonic handpiece based on the measured second voltage of the AC drive signal and the measured second current of the AC drive signal. 18. The system of claim 13 , 14 , or 16 , wherein the processor, by being configured to read from a memory integrated with the ultrasound handpiece in response to the ultrasound handpiece being connected to the console, indicates the The data of the no-load resistance is configured to determine the no-load resistance of the ultrasonic handpiece. 19. A method of probing patient tissue using an ultrasonic tool system comprising an ultrasonic handpiece having a tip for treating tissue of a patient and at least one driver, the tip being coupled to the driver and An AC drive signal is applied to the driver to vibrate the tip, the method comprising: supplying fluid to a distal region of the tip through at least a portion of the ultrasonic handpiece; at the distal region of the tip through at least a portion of the ultrasonic handpiece providing suction; providing an AC drive signal to the ultrasonic handpiece, the AC drive signal comprising a first component having a resonant frequency of the ultrasonic handpiece and a second component having a probing frequency less than the resonant frequency; measuring the AC drive at the tip a voltage and a current of the signal; calculating a resistance associated with the ultrasonic handpiece based on the measured voltage and the measured current; and providing at least one of an audible, visual, or tactile indication based on the calculated resistance. 20. The method of claim 19 , wherein providing at least one of an audible, visual, or tactile indication based on the calculated resistance comprises: identifying a property of the patient's tissue based on the calculated resistance; and providing an indication of the identified property. At least one of auditory, visual or tactile indications. 21. twenty one. The method of claim 19 or 20, wherein providing the AC drive signal to the ultrasound handpiece comprises: generating a base signal having a resonant frequency; and modulating the amplitude of the base signal in accordance with the probe frequency. 22. twenty two. The method of any one of claims 19-21, wherein the resonant frequency is about 25 kHz and the probing frequency is about 4 Hz. 23. twenty three. The method of any one of claims 19-22, wherein providing an AC drive signal to the ultrasonic handpiece comprises providing an AC drive signal to the ultrasonic handpiece that induces vibrations of the tip that are insufficient to ablate patient tissue. 24. twenty four. The method of claim 23 , wherein providing an AC drive signal to the ultrasonic handpiece that induces vibrations of the tip that are insufficient to ablate patient tissue comprises: providing the ultrasonic handpiece with an AC drive signal that induces peak-to-peak vibration at a distal region of the tip. AC drive signal for vibration with displacement less than or equal to 100 microns. 25. The method according to any one of claims 19-24, further comprising: determining that the ultrasound tool system is set to operate in a probing mode; In response to determining that the ultrasound tool system is configured to operate in a probing mode, providing an AC drive signal to the ultrasound handpiece; determining that the ultrasound tool system is configured to operate in an ablation mode; and in response to determining that the ultrasound tool system is configured to operate in an ablation mode, Another AC drive signal is provided to the ultrasonic handpiece that induces vibrations of the tip sufficient to ablate the patient's tissue. 26. The method of claim 25 , wherein providing the another AC drive signal to the ultrasonic handpiece to induce vibrations of the tip sufficient to ablate the patient's tissue comprises providing the ultrasonic handpiece to induce a peak at a distal region of the tip. - A vibrating AC drive signal with a peak displacement greater than 100 microns and less than or equal to 300 microns. 27. The method of claim 25 or 26, wherein the ultrasonic tool system further comprises a switch having a first setting and a second setting, and the method further comprises: monitoring a state of the switch to determine whether the switch is set to Whether the first setting or the second setting; determining that the switch is set to the first setting; in response to determining that the switch is set to the first setting, determining that the ultrasound tool system is set to operate in a probe mode; determining that the switch is set to the second setting; responding Upon determining that the switch is set to the second setting, it is determined that the ultrasound tool system is set to operate in the ablation mode. 28. The method of any one of claims 25-27, further comprising, in response to determining that the ultrasonic tool system is configured to operate in an ablation mode, supplying fluid to a distal region of the tip through at least a portion of the ultrasonic handpiece and Suction is provided at a distal region of the tip through at least a portion of the ultrasonic handpiece. 29. The method of any one of claims 19-28, wherein calculating a resistance associated with the ultrasonic handpiece based on the measured voltage and the measured current comprises: calculating based on the measured voltage and the measured current An equivalent current through a mechanical component of the ultrasonic handpiece; based on the calculated equivalent current through a mechanical component of the ultrasonic handpiece, a resistance associated with the ultrasonic handpiece is calculated. 30. The method of claim 29 , wherein calculating the resistance associated with the ultrasonic handpiece based on the calculated equivalent current through a mechanical component of the ultrasonic handpiece comprises: One amplitude, the calculated equivalent current through the mechanical part of the ultrasonic handpiece, the second amplitude at the probe frequency, and the difference between the measured voltage and the calculated equivalent current through the mechanical part of the ultrasonic handpiece at the probe frequency a phase difference at a frequency; and calculating a real part of an impedance of the ultrasonic handpiece based on the calculated first amplitude, the calculated second amplitude, and the calculated phase difference. 31. The method of any one of claims 19-30, wherein providing at least one of an audible, visual, or tactile indication based on the calculated resistance comprises: calculating the calculated resistance and the no-load resistance of the ultrasonic handpiece providing at least one of an auditory, visual or tactile indication based on the calculated difference. 32. The method of claim 31 , wherein providing at least one of an auditory, visual, or tactile indication based on the calculated difference comprises: identifying a property of the patient tissue based on the calculated difference; providing an auditory indication of the identified property. , at least one of visual or tactile indications. 33. The method according to any one of claims 19-32, wherein the ultrasound tool system further comprises a memory storing tissue property data indicating potential tissue properties, and for one of the potential tissue properties One or more values specific to each potential tissue property, and providing at least one of an audible, visual, or tactile indication based on the calculated resistance comprises: based on all values specific to the potential tissue property The one or more values and the calculated resistance identify one of the potential tissue properties indicated by the tissue property data as a property of the patient tissue; and provide an audible, visual, or at least one of the tactile indications. 34. The method of claim 33 , wherein identifying one of the potential tissue properties indicated by the tissue property data based on the one or more values specific to the potential tissue property and the calculated resistance comprises: calculating a difference between the calculated resistance and an unloaded resistance of the ultrasound handpiece; and identifying a potential tissue property indicated by the tissue property data based on the one or more values specific to the potential tissue property and the calculated difference The one underlying tissue property in . 35. The method of claim 31 , 32 or 34, further comprising determining the no-load resistance of the ultrasonic handpiece by: positioning the ultrasonic handpiece in an unloaded state; where the ultrasonic handpiece is positioned in an unloaded state Bottom: providing an AC drive signal to the ultrasonic handpiece; measuring a second voltage and a second current of the AC drive signal provided to the ultrasonic handpiece; and measuring a second voltage and a measured second current based on the AC drive signal Calculate the no-load resistance of the ultrasound handpiece. 36. The method of claim 31 , 32 or 34, further comprising determining the unloaded resistance of the ultrasound handpiece by reading data indicative of the unloaded resistance from a memory integrated with the ultrasound handpiece. 37. An ultrasonic tool system for probing patient tissue, the system comprising: an ultrasonic handpiece comprising a tip for treating a distal region of patient tissue and at least one driver, the tip being coupled to the driver and An AC drive signal is applied to the driver to vibrate the tip, the ultrasonic handpiece defines a first passage for providing suction at a distal region of the tip and a second passage for supplying fluid to the distal region of the tip. two channels; a console coupled to the ultrasonic handpiece and configured to generate the AC drive signal applied to the at least one driver of the ultrasonic handpiece to vibrate the tip of the ultrasonic handpiece, the console comprising : a first sensor for measuring the voltage of the AC drive signal; a second sensor for measuring the current of the AC drive signal; a processor coupled to the first and second sensors and configured to: providing to the at least one driver of the ultrasonic handpiece, the AC drive signal induces vibrations insufficient to ablate patient tissue at a distal region of the tip; measuring the voltage and current of the AC drive signal using the first and second sensors; and At least one of an audible, visual or tactile indication is provided based on the measured voltage and current. 38. The system of claim 37 , wherein the processor is configured to provide at least one of an audible, visual or tactile indication based on the measured voltage and current by being configured to: Identifying a property of the patient's tissue based on the measured voltage and current; providing at least one of an audible, visual, or tactile indication of the identified property. 39. The system of claim 38, wherein the processor is configured by being configured to perform to identify the characteristics of the patient's tissue based on the measured voltage and current: to calculate the equivalent current through the mechanical components of the ultrasonic handpiece based on the measured voltage and measured current; and to calculate the equivalent current through the mechanical components of the ultrasonic handpiece based on the calculated The equivalent current to identify the characteristics of the patient tissue. 40. The system of claim 39 , wherein the AC drive signal provided to the ultrasonic handpiece includes a first component having a resonant frequency of the ultrasonic handpiece and a second component having a probe frequency less than the resonant frequency, and the processor , configured to identify a characteristic of patient tissue based on the calculated equivalent current through a mechanical component of the ultrasonic handpiece by: calculating a first amplitude of the measured voltage at the probe frequency, calculating The second amplitude at the probe frequency of the calculated equivalent current through the mechanical part of the ultrasonic handpiece, and the phase at the probe frequency between the measured voltage and the calculated equivalent current through the mechanical part of the ultrasonic handpiece difference; identifying a characteristic of the patient tissue based on the calculated first amplitude, the calculated second amplitude, and the calculated phase difference. 41. The system according to any one of claims 37-40, wherein the processor is configured to provide an audible, visual or tactile indication based on the measured voltage and current by being configured to perform at least one of: identifying a characteristic of the patient tissue based on the measured voltage, the measured current, and the unloaded resistance of the ultrasonic handpiece; and providing at least one of an audible, visual, or tactile indication of the identified characteristic. 42. The system of any one of claims 37-41, wherein the console further comprises a memory coupled to the processor that stores tissue property data indicative of potential tissue properties, and for potential tissue One or more values specific to each of the potential tissue properties in the properties, and the processor is configured to provide an auditory sense based on the measured voltage and current by being configured to perform the following at least one of, a visual or a tactile indication: identifying one of the potential tissue properties indicated by the tissue property data based on the one or more values specific to the potential tissue property and the measured voltage and current , as a property of the patient tissue; and providing at least one of an audible, visual, or tactile indication of the identified property of the patient tissue. 43. The system of claim 42 , wherein the processor is configured to identify an ultrasonic handpiece caused by Said one of the potential tissue properties indicated by the tissue property data is said one of the potential tissue properties. 44. The system of claim 41 or 43, wherein the processor is configured to determine the no-load resistance of the ultrasonic handpiece by being configured to perform the following in response to the ultrasonic handpiece being connected to the console : When the ultrasonic handpiece is in a no-load state, an AC drive signal is provided to the ultrasonic handpiece; When the ultrasonic handpiece is in an unloaded state, the first and second sensors are used to measure the AC drive signal provided to the ultrasonic handpiece a second voltage and a second current of the AC drive signal; and calculating an unloaded resistance of the ultrasonic handpiece based on the measured second voltage and the measured second current of the AC drive signal. 45. The system of claim 41 or 43, wherein the processor, by being configured to read from a memory integrated with the ultrasound handpiece in response to the ultrasound handpiece being connected to the console, indicates no load resistance. data configured to determine the no-load resistance of the ultrasonic handpiece. 46. A method of probing tissue in a patient using an ultrasonic tool system comprising an ultrasonic handpiece, The ultrasonic handpiece has a tip for treating tissue of a patient and at least one driver, the tip is coupled to the driver and an AC drive signal is applied to the driver to vibrate the tip, the method comprising: handheld by ultrasound supplying fluid to the distal region of the tip by at least a portion of the piece; providing suction at the distal region of the tip through at least a portion of the ultrasonic handpiece; providing an AC drive signal to the ultrasonic handpiece, the AC drive signal being at the distal region of the tip Inducing insufficient vibration at the side region to ablate patient tissue; measuring voltage and current of the AC drive signal; and providing at least one of an audible, visual, or tactile indication based on the measured voltage and current. 47. The method of claim 46 , wherein providing at least one of an audible, visual, or tactile indication based on the measured voltage and current comprises: identifying a characteristic of the patient's tissue based on the measured voltage and current; providing the identified at least one of an auditory, visual or tactile indication of a characteristic of the 48. The method of claim 47, wherein identifying characteristics of the patient's tissue based on the measured voltage and current comprises: calculating an equivalent current through a mechanical component of the ultrasonic handpiece based on the measured voltage and the measured current; The calculated equivalent currents through the mechanical components of the ultrasound handpiece identify characteristics of the patient tissue. 49. The method of claim 48, wherein the AC drive signal provided to the ultrasonic handpiece includes a first component having a resonant frequency of the ultrasonic handpiece and a second component having a probe frequency less than the resonant frequency, and based on calculating Characterizing the patient tissue by calculating the equivalent current through the mechanical components of the ultrasonic handpiece includes: calculating the first amplitude of the measured voltage at the probe frequency, the calculated equivalent current through the mechanical component of the ultrasonic handpiece at the second amplitude at the probe frequency, and the phase difference at the probe frequency between the measured voltage and the calculated equivalent current through the mechanical part of the ultrasonic handpiece; based on the calculated first amplitude, the calculated first Two amplitude and phase differences are calculated to identify patient tissue properties. 50. The method of any one of claims 46-49, further comprising identifying a characteristic of the patient's tissue based on the measured voltage, the measured current, and the unloaded resistance of the ultrasound handpiece. 51. The method according to any one of claims 46-50, wherein the ultrasound tool system further comprises a memory storing tissue property data indicative of potential tissue properties, and for each of the potential tissue properties One or more values specific to a potential tissue property, and providing at least one of an audible, visual, or tactile indication based on the measured voltage and current includes: The one or more values and the measured voltage and current identify one of the potential tissue properties indicated by the tissue property data as a property of the patient tissue; providing an audible, At least one of visual or tactile indications. 52. The method of claim 51 , further comprising identifying a tissue indicated by tissue property data based on the one or more values specific to the underlying tissue property, the measured voltage and current, and the no-load resistance of the ultrasonic handpiece. The one of the potential tissue properties of . 53. The method of claim 50 or 52, further comprising determining the no-load resistance of the ultrasonic handpiece by: positioning the ultrasonic handpiece in an unloaded state; and when the ultrasonic handpiece is positioned in an unloaded state : providing an AC drive signal to the ultrasonic handpiece; measuring a second voltage and a second current of the AC drive signal provided to the ultrasonic handpiece; calculating an ultrasonic handpiece based on the measured second voltage and the measured second current of the AC drive signal The no-load resistance of the device. 54. The method of claim 50 or 52, further comprising determining the no-load resistance of the ultrasound handpiece by reading data indicative of the no-load resistance from a memory integrated with the ultrasound handpiece. 55. An ultrasonic tool system comprising: an ultrasonic handpiece comprising a tip having a distal region for treating tissue of a patient and at least one driver, the tip coupled to the driver and an AC drive signal applied to the driver vibrating the tip, the ultrasonic handpiece defining a first channel for providing suction at a distal region of the tip and a second channel for supplying fluid to the distal region of the tip; a console coupled to an ultrasonic handpiece and configured to generate the AC drive signal applied to the at least one driver of the ultrasonic handpiece to vibrate the tip of the ultrasonic handpiece; and a switch communicatively coupled to the console, the switch having a first A setting and a second setting, wherein, in response to the switch being set to the first setting, the console is configured to operate the ultrasound handpiece in an exploratory mode, and in response to the switch being set to the second setting, the console is configured to Configured to operate the ultrasound handpiece in ablation mode. 56. A method for operating an ultrasonic tool system comprising an ultrasonic handpiece having a tip for treating tissue of a patient and at least one driver, the tip coupled to the driver and An AC drive signal is applied to the driver to vibrate the tip, the ultrasonic tool system includes a switch having a first setting and a second setting, the method comprising: providing the AC drive signal to the at least one of the ultrasonic handpieces a driver to vibrate the tip of the ultrasonic handpiece; supply fluid to the distal region of the tip through at least a portion of the ultrasonic handpiece; provide suction at the distal region of the tip through at least a portion of the ultrasonic handpiece ; monitoring the state of the switch to determine whether the switch is set to the first setting or the second setting; determining that the switch is set to the first setting; in response to determining that the switch is set to the operating the ultrasound handpiece in an exploratory mode; determining that the switch is set to the second setting; and in response to determining that the switch is set to the second setting, operating the ultrasound handpiece in an ablation mode Ultrasound Handpiece. | An ultrasonic surgical aspirator whose console drives a vibrating tip with both a resonant-frequency and a sub-resonant probe component, while measuring voltage. The probe component lets the same tool sense tissue in addition to ablating/aspirating it. | R | "providing said AC drive signal to said at least one driver of an ultrasonic handpiece, said AC drive signal comprising a first component having a resonant frequency of the ultrasonic handpiece and a second component having a probe frequency less than said resonant frequency" (Claim 1) "calculate a resistance associated with the ultrasonic handpiece based on the measured voltage and measured current; and provide an audible, visual, or tactile indication based on the calculated resistance" (Claim 1) "based on the one or more values specific to the potential tissue property and the calculated resistance, identifying one of the potential tissue properties indicated by the tissue property data as a property of the patient tissue" (Claim 15) "the AC drive signal provided to the ultrasonic handpiece is configured to induce vibration of the tip insufficient to ablate patient tissue" (Claim 5) | 1 |  | 1 | [12 empty] | 1 |  |  |  |  | 1 | [10 empty] | STRYKER CORP | STRYKER CORP | 2021-06-04 | 2021 | 2023-03-14 | 2023 | 2020-06-04 | 2020 | CN20218042977 | AU2021283983A1¦CA3181229A1¦EP4161411A1¦EP4161411B1¦EP4523638A2¦EP4523638A3¦JP2023528485A¦JP2026048663A¦JP7777546B2¦US20230200835A1¦WO2021248062A1
405 | CN115969500A | Multifunctional PFA pulse ablation catheter assembly | The invention discloses a multifunctional PFA pulse ablation catheter assembly which comprises an ablation catheter. The ablation catheter comprises a catheter body; the ablation catheter head end integrally extends outwards from the catheter body to form a preset shape and is used for ablating surrounding cells; the mapping catheter penetrates through the inside of the ablation catheter, forms a preset shape, extends to the far end of the head end of the ablation catheter, and is used for recording changes of surrounding electric signals; a plurality of electrodes are arranged outside the ablation catheter and the mapping catheter, and the electrodes of the ablation catheter form a plurality of electric fields. According to the device, the mapping catheter is used for replacing a guide steel wire to guide and support the ablation catheter, meanwhile, a three-dimensional mapping system and other mapping catheters are replaced to evaluate a pulmonary vein isolation result after ablation, cost is reduced, and meanwhile, changes of cardiac cavity potential can be recorded during ablation. | 1. A multifunctional PFA pulse ablation catheter assembly, characterized in that: it comprises an ablation catheter (1); the ablation catheter (1) comprises a tube body (101); the ablation catheter head end (102), from the tube body (101) extending outward integrally to form a predetermined shape for ablation of surrounding cells; the mapping catheter (2) passes through the inside of the ablation catheter (1) and forms a predetermined shape, and extends to the head end of the ablation catheter ( 102) for recording changes in surrounding electrical signals; the external settings of the ablation catheter (1) and the mapping catheter (2) are all provided with a plurality of electrodes (3), and the electrodes (3) of the ablation catheter (1) Multiple electric fields are formed. 2. A kind of multifunctional PFA pulse ablation catheter assembly according to claim 1, is characterized in that: described ablation catheter tip (102) is annular, petal-shaped, divergent shape or ball net shape. 3. A kind of multifunctional PFA pulse ablation catheter assembly according to claim 1, is characterized in that: the head end of described mapping catheter (2) is ring-shaped. 4. A kind of multifunctional PFA pulse ablation catheter assembly according to claim 1, it is characterized in that: the number of electrodes outside the head end (102) of the ablation catheter is 8-30, and the electrodes on the head end of the mapping catheter (2) The quantity is 8-20pcs. 5. A kind of multifunctional PFA pulse ablation catheter assembly according to claim 1, is characterized in that: the length of described electrode (B) is 1-5mm, and described electrode (B) is wrapped in described ablation catheter (1) and described mapping The outside of the catheter (2). 6. A kind of multifunctional PFA pulse ablation catheter assembly according to claim 1, is characterized in that: the head end of described mapping catheter (2) is soft material. 7. A multifunctional PFA pulse ablation catheter assembly according to claim 1, characterized in that: the diameter of the mapping catheter (2) is smaller than the diameter of the lumen of the tube body (101). | A multifunctional PFA ablation catheter assembly: an ablation catheter with a preset distal shape, plus a mapping catheter inside it that also takes a defined shape and extends to the tip. The same device records signals (mapping) and delivers PFA energy. | R | "the mapping catheter (2) passes through the inside of the ablation catheter (1) and forms a predetermined shape, and extends to the head end of the ablation catheter (102) for recording changes in surrounding electrical signals" (Claim 1) "the electrodes (3) of the ablation catheter (1) form multiple electric fields" (Claim 1) "replaces a three-dimensional mapping system and other mapping catheters to evaluate a pulmonary vein isolation result after ablation… can record changes of cardiac cavity potential during ablation" (Abstract) | 1 | [10 empty] | 1 | [12 empty] | 1 | [6 empty] | SUZHOU XINXIAN MEDICAL TECH CO LTD | SUZHOU XINXIAN MEDICAL TECH CO LTD | 2022-06-09 | 2022 | 2023-04-18 | 2023 | 2022-06-09 | 2022 | CN202210653906 | [NA]
406 | CA3213212A1 | LASER SYSTEM AND METHOD FOR DETECTING AND PROCESSING INFORMATION | The present disclosure provides a laser system suitable for modification of a calcified blood vessel, comprising: a laser source; a feedback controller, configured to regulate a dosimetry of the laser source to produce spatially and/or temporally modulated laser light; a catheter comprising a first optical delivery element, the first optical delivery element configured to guide the modulated laser light to an in-vivo object in the blood vessel; and a detecting element, configured to detect one or more physical, chemical, mechanical and/or dimensional characteristics of an area of the in-vivo object in real-time, wherein the feedback controller is configured to process the real-time detected information pertaining to the one or more physical, chemical, mechanical and/or dimensional characteristics of the area in real-time, wherein the feedback controller is further configured to regulate in real-time the dosimetry of the laser source based on the real-time-detected information for a controlled formation of a porous structure and/or a zone of denaturized tissue in the in-vivo object. | 1. i. A laser system suitable for modification of a calcified blood vessel (201), comprising a laser source (101) a feedback controller (106), configured to regulate a dosimetry of the laser source (101) to produce spatially and/or temporally modulated laser light a catheter (103) comprising a first optical delivery element (102), the first optical delivery element (102) configured to guide the modulated laser light to an in-vivo object (202) in the blood vessel (201); and a detecting element (105), configured to detect one or more physical, chemical, mechanical and/or dimensional characteristics of an area of the in-vivo object in realtime, wherein the feedback controller (106) is configured to process the real-time detected information pertaining to the one or more physical, chemical, mechanical and/or dimensional characteristics of the area in real-time, wherein the feedback controller (106) is further configured to regulate in real-time the dosimetry of the laser source (101) based on the real-time-detected information for a controlled formation of a porous structure and/or a zone of denaturized tissue in the in-vivo object (202). 2. The laser system of claim 1, wherein the one or more physical, chemical, mechanical and/or dimensional characteristics comprise at least one of: a position of the in-vivo object (202) a composition of the in-vivo object (202) a dimension of the in-vivo object (202) a temperature of the in-vivo object (202) and/or of an environment of the in-vivo object (202) a stress distribution of the area of the in-vivo object (202) a stress distribution near a vessel wall (202a) of the blood vessel (201) a light scattering induced by the area of the in-vivo object (202) a conductivity of the in-vivo object (202) a Young’s modulus of the area of the in-vivo object (202) a characteristic pertaining to a porous structure and/or a zone of denaturized tissue on the in-vivo object (202) a lumen area of the blood vessel (201) Date Reçue/Date Received 2023-09-20 a compliance of the blood vessel (201) a strength of the in-vivo object (202); and a plasticity threshold of the in-vivo object (202). 3. The laser system of claims 1 or 2, wherein the feedback controller (106) is configured to control, based on the real-time detected information, a distance between the optical delivery element (102) and the in-vivo object (202) in the course of an irradiation by the laser source. 4. The laser system of any one of the preceding claims, wherein the first optical delivery element (102) comprises a bundle of optical fibers and/or is configured to multiplex a plurality of laser outputs of the laser source (101) into one fiber at an input of the first optical delivery element (102). 5. The laser system of any one of the preceding claims, further comprising a balloon (104), configured to be inflated and deflated in the bloodvessel (201); wherein the feedback controller (106) is further configured to control a gas pressure in the balloon (104), and/or to implement a desired positioning of the balloon (104) in realtime based on the real-time detected information. 6. The laser system of any one of the preceding claims, wherein the feedback controller (106) comprises and/or is coupled to a remote highperformance computer, a hybrid quantum-classical computational facility, and/or a quantum computer (io6d); and/or the feedback controller (106) comprises and/or is connected to a storage device, the storage device storing an offline settings table, wherein the settings table is calculated by a remote high-performance computer, a remote hybrid quantum-classical computational facility, and/or a remote quantum computer (io6d). 7. A method for detecting and processing information, comprising a) detecting one or more physical, chemical, mechanical and/or dimensional characteristics of an area of an in-vivo object (202) in a calcified blood vessel (201); and b) processing the detected information pertaining to the physical, chemical, mechanical and/or dimensional characteristics to acquire a property of a porous structure formation and/or a formation of a zone of denaturized tissue in the in-vivo object (202), Date Reçue/Date Received 2023-09-20 wherein the information pertaining to the physical, chemical, mechanical and/or dimensional characteristics of the in-vivo object (202) area is detected and processed in real-time during the porous structure formation and/or the formation of the zone of denaturized tissue. 8. The method of claim 7, wherein the processing of the detected information comprises generating a value for a dosimetry of a laser source (101) in real-time based on the detected information pertaining to the physical, chemical, mechanical and/or dimensional characteristics of the in-vivo object (202) area, wherein the porous structure formation and/or the formation of the zone of denaturized tissue is induced by temporally and/or spatially modulated laser light generated by the laser source (101). 9. The method of claim 8, wherein the generating of the value for the dosimetry of a laser source (101) comprises: generating a time interval between two laser pulses to be longer than a time it takes for a fluid to fill the porous structure after a porous structure formation, wherein the time interval is generated based on the pore size distribution of the porous structure. 10. The method of any one of the claims 7 to 9, wherein the detecting of the physical, chemical, mechanical and/or dimensional characteristics of the area of the in-vivo object comprises detecting a temperature in the area, wherein the value for the dosimetry of the laser source (101) is generated if the temperature is within a predetermined range. 11. The method of any one of claims 7 to 10, wherein the physical, chemical, mechanical and/or dimensional characteristics comprises a characteristic of a scattered light, and the processing of the detected information further comprises: calculating a pore size distribution of the porous structure based on the scattered light. 12. The method of any one of claims 7 to 11, wherein the processing of the detected information further comprises calculating a stress distribution and/or a temperature distribution in the area of the in-vivo object (202); and/or mapping a stress distribution to a temperature distribution and/or evaluating the correlation between a stress distribution and a temperature distribution. Date Reçue/Date Received 2023-09-20 13- The method of any one of daims 7 to 12, further comprising detecting a physical, chemical, mechanical and/or dimensional characteristic of the area of the in-vivo object (202) in the calcified blood vessel (201) before, during, and/or after the porous structure formation and/or the formation of the zone of 5 denaturized tissue in the in-vivo object (202) processing the detected information pertaining to the physical, chemical, mechanical and/or dimensional characteristic of the area of the in-vivo object (202) in the calcified blood vessel (201) before, during and/or after the porous structure formation and/or the formation of the zone of denaturized tissue in the in-vivo object (202) to 10 identify a location where a stress is higher than a predetermined value. 14. The method of any one of the daims 7 to 13, further comprising detecting a physical, chemical, mechanical and/or dimensional characteristic of the area of an in-vivo object (202) on the blood vessel wall (202a) after the blood vessel 15 wall (202a) has undergone a mechanical action. 15 The method of any one of the claims 7 to 14, wherein the processing of the detected information is performed in real time using a remote high-performance computer, a hybrid quantum-classical computational facility, and/or a quantum computer (io6d). Date Reçue/Date Received 2023-09-20 | A laser system for modifying calcified blood vessels using a laser source, a feedback controller that produces spatially/temporally modulated light, a catheter optical delivery element, and a detector. The detector's signal feeds back into laser dosing for safer plaque modification. | R | "a detecting element, configured to detect one or more physical, chemical, mechanical and/or dimensional characteristics of an area of the in-vivo object in real-time" (Claim 1) "the feedback controller is further configured to regulate in real-time the dosimetry of the laser source based on the real-time-detected information for a controlled formation of a porous structure and/or a zone of denaturized tissue in the in-vivo object" (Claim 1) "calculating a stress distribution and/or a temperature distribution in the area of the in-vivo object; and/or mapping a stress distribution to a temperature distribution and/or evaluating the correlation between a stress distribution and a temperature distribution" (Claim 12) |  | 1 | [11 empty] | 1 | [17 empty] | TERRA QUANTUM AG | TERRA QUANTUM AG | 2023-09-20 | 2023 | 2024-03-23 | 2024 | 2022-09-23 | 2022 | CA20233213212 | AU2023229530A1¦AU2023229530B2¦AU2023229602A1¦AU2023229602B2¦AU2023229607A1¦AU2023229607B2¦CA3213220A1¦CA3213431A1¦CN117716325A¦CN117752412A¦CN117752491A¦CN117752954A¦CN117769694A¦EP4342407A1¦EP4342408A1¦EP4342435A1¦EP4377769A1¦EP4377771A1¦JP2024046744A¦JP2024046746A¦JP2024046747A¦JP7755322B2¦JP7807088B2¦KR20240041830A¦KR20240041831A¦KR20240042353A¦TW202319888A¦TW202324065A¦US12096112B2¦US12132983B2¦US12501150B2¦US20230031871A1¦US20230032467A1¦US20240099771A1¦US20240099888A1¦US20240130789A1¦US20250024138A1¦US20250030940A1¦WO2023009806A1¦WO2023009853A1
407 | CN119700291A | CT-guided tumor ablation dose regulation and control method and system | The invention relates to the technical field of tumor treatment, and discloses a CT-guided tumor ablation dose regulation and control method and system, and the method comprises the steps: obtaining the three-dimensional contour data and gray value distribution of a tumor through CT scanning, calculating the volume and average density of the tumor, and obtaining the initial ablation dose. Initial parameters of ablation equipment are set, an ablation program is started, data are updated at specific time intervals in the ablation process, and the real-time ablation dosage is calculated. And comparing the real-time ablation dose with the initial ablation dose, and adjusting the power of ablation equipment according to a specific formula and a comparison result. The system comprises an imaging and parameter acquisition module, an initial dose calculation module, an ablation equipment control module, a real-time monitoring module, a dose comparison and regulation module and the like. The ablation dosage can be accurately regulated and controlled according to the real-time condition of a tumor, the ablation treatment effect is improved, the damage risk to normal tissue is reduced, an alarm can be given to prompt an operator when the dosage is abnormal, and the high clinical application value is achieved. | 1. L A method for controlling the dose of tumor ablation under CT guidance, characterized in that the method comprises: starting a CT scanning device to perform three-dimensional imaging of the tumor site, obtaining three-dimensional contour data of the tumor, and calculating the tumor volume based on the contour data; calculating the average density of the tumor through the gray value distribution of the CT image; and calculating the initial ablation dose based on the tumor volume and the average density. Add calculation; setting the initial parameters of the ablation device according to the initial ablation dose, starting the ablation program, and during the ablation process, starting the CT scanning device again at a specific interval of time nine to update the three-dimensional contour data and gray value distribution of the tumor, repeating the above volume, density and dose calculation steps to obtain the real-time ablation dose Oyeoi; comparing the real-time ablation dose Oyeoi with the initial ablation dose Dinit, and when D3I > Dinit XQ +. ), reducing the power of the ablation device, where & is the preset upper limit floating ratio of the dose; when D3I < Dinit XQ +.init When X(1-(3)), the power of the ablation device is increased, and X is the preset floating ratio of the lower limit of the dose. 2.The tumor ablation dose control method according to claim 1 is characterized in that the tumor volume V is calculated according to the contour data through the formula V = £1 Si × hi, where Si is the cross-sectional area of the i-th layer, hi is the distance between two adjacent cross-sectional layers, and η is the number of layers. 3The method for controlling the dose of tumor ablation according to claim 1 is characterized in that the average density of the tumor is calculated by combining the gray value distribution of the CT image with the formula: Pj is the density value corresponding to the jth pixel point, and m is the total number of pixels selected; the method for selecting the pixel points is: taking the geometric center of the tumor as the center and r as the radius to draw a circle, and selecting the pixel points in the circle to participate in the calculation, and r is determined according to the maximum diameter R of the tumor through the formula r = and. 4. The tumor ablation dose control method according to claim 1 is characterized in that the formula for calculating the initial ablation dose Dmit is: Dinit - k XVX ρ; where Α; is the empirical coefficient, the value of Α; is related to the tissue type of the tumor, and different types of tumors correspond to different constant values. 5. The tumor ablation dose control method according to claim 1 is characterized in that the step of setting the initial parameters of the ablation device according to the initial ablation dose includes: calculating the initial power R value t according to the energy conversion efficiency of the ablation device through the formula Pinit = 4, where AT is the preset initial ablation time; setting the initial coolant flow rate according to the heat dissipation characteristics of the ablation needle, and the flow rate and power satisfy the relationship of CX = CX, where c is the heat dissipation coefficient. 6.The tumor ablation dose control method according to claim 1 is characterized in that in the step of using force to turn on the CT scanning device to update data at each specific time interval, the value of force is related to the growth rate of the tumor, and the faster the growth rate, the smaller the value of force. 7The method for controlling the dose of tumor ablation according to claim 1 is characterized in that after the step of reducing or increasing the power of the ablation device, it further comprises: continuing to monitor the real-time ablation dose reoi, and if the monitoring results show that reoi still exceeds or falls below the corresponding dose range after N consecutive times, an alarm is issued to prompt the operator, and N is the preset alarm threshold number. 8.A CT-guided tumor ablation dose control system, characterized in that it includes the following modules: an imaging and parameter acquisition module, which turns on a CT scanning device to perform three-dimensional imaging of the tumor site, obtains three-dimensional contour data of the tumor, calculates the tumor volume based on the contour data, and calculates the average density of the tumor through the grayscale value distribution of the CT image; an initial dose calculation module, which calculates the initial ablation dose Dmit using a preset formula based on the tumor volume and average density; the formula for calculating the initial ablation dose Dmit is: ^init - k XVX ρ·, where V is the tumor volume, 0 is the average density of the tumor, Α; is an empirical coefficient, and the value of Α; is related to the tissue type of the tumor, and different types of tumors correspond to different constant values; an ablation device control module, which sets the initial parameters of the ablation device according to the initial ablation dose Dinit and starts the ablation program; a real-time monitoring module, which turns on the CT scanning device again at a specific interval to update the three-dimensional contour data and grayscale value distribution of the tumor, and repeats the above volume, density and dose calculation steps to obtain the real-time ablation dose OreaZ; a dose comparison and control module, which converts the real-time ablation dose Dreai Compared with the initial ablation dose Dmit, when Dreal > Dinit X (1 + .), reduce the power of the ablation device, where & is the preset upper dose floating ratio; when Dreai < Dinit When X(1-(3)), the power of the ablation device is increased, and X(1-(3)) is the floating ratio of the preset lower limit of the dose. 9. The tumor ablation dose control system according to claim 8 is characterized in that the ablation device control module includes a power setting submodule, which calculates the initial power Pinit according to the energy conversion efficiency of the ablation device through the formula Pinit = ΔT X (7, where AT is the preset initial ablation time; according to the heat dissipation characteristics of the ablation needle, sets the initial coolant flow rate Gt, the flow rate and power satisfy the relationship of Gt = c X t, where c is the heat dissipation coefficient. 10.The tumor ablation dose control system according to claim 8 is characterized in that the imaging and parameter acquisition module: calculates the tumor volume V according to the contour data through the formula V = Si × hi, where Si is the cross-sectional area of the i-th layer, % is the spacing between two adjacent cross-sectional layers, and η is the number of layers; calculates the average tumor density O through the gray value distribution of the CT image in combination with the formula o = £Taozhu, where spoon is the density value corresponding to the j-th fm pixel point, and m is the total number of selected pixel points; the method for selecting pixel points is: draw a circle with the geometric center of the tumor as the center and r as the radius, select the pixel points within the circle to participate in the calculation, and r is determined according to the maximum diameter R of the tumor through the formula r = and. | A CT-guided tumor-ablation dose-control method that calculates tumor volume and density from CT contour and gray-value data to set the initial ablation dose. The system updates parameters at intervals during ablation based on new CT data. | R | "starting a CT scanning device to perform three-dimensional imaging of the tumor site, obtaining three-dimensional contour data of the tumor, and calculating the tumor volume based on the contour data; calculating the average density of the tumor through the gray value distribution of the CT image" (Claim 1) "starting the CT scanning device again at a specific interval of time… to update the three-dimensional contour data and gray value distribution of the tumor, repeating the above volume, density and dose calculation steps to obtain the real-time ablation dose Dreal" (Claim 1) "comparing the real-time ablation dose Dreal with the initial ablation dose Dinit, and when Dreal > Dinit × (1 + δ), reducing the power of the ablation device" (Claim 1) "the formula for calculating the initial ablation dose Dinit is: Dinit = k × V × ρ; where k is the empirical coefficient, the value of k is related to the tissue type of the tumor" (Claim 4) | [8 empty] | 1 | [11 empty] | 1 | [7 empty] | 1 |  |  | THE SIXTH AFFILIATED HOSPITAL SUN YAT SEN UNIV | THE SIXTH AFFILIATED HOSPITAL SUN YAT SEN UNIV | 2025-02-21 | 2025 | 2025-03-28 | 2025 | 2025-02-21 | 2025 | CN202510196901 | [NA]
408 | CN120053073A | Visual localization and ablation method and system for cardiovascular balloon | The invention discloses a visual localization and ablation method and system for a cardiovascular balloon. Acquiring a real-time image sequence through a synchronous multi-view image acquisition technology; through a multi-view image reconstruction algorithm, a three-dimensional blood vessel model is obtained through conversion; in the three-dimensional blood vessel model, identifying an opening position and morphological characteristics of a target blood vessel, and extracting to obtain a blood vessel lesion area for positioning an ablation target; the optimal positioning point coordinates of the cardiovascular balloon are obtained through calculation in combination with hemodynamic simulation, and the optimal path is planned; adjusting the position of the cardiovascular balloon using a navigation algorithm; real-time parameters in the ablation process are collected, and the ablation parameters are monitored and adjusted and controlled in real time. The system comprises an image acquisition module, a two-dimensional extraction module, a three-dimensional modeling module, a vasculopathy area extraction module, an optimal path planning module, a navigation adjustment module and a parameter adjustment module. Modern imaging, hemodynamic simulation and electromagnetic positioning are combined, and dynamic control over the ablation process is achieved. | 1. L A visual positioning and ablation method for a cardiovascular balloon, characterized in that it includes: using a multi-channel imaging system to obtain a real-time image sequence through synchronous multi-view image acquisition technology; using an edge detection algorithm to extract the main contours of the blood vessel from the real-time image sequence to obtain a two-dimensional blood vessel image sequence; converting the two-dimensional blood vessel image sequence into a three-dimensional blood vessel model through a multi-view image reconstruction algorithm; in the three-dimensional blood vessel model, based on a deep learning classification network and an image generation adversarial network technology, identifying the opening position and morphological characteristics of the target blood vessel, and extracting the vascular lesion area used to locate the ablation target; combining hemodynamic simulation, based on the three-dimensional blood vessel model and the vascular lesion area, calculating the coordinates of the optimal positioning point of the cardiovascular balloon, and planning the optimal path of the cardiovascular balloon from the current position to the coordinates of the optimal positioning point; obtaining the real-time position information of the cardiovascular balloon through electromagnetic positioning, calculating the relative spatial relationship between the cardiovascular balloon and the optimal path, and adjusting the position of the cardiovascular balloon using a navigation algorithm; collecting real-time parameters during the ablation process, monitoring the ablation parameters and adjusting and controlling them in real time. 2.The visual positioning and ablation method for cardiovascular balloons according to claim 1 is characterized in that the multi-channel imaging system is used to obtain a real-time image sequence through synchronous multi-view image acquisition technology, including: obtaining cross-sectional images of blood vessels and surrounding tissues through a CT imaging channel; obtaining images of the thickness and blood flow characteristics of the blood vessel wall through an MRI imaging channel; obtaining direct-view images of the interior of the cardiovascular system in real time through a flexible endoscope imaging channel; setting the acquisition frequency and image resolution of each imaging channel, synchronously recording the images acquired by each imaging channel according to timestamps through an image sensor, aligning multi-channel data, and obtaining a real-time image sequence. 3The visual positioning and ablation method for cardiovascular balloon according to claim 1 is characterized in that the step of extracting the main contour of the blood vessel from the real-time image sequence using an edge detection algorithm to obtain a two-dimensional blood vessel image sequence comprises: performing two-dimensional directional gradient calculation on the images in the real-time image sequence based on a Sobel edge detection algorithm, LL The horizontal gradient of each pixel in the X direction is Gx(x,y)= +, α=-10=-1 1I The vertical gradient of each pixel in the γ direction Gj,(X,D) = £ + a,y + hy Ky, where Kχ (a, b) is the Sobel convolution kernel in the X direction, Ky (a, b) is the Sobel convolution kernel in the Y direction, the Sobel convolution kernel is designed according to the thickness of the blood vessel, and I (x+a,y+b) is the pixel value of the current position in the image; Calculate the gradient amplitude of each pixel point G(x,y) = jG;(x,y) + G;(Ky), which is used to indicate the change intensity of the pixel point; Calculate the gradient direction of each pixel point 6(x,y)=atan2(Gy(x,y),G* (x,y)), which is used to indicate the direction of the edge; According to the edge strength and morphological characteristics of the blood vessel, set the high threshold T and the low threshold £; The pixel point with a gradient amplitude greater than the high threshold is a valid edge and the pixel point is retained; For the pixel point with a gradient amplitude less than the high threshold but greater than the low threshold, determine whether the pixel point is connected to a strong edge. If connected, the pixel point is considered to be a valid edge and retained, if not connected, the pixel point is not a valid edge and is removed The invention further comprises the following steps: removing pixels whose gradient amplitude is less than the low threshold value and not belonging to the edge, extracting the retained pixels and obtaining a binary image, extracting the contour of the blood vessel from the binary image using a contour extraction algorithm and obtaining a two-dimensional blood vessel image sequence containing the main contour of the blood vessel. 4. The visual positioning and ablation method for cardiovascular balloon according to claim 1 is characterized in that the converting the two-dimensional vascular image sequence into a three-dimensional vascular model through a multi-view image reconstruction algorithm comprises: extracting feature points of the blood vessels at multiple viewpoints from the two-dimensional vascular image sequence to represent the key feature area of the blood vessels; applying a beam normal vector optimization algorithm according to the feature points extracted from each viewpoint, projecting the two-dimensional feature points into a three-dimensional space, and optimizing to obtain a three-dimensional point cloud coordinate symbol ψ={ear,£,...,dagger}, wherein the three-dimensional point cloud coordinates conform to the projection in all viewpoints nmnm n n, and the optimization goal is to minimize the error function ψ=ΣΣ&,·=ΣΣ¦¦α-λ·^¦¦, wherein η is the total number of feature points, m is the total number of viewpoints, ej=¦=(x"x'i)2+(y"y'i)2 is the projection error of each feature point, Pi=(Xi,y) is the two-dimensional coordinate of the feature point projected by the blood vessel in the jth viewpoint, Pj is the camera projection matrix representing the mapping relationship from the three-dimensional space to the two-dimensional image plane, Pi = (Xi, Yi, Z) is the three-dimensional coordinate point corresponding to each two-dimensional feature point, P': = &':,y') is the projection position of the three-dimensional point Pj at the jth viewing angle; the three-dimensional point cloud coordinates of each viewing angle are used to generate a three-dimensional point cloud model representing the spatial distribution of the blood vessel; the three-dimensional point cloud models of each viewing angle are merged and surface reconstructed to generate a three-dimensional surface model of the blood vessel. 5.The visual positioning and ablation method for cardiovascular balloon according to claim 1 is characterized in that, in the three-dimensional blood vessel model, based on the deep learning classification network and the image generation adversarial network technology, the opening position and morphological features of the target blood vessel are identified, and the vascular lesion area used to locate the ablation target is extracted, which includes: voxelizing the three-dimensional blood vessel model to obtain three-dimensional image data; designing a deep learning classification network model, the three-dimensional image data is used as the input of the deep learning classification network model, and is used to represent the grayscale or intensity information of the blood vessel in three-dimensional space; using the labeled data set to train the deep learning classification network model, and identifying the blood vessel opening and morphological features by optimizing the objective function; the output of the deep learning classification network model is the classification probability of each voxel or blood vessel area, indicating whether the point belongs to the target blood vessel opening area; extracting blood vessel morphological features from the intermediate feature map and label map of the deep learning classification network model, the blood vessel morphological features include at least one of the blood vessel branch angle, blood vessel diameter, and curvature; designing a generator and a discriminator of the image generation adversarial network model, the generator receives the three-dimensional image data,Generate a vascular lesion area, the discriminator distinguishes the area generated by the generator from the real vascular lesion area, and outputs the vascular morphology and the spatial distribution of the lesion area; calculate the position of the lesion area in the three-dimensional space based on the extracted lesion area and the spatial distribution, and obtain the vascular lesion area for locating the ablation target in the three-dimensional vascular model. 6The visual positioning and ablation method for cardiovascular balloon according to claim 1 is characterized in that the combination of hemodynamic simulation, based on the three-dimensional vascular model and the vascular lesion area, calculates the cardiovascular balloon Optimal positioning point coordinates, planning the best path of the cardiovascular balloon from the current position to the optimal positioning point coordinates includes: establishing a hemodynamic simulation equation xi ©7 dt + ν·Γν =-"Death+bad2Ya+ 17, where ρ is the density of blood, V is the velocity field, e is the pressure, μ is the viscosity of blood, F is the external force, and P is the gradient operator; selecting a candidate positioning point R according to the three-dimensional vascular model and the vascular lesion area, the candidate positioning point R is located at the front end of the vascular lesion area or close to the vascular lesion area; calculating the optimal positioning point coordinates of the cardiovascular balloon &must=argmin, where g (IQ{R)\v R }), the optimal positioning point is the point where the balloon has the least effect on blood flow when passing through the blood vessel, Q(R) is the flow rate of the candidate positioning point, and Fe is the pressure gradient of the candidate positioning point; a blood flow disturbance model during the movement of the balloon is established, and the blood flow disturbance yr is calculated for any path point, wherein yr is the blood flow velocity at the path point 4, yr is the velocity of the balloon, and yritis the pressure gradient at the path point 4; an optimization function η is established with the goal of minimizing the blood flow disturbance 3Qu and optimizing the path length Distance (r5, q) Lputh=Σ·39 +"·Distancc(7^1；/J), where is the total path cost, indicating the total cost of the balloon from the current position u=L to the optimal positioning point, and W] is the distance weight coefficient; Solve the optimization function, calculate the total cost through the path selection algorithm, and the cost update formula is c(rj =min(C (rv) ,C (ru) +6Qu+w2 Distance (4,q)), where q is the adjacent path node of the path point 4, C (rv) is the cost of node q, Distance (ru,rv) is the geometric distance between node q and the path point 4, indicating the physical distance of the balloon from one node to another, W2 is the distance weight coefficient, and the path with the least cost is selected as the final path to obtain the best path of the cardiovascular balloon from the current position to the coordinates of the optimal positioning point. 7The visual positioning and ablation method for cardiovascular balloon according to claim 6 is characterized in that after obtaining the optimal path, the optimal path is smoothed to obtain the smoothed optimal path. r(0 - , where is the B-spline basis function and 4 is the control point of path point 4. M sentence 8.The visual positioning and ablation method for a cardiovascular balloon according to claim 1 is characterized in that the real-time position information of the cardiovascular balloon is obtained by electromagnetic positioning, the relative spatial relationship between the cardiovascular balloon and the optimal path is calculated, and the position of the cardiovascular balloon is adjusted using a navigation algorithm, including: embedding a micro sensor in the balloon catheter, detecting electromagnetic field signals and position information in real time, and sending them outward; receiving the electromagnetic field signal and the position information, and calculating the spatial coordinates Psens°r = (X, γ, ζ) and the direction angle e of the cardiovascular balloon; ' sensor ' sensor ' sensor'1χ 1 7 / τ'·sensor ' Calculate the position error Δ P (t) and direction error Δ Θ; Calculate the adjustment amount of the direction and movement speed of the cardiovascular balloon, and use the proportional-integral-differential control algorithm to obtain an adjustment control signal for controlling the movement of the balloon, wherein the formula of the proportional-integral-differential control algorithm is w(0 = Κρ'ΚΡ© + &· + , u (t) is the control signal, Kp is the proportional coefficient, valley is the differential coefficient, 0 is the integral gain coefficient, δρ(ό is the position error, acorpse«) is the rate of change of the position error. 9.The visual positioning and ablation method for cardiovascular balloons according to claim 1 is characterized in that the real-time parameters collected during the ablation process, the monitoring and real-time adjustment and control of the ablation parameters include: collecting the temperature, pressure and energy transmission conditions during the ablation process; real-time monitoring of whether the temperature reaches the preset ablation range, whether the balloon pressure exceeds the normal range, and whether the energy transmission rate exceeds the normal range; adjusting the control signal based on the preset threshold through the real-time monitoring of various parameters to control the ablation strategy. 10. A visual positioning and ablation system for cardiovascular balloons, characterized in that it comprises: an image acquisition module, used to adopt a multi-channel imaging system to acquire a real-time image sequence through a synchronous multi-view image acquisition technology; a two-dimensional extraction module, used to extract the main contours of the blood vessel from the real-time image sequence using an edge detection algorithm to obtain a two-dimensional blood vessel image sequence; a three-dimensional modeling module, used to convert the two-dimensional blood vessel image sequence into a three-dimensional blood vessel model through a multi-view image reconstruction algorithm; a vascular lesion area extraction module, used to identify the opening position and morphological characteristics of the target blood vessel in the three-dimensional blood vessel model based on a deep learning classification network and an image generation adversarial network technology, and extract the vascular lesion area for locating the ablation target; an optimal path planning module, used to calculate the optimal positioning point coordinates of the cardiovascular balloon based on the three-dimensional blood vessel model and the vascular lesion area in combination with hemodynamic simulation, and plan the optimal path of the cardiovascular balloon from the current position to the optimal positioning point coordinates; a navigation adjustment module, used to obtain the real-time position information of the balloon catheter through electromagnetic positioning, calculate the relative spatial relationship between the cardiovascular balloon and the optimal path, and adjust the position of the cardiovascular balloon using a navigation algorithm; The parameter adjustment module is used to collect real-time parameters during the ablation process, monitor the ablation parameters and adjust and control them in real time. | A cardiovascular-balloon ablation system that uses synchronous multi-view image acquisition and a 3D reconstruction algorithm to build a vessel model, then identifies the target opening and balloon position. The 3D view drives precise balloon positioning and ablation. | R | Imaging | [7 empty] | 1 | 1 | [12 empty] | 1 |  |  | 1 | [6 empty] | TIANJIN INTELLIGENT HEALTH MEDICAL TECH CO LTD¦TIANJIN YUANSHAN MEDICAL TECH CO LTD | TIANJIN INTELLIGENT HEALTH MEDICAL TECH CO LTD¦TIANJIN YUANSHAN MEDICAL TECH CO LTD | 2025-02-10 | 2025 | 2025-05-30 | 2025 | 2025-02-10 | 2025 | CN202510145311 | [NA]
409 | WO2025263780A1 | HIGH-FREQUENCY THERMAL COAGULATION CONTROL DEVICE AND SYSTEM USING THREE-DIMENSIONAL BRAINWAVE ELECTRODE BASED ON REAL-TIME TEMPERATURE MONITORING | The present invention may be characterized by: communicating with a plurality of first sensors which are provided on one side of the inside of a rod-type electrode member in a high-frequency thermal coagulation device and detect brain waves of a high-frequency irradiation site in the brain, the rod-type electrode member being inserted into the high-frequency irradiation site and including a plurality of electrodes; receiving brain wave information acquired through the plurality of first sensors; communicating with a plurality of second sensors which are provided on the other side of the inside of the rod-type electrode member in the high-frequency thermal coagulation device and detect the temperature of the high-frequency irradiation site; receiving temperature information acquired through the plurality of second sensors; and when the plurality of electrodes reach the high-frequency irradiation site, controlling an energy supply unit so that a current is applied to the plurality of electrodes at a preset current intensity in connection with the brain wave information and the temperature information. | 1. A communication unit for performing communication with a high-frequency thermal coagulation device; and a processor for performing operations related to controlling the high-frequency thermal coagulation.Including, the communication unit, which is inserted into a high-frequency irradiation site of the brain among the high-frequency thermal coagulation devices and is provided on one side of the inside of a rod-type electrode member including a plurality of electrodes, performs communication with a plurality of first sensors that detect brain waves of the high-frequency irradiation site, and which is provided on the other side of the inside of the rod-type electrode member among the high-frequency thermal coagulation devices and detects the temperature of the high-frequency irradiation site, and performs communication with an energy supply unit that selectively applies current at a preset current intensity to the plurality of electrodes among the high-frequency thermal coagulation devices, and the processor receives brain wave information acquired through the plurality of first sensors, receives temperature information acquired through the plurality of second sensors, and controls the energy supply unit so that when the plurality of electrodes reach the high-frequency irradiation site, the current is applied to the plurality of electrodes at a preset current intensity in connection with the brain wave information and the temperature information.High frequency thermal coagulation control device. 2. In the first paragraph, the high-frequency thermal coagulation control device, characterized in that the processor receives the brain wave information through the plurality of first sensors provided corresponding to the plurality of electrodes on one inner side of the electrode member. 3. In the second paragraph, the high-frequency thermal coagulation control device, characterized in that the processor receives the brain wave information through the plurality of first sensors respectively provided corresponding to the plurality of electrodes at preset equal intervals. 4. In the second paragraph, the high-frequency thermal coagulation control device, characterized in that the processor receives the brain wave information through the plurality of first sensors respectively provided corresponding to the plurality of electrodes at different preset intervals. 5. In the second paragraph, the high-frequency thermal coagulation control device, characterized in that the processor receives the temperature information through the plurality of second sensors provided corresponding to the plurality of electrodes among the inner other sides of the electrode member. 6. In the fifth paragraph, the high-frequency thermal coagulation control device, characterized in that the processor receives the temperature information through the plurality of second sensors respectively provided corresponding to the plurality of electrodes at preset equal intervals. 7. In the fifth paragraph, the high-frequency thermal coagulation control device, characterized in that the processor receives the temperature information through the plurality of second sensors respectively provided corresponding to the plurality of electrodes at different preset intervals. 8. In the first paragraph, the processor controls the energy supply unit so that, when the plurality of electrodes reach the high-frequency irradiation site, a first current of the monopolar type and a second current of the bipolar type are selectively applied to the plurality of electrodes at a preset current intensity linked to the brain wave information and the temperature information, a high-frequency thermal coagulation control device. 9. In the first paragraph, the processor controls the energy supply unit so that, when the plurality of electrodes reach the high-frequency irradiation site, a first current of the monopolar type and a second current of the bipolar type are alternately applied to the plurality of electrodes at a preset current intensity linked to the brain wave information and the temperature information, a high-frequency thermal coagulation control device. 10. In the first paragraph, the processor controls the energy supply unit so that, when the plurality of electrodes reach the high-frequency irradiation site, a first current of the monopolar type and a second current of the bipolar type are sequentially applied to the plurality of electrodes at a preset current intensity linked to the brain wave information and the temperature information, a high-frequency thermal coagulation control device. 11. A method for controlling high-frequency thermal coagulation of a high-frequency thermal coagulation control device, comprising: a step of performing communication with a plurality of first sensors that detect brain waves of the high-frequency irradiation area, which are inserted into a high-frequency irradiation area of the brain among high-frequency thermal coagulation devices and are provided on one side of the inside of a rod-type electrode member including a plurality of electrodes; A step of receiving brain wave information acquired through the plurality of first sensors; A step of performing communication with a plurality of second sensors that are provided on the inner side of the rod-type electrode member of the high-frequency thermal coagulation device and detect the temperature of the high-frequency irradiation area; A method comprising: a step of receiving temperature information acquired through the plurality of second sensors; and a step of controlling the energy supply unit so that, when the plurality of electrodes reach the high-frequency irradiation site, the current is applied to the plurality of electrodes at a preset current intensity linked to the brain wave information and the temperature information. 12. In the 11th paragraph, the control step is characterized in that, when the plurality of electrodes reach the high-frequency irradiation site, the energy supply unit is controlled so that a first current of the monopolar type and a second current of the bipolar type are selectively applied to the plurality of electrodes at a preset current intensity linked to the brain wave information and the temperature information. 13. In the 11th paragraph, the control step is characterized in that, when the plurality of electrodes reach the high-frequency irradiation site, the energy supply unit is controlled so that a first current of the monopolar type and a second current of the bipolar type are alternately applied to the plurality of electrodes at a preset current intensity linked to the brain wave information and the temperature information. 14. In the 11th paragraph, the control step is characterized in that, when the plurality of electrodes reach the high-frequency irradiation site, the energy supply unit is controlled so that a first current of the monopolar type and a second current of the bipolar type are sequentially applied to the plurality of electrodes at a preset current intensity linked to the brain wave information and the temperature information. 15. A high-frequency thermal coagulation control system comprising a high-frequency thermal coagulation device and a high-frequency thermal coagulation control device for controlling the high-frequency thermal coagulation device, wherein the high-frequency thermal coagulation device comprises: a rod-type electrode member inserted into a high-frequency irradiation site of the brain and including a plurality of electrodes; A plurality of first sensors provided on one inner side of the electrode member of the above-mentioned rod type and detecting brain waves of the high-frequency irradiation area; A system comprising: a plurality of second sensors provided on the inner side of the rod-type electrode member and detecting the temperature of the high-frequency irradiation area; and an energy supply unit selectively applying current to the plurality of electrodes; wherein the high-frequency thermal coagulation control device receives brain wave information acquired through the plurality of first sensors, receives temperature information acquired through the plurality of second sensors, and controls the energy supply unit so that, when the plurality of electrodes reach the high-frequency irradiation area, the current is applied to the plurality of electrodes at a preset current intensity linked to the brain wave information and the temperature information. | A high-frequency thermal coagulation system using a rod electrode with built-in sensors that detect brain waves at the irradiation site. Real-time brainwave info during thermocoagulation guides parameters and monitors safety. | R | "a rod-type electrode member inserted into a high-frequency irradiation site of the brain and including a plurality of electrodes; A plurality of first sensors provided on one inner side of the electrode member of the above-mentioned rod type and detecting brain waves of the high-frequency irradiation area; A plurality of second sensors provided on the inner side of the rod-type electrode member and detecting the temperature of the high-frequency irradiation area" (Claim 15) "the processor receives brain wave information acquired through the plurality of first sensors, receives temperature information acquired through the plurality of second sensors, and controls the energy supply unit so that, when the plurality of electrodes reach the high-frequency irradiation site, the current is applied to the plurality of electrodes at a preset current intensity in connection with the brain wave information and the temperature information" (Claim 1) "a first current of the monopolar type and a second current of the bipolar type are alternately applied to the plurality of electrodes at a preset current intensity linked to the brain wave information and the temperature information" (Claim 9) | 1 |  |  |  | 1 | [24 empty] | 1 |  | UIF UNIV INDUSTRY FOUNDATION YONSEI UNIV | UIF UNIV INDUSTRY FOUNDATION YONSEI UNIV | 2025-04-14 | 2025 | 2025-12-26 | 2025 | 2024-06-18 | 2024 | WO2025KR05022 | KR20250178162A
410 | US20250311929A1 | SYSTEM AND METHOD OF DUAL-MODE EAT/US-GUIDED ELECTROPORATION | The present embodiments relate generally to medical care and more particularly to a new imaging modality referred to herein as electroacoustic tomography (EAT) that not only depicts electrical field energy distribution in real time, but also enables clear discrimination between IRE and RE zones in situ during the treatment. According to some aspects, EAT exploits the phenomenon that the amplitude of acoustic emission generated by an electric field is proportional to the electrical energy deposition in tissue. After detecting these acoustic waves with ultrasound transducers, an image of the electric field distribution can be reconstructed in real-time. This will allow real time monitoring the IRE and provides feedback control of the treatment. Some embodiments are directed to a novel EAT combined with ultrasound (EAT/US) guided IRE intervention that permits a) real time ultrasound image-guided needle placement, and b) intraoperational discrimination of IRE and RE zones during the treatment of cancer. | 1. A system for electroacoustic tomography (EAT) including: a device adapted to receive ultrasound signals reflected from transmitted ultrasound waves; a device adapted to receive electro-acoustic signals generated from an electrical pulse generation device; an EAT imaging unit adapted to generate an electroacoustic image based on the received electro-acoustic signals; an ultrasound imaging unit adapted to generate an ultrasound image based on the received ultrasound signals; and an image combining device adapted to receive and combine the generated electroacoustic image and ultrasound image. 2. The system of claim 1, further comprising: a device adapted to generate an electrical field induced acoustic signal; and a device adapted to transmit ultrasound waves. 3. The system of claim 1, further comprising a trigger sequencing logic device for switching a connection of an ultrasound transducer between ultrasound transmission/reception and electro-acoustic signal acquisition. 4. The system of claim 3, wherein the trigger sequencing logic device is configured to perform switching such that the ultrasound signal and the electroacoustic signal are received in different time windows. 5. The system of claim 1, wherein the ultrasound imaging unit includes: a transducer adapted to receive ultrasound signals in an ultrasound reception window controlled by a flip-flop sequencing logic switch; an ultrasonic beamformer adapted to produce a series of radio frequency signals from the ultrasonic signals; and an ultrasound image construction unit adapted to produce the ultrasound image from the radio frequency signals. 6. The system of claim 1, wherein the EAT imaging unit includes: an acquisition stage adapted to collect an electroacoustic signal in response to an electrical pulse signal trigger; a multi-channel pre-amplification stage adapted to amplify the electroacoustic signal; an averaging/filtering stage adapted to process the amplified electroacoustic signal; and an EAT image construction unit adapted to perform electro-acoustic imaging reconstruction. 7. The system of claim 1, wherein the image combiner uses spatially resampling to produce averaged electroacoustic imaging reconstructions so that the electroacoustic image can be superimposed on top of the ultrasound image. 8. The system of claim 7, wherein the image combiner produces combined image frames for display. 9. A method for electroacoustic tomography (EAT) including: depicting electrical field energy distribution in real time; enabling clear discrimination between irreversible electroporation (IRE) zones and reversible electroporation (RE) zones in situ during treatment, including exploiting the phenomenon that the amplitude of acoustic waves generated by an electric field is proportional to the electrical energy deposition in tissue; and after detecting the acoustic waves with ultrasound transducers, constructing an image of the electric field distribution in real-time. 10. The method of claim 9, further comprising enabling real time ultrasound image-guided needle placement. 11. A method for electroacoustic tomography (EAT) including: receiving ultrasound signals reflected from transmitted ultrasound waves; receiving electro-acoustic signals generated from an electrical pulse generation device; generating, by an EAT imaging unit, an electroacoustic image based on the received electro-acoustic signals; generating, by an ultrasound imaging unit, an ultrasound image based on the received ultrasound signals; and receiving, by an image combining device, the generated electroacoustic image and ultrasound image and combining them to generate a combined image. 12. The method of claim 11, further comprising: generating an electrical field induced acoustic signal; and transmitting ultrasound waves. 13. The method of claim 11, further comprising switching a connection of an ultrasound transducer between ultrasound transmission/reception and electro-acoustic signal acquisition. 14. The method of claim 13, wherein switching includes performing the switching such that the ultrasound signal and the electroacoustic signal are received in different time windows. 15. The method of claim 11, wherein generating the ultrasound image includes: receiving ultrasound signals in an ultrasound reception window controlled by a flip-flop sequencing logic switch; producing a series of radio frequency signals from the ultrasonic signals; and producing the ultrasound image from the radio frequency signals. 16. The method of claim 11, wherein generating the electroacoustic image includes: collecting an electroacoustic signal in response to an electrical pulse signal trigger; amplifying the electroacoustic signal; processing the amplified electroacoustic signal in an averaging/filtering stage; and performing electro-acoustic imaging reconstruction based on the processed electroacoustic signal. 17. The method of claim 11, wherein combining includes using spatially resampling to produce averaged electroacoustic imaging reconstructions so that the electroacoustic image can be superimposed on top of the ultrasound image. 18. The method of claim 17, further including producing combined image frames for display. | A dual-mode electroacoustic-tomography (EAT)/ultrasound electroporation system: EAT shows electric-field energy distribution in real time and discriminates IRE versus reversible electroporation zones. The acoustic-emission amplitude scales with deposited electrical energy. | R | "depicting electrical field energy distribution in real time; enabling clear discrimination between irreversible electroporation (IRE) zones and reversible electroporation (RE) zones in situ during treatment, including exploiting the phenomenon that the amplitude of acoustic waves generated by an electric field is proportional to the electrical energy deposition in tissue; and after detecting the acoustic waves with ultrasound transducers, constructing an image of the electric field distribution in real-time" (Claim 9) "a trigger sequencing logic device for switching a connection of an ultrasound transducer between ultrasound transmission/reception and electro-acoustic signal acquisition" (Claim 3) "the image combiner uses spatially resampling to produce averaged electroacoustic imaging reconstructions so that the electroacoustic image can be superimposed on top of the ultrasound image" (Claim 7) |  |  | 1 | 1 | [7 empty] | 1 | [10 empty] | 1 | [5 empty] | 1 |  |  | UNIV CALIFORNIA | UNIV CALIFORNIA | 2023-05-30 | 2023 | 2025-10-09 | 2025 | 2022-05-31 | 2022 | US202318869803 | WO2023235332A1
411 | US12072283B2 | System, method, computer-accessible and apparatus for providing near-infrared spectroscopy for anatomical mapping of the epicardium | An exemplary catheter can be provided, which can include, for example a source fiber(s) configured to (i) receive a near infrared spectroscopic (NIRS) radiation, and (ii) provide the NIRS radiation to a portion(s) of a sample(s), a detection fiber(s) configured to receive a return radiation from the sample(s) that can be based on the NIRS radiation that was provided to the portion(s) of the sample(s), and an ablation electrode(s) configured to ablate the sample(s) based on the return radiation. The source fiber(s), the detection fiber(s), and the ablation electrode(s) can be integrated into the single sheath. The ablation electrode(s) can be a radiofrequency ablation electrode. | 1. A catheter comprising: at least one source fiber configured to (i) receive a near infrared spectroscopic (NIRS) radiation, and (ii) provide the NIRS radiation to at least one portion of at least one sample; at least one detection fiber configured to receive a return radiation from the at least one sample that is based on the NIRS radiation that was provided to the at least one portion of the at least one sample; and at least one ablation electrode configured to ablate the at least one sample based on the return radiation. 2. The catheter of claim 1, further comprising a single sheath, wherein the at least one source fiber, the at least one detection fiber, and the at least one ablation electrode are integrated into the single sheath within the catheter. 3. The catheter of claim 1, wherein the at least one ablation electrode is a radiofrequency ablation electrode. 4. The catheter of claim 1, wherein a configuration of the at least one source fiber and the at least one detection fiber are spaced apart at a specific distance from one another. 5. The catheter of claim 4, wherein the specific distance is between about 0 mm to about 4 mm. 6. The catheter of claim 1, further comprising at least one broadband light source configured to generate the NIRS radiation being received by the at least one source fiber. 7. The catheter of claim 1, further comprising at least one continuous-wave light source configured to generate the NIRS radiation being received by the at least one source fiber. 8. The catheter of claim 1, further comprising a charge coupled device configured to (i) receive the return radiation from the at least one detection fiber, and (ii) generate imaging information based on the return radiation, wherein the imaging information is associated with the at least one portion of the at least one sample. 9. The catheter of claim 1, wherein: the at least one source fiber includes at least six source fibers, the at least one detection fiber includes at least six detection fibers, and the at least one ablation electrode includes at least twelve ablation electrodes, and wherein a combination of the at least six source fibers and the at least six detection fibers facilitates a parallel or simultaneous detection of information regarding the at least one portion of the at least one sample. 10. The catheter of claim 9, wherein: one of the at least six source fibers and one of the at least six detection fibers form a pair of fibers, one of the at least twelve ablation electrodes is disposed on a first side of the pair of fibers, another one of the at least twelve ablation electrodes is disposed on a second side of the pair of fibers, and the first side is disposed opposite to the second side with respect to the pair of fibers. 11. A method for ablating at least one portion of at least one sample, comprising: generating at least one near infrared spectroscopic (NIRS) radiation; providing the at least one NIRS radiation to the at least one portion using at least one source fiber provided in a catheter; receiving at least one return radiation from the at least one portion that is based on the at least one NIRS radiation that was provided to the at least one portion of the at least one sample using at least one detection fiber provided in the catheter; and ablating the at least one portion using at least one ablation electrode provided in the catheter based on the at least one return radiation. 12. The method of claim 11, wherein the at least one ablation electrode includes at least one radiofrequency ablation (RFA) electrode integrated into the catheter. 13. The method of claim 11, further comprising determining a three-dimensional position of the catheter with respect to the at least one portion. 14. The method of claim 11, further comprising determining a contact between the catheter and the at least one portion. 15. The method of claim 11, wherein the at least one sample includes a heart, and further comprising generating at least one epicardial map of the heart based on the at least one return radiation. 16. The method of claim 15, wherein the at least one epicardial map includes at least one of (i) at least one coronary vessel in the heart, (ii) fat in the heart, (iii) scar tissue in the heart, or (iv) fibrosis in the heart. 17. The method of claim 16, further comprising determining a location of the at least one coronary vessel in the heart by determining a hemoglobin concentration in the heart. 18. The method of claim 17, further comprising determining the hemoglobin concentration using a spectral unmixing procedure. 19. The method of claim 18, wherein the spectral unmixing procedure is an inverse Monte Carlo procedure. 20. The method of claim 11, further comprising generating at least one optical substrate map of at least one portion of the at least one sample based on the at least one return radiation. 21. The method of claim 20, further comprising automatically assessing at least one substrate using the at least one optical substrate map, wherein the at least one substrate includes a pulmonary vein, a coronary artery, fat, fibrosis or a scar. 22. The method of claim 20, further comprising: tracking a position of the catheter at or within the tissue; and interpolating the at least one optical substrate map based on the at least one return radiation as a function of the tracked position. 23. The method of claim 20, wherein the generating of the at least one optical substrate map includes tracking at least one input. 24. The method of claim 11, wherein the at least one source fiber includes a plurality of source fibers, wherein the at least one detection fiber includes a plurality of detection fibers, and wherein the receiving of at least one return radiation includes a receipt of plurality of radiations, each from a respective one of the plurality of detection fibers, facilitating a parallel or simultaneous detection of information regarding the at least one portion of the at least one sample. 25. The method of claim 24, wherein the information is simultaneously detected from different locations on or in the at least one portion of the at least one sample. 26. The method of claim 11, wherein the source and detection fibers extend in the catheter for a majority of a length thereof. 27. A non-transitory computer-accessible medium having stored thereon computer-executable instructions for causing an ablation of at least one portion of at least one sample, wherein, when a computing arrangement executes the instructions, the computing arrangement is configured to perform procedures comprising: receiving information related to a backscattered radiation from the at least one portion that is based on at least one near infrared spectroscopic radiation provided to the at least one sample by at least one source fiber provided in a catheter, wherein the received information is based on at least one return radiation at least one detection fiber provided in the catheter; generating at least one epicardial map of the at least one portion based on the information; determining a first location of at least one coronary vessel using the information; and causing the ablation of the at least one portion by at least one ablation electrode provided in the catheter at a second location that excludes the at least one coronary vessel. | A NIRS catheter for epicardial mapping with source fibers delivering near-infrared light, detection fibers reading return light, and an ablation electrode that uses that NIRS info. The catheter ablates the same area it characterizes optically. | R | "at least one source fiber configured to (i) receive a near infrared spectroscopic (NIRS) radiation, and (ii) provide the NIRS radiation to at least one portion of at least one sample… at least one ablation electrode configured to ablate the at least one sample based on the return radiation" (Claim 1) "the at least one sample includes a heart, and further comprising generating at least one epicardial map of the heart based on the at least one return radiation" (Claim 15) "the at least one epicardial map includes at least one of (i) at least one coronary vessel in the heart, (ii) fat in the heart, (iii) scar tissue in the heart, or (iv) fibrosis in the heart" (Claim 16) "automatically assessing at least one substrate using the at least one optical substrate map, wherein the at least one substrate includes a pulmonary vein, a coronary artery, fat, fibrosis or a scar" (Claim 21) |  | 1 | [8 empty] | 1 | [9 empty] | 1 |  |  |  | 1 | [6 empty] | UNIV COLUMBIA | THE TRUSTEES OF COLUMBIA UNIVERSITY IN THE CITY OF NEW YORK | 2020-06-19 | 2020 | 2024-08-27 | 2024 | 2020-06-19 | 2020 | US202016906429 | US20210396661A1
412 | US12171524B2 | Devices, systems, and methods for imaging in certain endoscopic environments | Systems and methods for multi-modal imaging using an endoscope having an instrument channel, where the imaging is achieved without using the channel, are disclosed. The systems can include a multi-modal imaging paddle housing couple to a distal end of the endoscope. The housing can receive at least two imaging probes. The imaging probes can be an angle-resolved low-coherence interferometry probe (a/LCI) and an optical coherence tomography (OCT) probe. The housing can be scaled and positioned to be visible via the endoscope camera. The system and method can include locating the housing in a region of interest using the endoscope camera, acquiring OCT measurements to identify targets, and then acquiring a/LCI measurements at the identified targets. | 1. A system for multi-modal imaging using an endoscope having an instrument channel, wherein the multi-modal imaging is achieved without using the instrument channel, the system comprising: an endoscope selected from the group consisting of an anoscope, a colonoscope, a colposcope, an esophagoscope, a gastroscope, a laryngoscope, and a sigmoidoscope, the endoscope having a video camera positioned at a distal end for providing video images; a multi-modal imaging paddle housing coupled to the distal end of the endoscope, the multi-model imaging paddle housing occupying a retracted position adjacent to the distal end of the endoscope, the multi-modal imaging paddle housing optionally adapted to be movable between the retracted position and an extended position that is extended axially and/or radially from the distal end of the endoscope, the multi-modal imaging paddle housing adapted to receive a portion of at least two imaging probes, the multi-modal imaging paddle housing having a single imaging window, through which each of the at least two imaging probes acquire images, or two imaging windows, each through which a respective one of the at least two imaging probes acquires images; an angle-resolved low-coherence interferometry (a/LCI) probe having at least a distal portion received within the multi-modal imaging paddle housing; an optical coherence tomography (OCT) probe having at least a distal portion received within the multi-modal imaging paddle housing; and an imaging controller operatively coupled to the a/LCI probe and the OCT probe, wherein the multi-modal imaging paddle housing has a shape and a location in the retracted position, wherein the shape and the location combine to make the multi-modal imaging paddle housing visible in the video images provided by the video camera, such that a user can visually guide the multi-modal imaging paddle housing to a region of interest, wherein the a/LCI probe and the OCT probe are electronically coupled to the imaging controller and optically coupled to respective light sources and optical detectors via wires and optical fibers that are coupled to an exterior surface of the endoscope, wherein the a/LCI probe operates within a first wavelength range and the OCT probe operates in a second wavelength range, wherein the first and second wavelength ranges are different and non-overlapping, wherein the single imaging window transmits at least 95% of light at each wavelength in the first wavelength range and the second wavelength range, wherein one of the two imaging windows transmits at least 95% of light at each wavelength in the first wavelength range, wherein the other of the two imaging windows transmits at least 95% of light at each wavelength in the second wavelength range, wherein the instrument channel is not occupied by the a/LCI probe or the OCT probe. 2. The system of claim 1, wherein the multi-modal imaging paddle housing further comprises one or more radiofrequency ablation electrodes, wherein the one or more radiofrequency ablation electrodes are coupled to external power and controls via the wires that are coupled to the exterior surface of the endoscope. 3. The system of claim 1, wherein the multi-modal imaging paddle housing including a coupling cuff and is coupled to the distal end of the endoscope via the coupling cuff. 4. The system of claim 1, wherein the first wavelength range includes 840 nm within it and the second wavelength range includes 1310 nm within it. 5. The system of claim 1, wherein the a/LCI probe includes a reflection-only three-optic rotator prism and a two-axis scanning mirror. 6. The system of claim 5, wherein the controller is adapted to direct the a/LCI probe and the respective light sources and optical detectors to acquire A-scans by varying an angle of the reflection-only three-optic rotator prism. 7. The system of claim 1, wherein the a/LCI probe is coupled to the respective light sources and optical detectors via a pathlength-matched linear fiber array comprising a plurality of single-mode optical fibers, via coherent fiber bundles, or via a single multi-mode fiber. 8. The system of claim 1, wherein the OCT probe is a rotating probe. | An endoscopic multi-modal imaging accessory that attaches to the distal end without using the instrument channel, holding two probes (e.g., a/LCI and OCT). It enables dual-mode optical imaging without sacrificing the working channel. | R | "an angle-resolved low-coherence interferometry (a/LCI) probe having at least a distal portion received within the multi-modal imaging paddle housing; an optical coherence tomography (OCT) probe having at least a distal portion received within the multi-modal imaging paddle housing" (Claim 1) "the multi-modal imaging paddle housing further comprises one or more radiofrequency ablation electrodes, wherein the one or more radiofrequency ablation electrodes are coupled to external power and controls" (Claim 2) "the first wavelength range includes 840 nm within it and the second wavelength range includes 1310 nm within it" (Claim 4) |  | 1 | [6 empty] | 1 |  | 1 | [16 empty] | 1 |  |  |  | UNIV DUKE¦UNIV NORTH CAROLINA CHAPEL HILL | THE UNIVERSITY OF NORTH CAROLINA AT CHAPEL HILL | 2022-01-31 | 2022 | 2024-12-24 | 2024 | 2021-01-29 | 2021 | US202217589333 | US20220240782A1
413 | EP4106625A1 | SYSTEM AND METHOD FOR GUIDING DIRECTION TO AND TREATING TARGETS FOR ABNORMAL BIOLOGICAL RHYTHMS | Ablation catheters for treating electrical rhythm disorders include an array of sensor electrodes for detecting electrical signals to determine the location of target areas for treatment. If the catheter is not optimally positioned at the target area, the controller uses the detected signals to guide movement of the catheter toward the target area. Once proper positioning is confirmed, the controller activates the ablation components within the catheter to deliver energy to modify tissue at the target region. | 1. Claims of equivalent WO 2021168380 A1 CLAIMS: 1. A system for treating a heart rhythm disorder, the device comprising: a catheter configured to be placed in contact with a tissue surface, the catheter comprising: a flexible body having a contact surface; an array of sensor electrodes arranged within the flexible body, each sensor electrode having a conductive surface substantially flush with the contact surface, each sensor electrode configured to detect electrical signals from the tissue surface; and one or more treatment elements configured to deliver energy to the tissue surface; a plurality of conductors, each conductor having a distal end connected to one of a sensor electrode and the one or more treatment element; a controller in communication with proximal ends of the plurality of conductors, the controller comprising a processor configured to: receive the detected electrical signals; determine a location of a target region of a heart rhythm disorder; determine whether the catheter is optimally positioned at the target region, and, if not optimally positioned, to compute directionality to the target region and generate movement instructions to move the catheter toward the target region; and after determining that the catheter is optimally positioned, generate treatment signals to activate the one or more treatment elements to modify tissue in the target region. 2. The system of claim 1, wherein the flexible layer is generally planar and has a shape selected from a group consisting of a rectangle, an ellipse, and an annulus. 3. The system of claim 1 or claim 2, further comprising an elongated hollow shaft having a distal end, a proximal end, and a length, wherein the catheter is disposed at the distal end, the controller is disposed at the proximal end, and the plurality of conductors is retained within and extends the length the shaft, wherein the distal end of the shaft is manipulable from the proximal end. 4. The system of claim 3, further comprising a shaft motor configured to steer the distal end of the shaft in response to movement instructions generated by the controller. 5. The system of claim 3 or claim 4, further comprising a sheath slidably disposed on the shaft, the sheath having an interior volume configured to retain the catheter in a folded condition until the catheter is deployed by sliding the sheath away from the distal end of the shaft. 6. The system of any one of claims claim 2- 5, further comprising irrigant pores formed in the flexible body, the irrigant pores in fluid communication with an irrigant source associated with the controller, wherein the irrigant source is configured to deliver irrigant through the irrigant pores to tissue at the target region in conjunction with activation of the array of treatment elements. 7. The system of any one of claims 1-6, wherein the one or more treatment elements comprise an array of ablation electrodes, and wherein a subset of the plurality of conductors connected to the one or more treatment elements are electrical conductors configured to deliver electromagnetic energy to each ablation electrode. 8. The system of claim 7, wherein the array of sensor electrodes and the array of ablation electrodes are uniformly distributed around the contact surface. 9. The system of claim 7, wherein ablation electrodes of the array of ablation electrodes are evenly interspersed among the array of sensor electrodes. 10. The system of claim 7 wherein the processor is further configured to: determine a size of the target region based on the detected electrical signals; identify one or more ablation electrodes of the array of ablation electrodes based on at least the size and the location of the target region; and activate the identified ablation electrodes. 11. The system of any one of claims 1-10, wherein the processor determines the location of the target region by: generating a directionality map of heart rhythms based on the detected electrical signals, the directionality map describing pathways of heart rhythms; generating a guidance direction in which to move the flexible body towards the target region, and integrating the directionality map to determine the location of the target region . 12. The system of claim 11, wherein the directionality map is generated by applying a trained machine learning model to the electrical signals, wherein the machine learning model is trained on training examples comprising electrical signals of a human heart and known target regions of the heart rhythm disorder. 13. The system of any one of claims 7-10, wherein each ablation electrode is configured to emit a distinct waveform. 14. The system of claim 13, wherein the controller is configured to separately address one or more subsets of ablation electrodes of the array, and wherein the treatment signals comprise a first signal to a first subset of ablation electrodes to emit a first waveform and a second signal to a second subset of ablation electrodes to emit a second waveform. 15. The system of any one of claims 1-14, wherein the array of sensor electrodes comprises at least four electrodes. 16. The system of any one of claims 1-15, wherein the sensor electrodes are configured to deliver ablation energy so that the one or more treatment elements comprise the array of sensor electrodes. 17. The system of any one of claims 1-5, wherein the one or more treatment elements comprise one or more coolant chambers formed within the flexible body and configured for retaining a coolant, and wherein the plurality of conductors comprises a subset of conductors configured to direct a coolant fluid from a coolant source to the one or more coolant chambers to deliver freezing energy to tissue at the target region. 18. The system of claim 17, wherein the flexible body has a thermally conductive material incorporated therein to enhance conduction of freezing energy to tissue in contact with the contact surface. 19. The system of any one of claims 1-5, wherein the one or more treatment elements comprises an array of cryoablation loci formed within the flexible body, and wherein the plurality of conductors comprises a subset of conductors configured to direct a coolant fluid from a coolant source to the cryoablation loci in response to treatment signals from the controller to deliver freezing energy to tissue at the target region. 20. The system of any one of claims 1-6, wherein the one or more treatment elements comprise an array of targeting fiducials distributed within the flexible body, the targeting fiducials configured for guiding delivery of ablation energy from one or more external ablation energy sources. 21. The system of any one of claims 1-20, further comprising a contact sensor configured to determine whether the contact surface and the tissue surface are in adequate contact and, if not, to provide a signal to the controller to guide movement of the flexible body to provide improved contact. 22. A method for treating a heart rhythm disorder, the method comprising: detecting electrical signals of a heart using the array of sensor electrodes of the system of any one of claims 1-21; generating a directionality map describing pathways of heart rhythms based on the detected electrical signals; integrating the directionality map to determine one of: (i) a location of a target region of the heart rhythm disorder in the directionality map, and(ii) a guidance direction to the target region of the heart rhythm disorder outside the directionality map; determining whether the flexible body is optimally positioned at the target region according to the directionality map; and responsive to determining optimal positioning, activating the one or more treatment elements of the system to deliver energy to modify tissue at the determined location of the target region. 23. The method of claim 22, wherein generating the directionality map comprises applying a trained machine learning model to the detected electrical signals, wherein the machine learning model is trained on training examples comprising electrical signals of a human heart and known locations of one or more target regions of a heart rhythm disorder. 24. The method of claim 22, further comprising: responsive to determining the guidance direction to the target region outside the directionality map, steering the device to a subsequent position in the guidance direction to the target region. 25. The method of claim 22, further comprising: detecting subsequent electrical signals of the heart with the plurality of sensing electrodes after steering the device to the subsequent position; generating a subsequent directionality map describing pathways of heart rhythms based on the subsequent electrical signals; integrating the subsequent directionality map to determine one of: (i) the location of a target region of the heart rhythm disorder in the subsequent directionality map, and(ii) a subsequent guidance direction to the target region of the heart rhythm disorder outside the subsequent directionality map; and responsive to determining the subsequent guidance direction to the source region outside the directionality map, steering the flexible body to a third position in the subsequent guidance direction to the target region. 26. The method of any of claim 22, further comprising, responsive to determining the direction of the target region outside the directionality map, providing a notification on an electronic display to move the device in the direction of the target region. 27. The method of claim 22, further comprising responsive to determining the location of the target region, identifying the one or more ablation components within a threshold proximity to the location of the target region. 28. The method of claim 27, further comprising, responsive to determining the location of the target region determining a size of the target region based on the directionality map, wherein identifying the one or more ablation components is further based on at least the size of the target region. 29. The method of claim 22, wherein each ablation component is an ablation electrode, wherein instructing the one or more ablation components comprises actuating one or more ablation electrodes to emit one or more electromagnetic waves that modify tissue at the target region of the heart rhythm disorder. 30. The method of claim 29, wherein each ablation electrode is adjustable to generate a distinct waveform, wherein activating the one or more ablation electrodes includes: instructing a first ablation electrode to generate a first waveform; and instructing a second ablation electrode to generate a second waveform. 31. The method of claim 22, wherein each ablation component is a cryoablation loci, wherein instructing the one or more ablation components comprises directing coolant to one or more cryoablation loci that modify tissue at the target region of the heart rhythm disorder. 32. The method of claim 22, further comprising, responsive to determining the location of the target region, directing coolant into a chamber on the device cooling a portion of the heart inclusive of the target region. 33. The method of claim 22, further comprising, responsive to determining the location of the target region, venting irrigant from one or more irrigation pores on the device onto the heart to cool a portion of the heart inclusive of the target region. 34. The method of claim 22, further comprising detecting subsequent electrical signals of the heart with the plurality of sensing electrodes after ablation of the target region; and determining whether the heart rhythm disorder persists based on the subsequent electrical signals. 35. The method of claim 22, further comprising, responsive to determining that the heart rhythm disorder persists: generating a subsequent directionality map describing pathways of heart rhythms based on the subsequent electrical signals; integrating the subsequent directionality map to determine one of: a location of a second target region of the heart rhythm disorder in the subsequent directionality map, and a guidance direction to the second target region of the heart rhythm disorder outside the subsequent directionality map; and responsive to determining the location of the second target region, instructing one or more ablation components of a plurality of ablation components on the device to modify tissue at the second target region at the determined location of the second target region. | A cardiac-rhythm ablation catheter whose sensor electrode array detects electrical signals to find the target region and guides catheter movement toward it. Once positioned, the controller activates onboard ablation components to treat the rhythm disorder. | R | [0056] The processor may be further configured to determine a size of the target region based on the detected electrical signals, identify one or more ablation electrodes of the array of ablation electrodes based on at least the size and the location of the target region; and activate the identified ablation electrodes. | 1 | [13 empty] | 1 | [6 empty] | 1 | [9 empty] | UNIV LELAND STANFORD JUNIOR | UNIV LELAND STANFORD JUNIOR | 2021-02-20 | 2021 | 2022-12-28 | 2022 | 2020-02-20 | 2020 | EP20210756920 | AU2021224243A1¦AU2021224243B2¦CA3171471A1¦CN115768346A¦EP4106625A4¦IL295739A¦IL295739B1¦JP2023517284A¦US12193728B2¦US20210259765A1¦WO2021168380A1
414 | CN121015128A | Bimodal endoscopic probe with OCT (optical coherence tomography) imaging and plasma ablation functions | The invention provides a dual-mode endoscopic probe with OCT imaging and plasma ablation. The dual-mode endoscopic probe comprises a probe shell, a plasma ablation electrode device and an imaging device. The probe shell is provided with an imaging channel and an ablation channel which are arranged in parallel, the imaging device is arranged in the imaging channel, and the plasma ablation electrode device is arranged in the ablation channel; a coaxial needle type plasma ablation electrode is coupled with OCT imaging, plasma ablation is a minimally invasive precise ablation technology and is matched with optical coherence tomography, tomography is carried out on polyp and surrounding intestinal wall tissue before and after excision, the excision depth is judged, proper plasma excitation voltage and frequency are set, and excision time is shortened. Complications caused by incomplete polyp cutting or excessive polyp cutting are avoided, and accurate ablation cutting is achieved. | 1 A dual-modal endoscopic probe with OCT imaging and plasma ablation capabilities is characterized by comprising a probe housing, a plasma ablation electrode device, and an imaging device; the probe housing has parallelly arranged imaging channels and ablation channels, the imaging device is disposed within the imaging channel, and the plasma ablation device is disposed within the ablation channel; the imaging device includes a gradient refractive index lens, an optical fiber, a piezoelectric bicrystalline wafer, and a fixing device; the gradient refractive index lens is disposed at the distal end of the imaging channel; the optical fiber is located behind the gradient refractive index lens and is fixed to the centerline of the upper surface of the piezoelectric bicrystalline wafer, with one end of the optical fiber extending out of the surface of the piezoelectric bicrystalline wafer towards the gradient refractive index lens to form a cantilever beam structure; the piezoelectric bicrystalline wafer is fixed to the proximal end of the imaging channel by the fixing device. 2 According to claim 1, the dual-modal endoscopic probe with OCT imaging and plasma ablation is characterized in that the plasma ablation device includes a movable electrode, a heat shrink tubing, and a loop electrode; the outer layer of the movable electrode is the heat shrink tubing, and the outer layer of the heat shrink tubing wraps around the loop electrode; wherein, the end of the movable electrode extends beyond one end of the heat shrink tubing and bends towards the imaging side. 3 According to claim 1, the dual-modal endoscopic probe with OCT imaging and plasma ablation is characterized in that the gradient refractive index lens has an 8° end face and a gradient index of 0.327 mm.1。 4 The dual-mode endoscopic probe with OCT imaging and plasma ablation as described in claim 1 is characterized in that the optical fiber is a single-mode optical fiber, and the end face of the optical fiber has an 8° bevel. 5 According to claim 1, the dual-modal endoscopic probe with OCT imaging and plasma ablation is characterized in that the piezoelectric bicrystalline wafer is composed of two piezoelectric ceramics with the same polarization direction stacked together. 6 According to claim 1, the dual-modal endoscopic probe with OCT imaging and plasma ablation is characterized in that the fixing device consists of two opposing fixing plates, the inner surfaces of the two fixing plates are provided with slots, the slots are adapted to the sides of the piezoelectric bicrystalline wafer, and the outer surfaces of the two fixing plates are arc surfaces adapted to the inner surface of the imaging channel; the two sides of the piezoelectric bicrystalline wafer are sandwiched between the two fixing plates, and the piezoelectric bicrystalline wafer is adhered to the imaging channel by optical adhesive. 7 According to claim 1, the dual-modal endoscopic probe with OCT imaging and plasma ablation is characterized in that the probe housing is made of medical ABS material, and the junction of the gradient refractive index lens and the probe housing is coated with ultraviolet curing adhesive. 8 The dual-modal endoscopic probe with OCT imaging and plasma ablation according to claim 1 is characterized in that the surface of the gradient refractive index lens is coated with an anti-interference coating of inert tin oxide. | A dual-mode endoscopic probe combining OCT imaging and plasma ablation in parallel channels of the housing — a coaxial needle-type plasma electrode for ablation and an OCT probe for imaging side by side. It minimizes invasiveness while doing imaging-guided plasma ablation. | R | "the probe housing has parallelly arranged imaging channels and ablation channels, the imaging device is disposed within the imaging channel, and the plasma ablation device is disposed within the ablation channel" (Claim 1) "tomography is carried out on polyp and surrounding intestinal wall tissue before and after excision, the excision depth is judged, proper plasma excitation voltage and frequency are set, and excision time is shortened" (Abstract) "complications caused by incomplete polyp cutting or excessive polyp cutting are avoided, and accurate ablation cutting is achieved" (Abstract) |  | 1 | [15 empty] | 1 (plasma ablation) | [9 empty] | 1 |  |  |  | UNIV SHANGHAI SCIENCE & TECH | UNIV SHANGHAI SCIENCE & TECH | 2025-07-18 | 2025 | 2025-11-28 | 2025 | 2025-07-18 | 2025 | CN202510990859 | [NA]
415 | CN119326492A | Control system for ablation of duodenum through pulsed electric field | The invention relates to the field of medical physics, and discloses a control method and system for pulse electric field ablation of duodenum, and the method comprises the steps: determining a multi-target tissue, needing pulse electric field ablation, of the duodenum corresponding to a target user, collecting a target tissue image of the multi-target tissue, and constructing a multi-tissue three-dimensional model of the multi-target tissue; analyzing tissue parameters of the multi-target tissue, and analyzing electric field intensity and pulse parameters required for ablation of the multi-target tissue; multiple electrodes of the ablation equipment are configured, and multiple electrode implantation sites of the multiple electrodes are determined; constructing a digestive tract model from the corresponding oral cavity to the duodenum of a target user, analyzing an ideal transmission path of multiple electrodes, analyzing a path deviation coefficient of a simulation transmission path and the ideal transmission path, correcting the simulation transmission path, transmitting the multiple electrodes to a multi-electrode implantation site, and then performing multi-electrode implantation. And pulse electric field ablation of multi-target tissues is realized by utilizing multiple electrodes. The precise ablation effect of the pulsed electric field can be improved. | 1.A control method for pulsed electric field ablation of the duodenum, characterized in that the method comprises: determining multiple target tissues of the duodenum corresponding to the target user that need to be subjected to pulsed electric field ablation, acquiring target tissue images of the multiple target tissues, registering the target tissue images to obtain registered images, and constructing a multi-tissue three-dimensional model of the multiple target tissues based on the registered images; analyzing the tissue types of the multiple target tissues, analyzing the tissue parameters of the multiple target tissues based on the tissue types, and analyzing the electric field strength and pulse parameters required for ablating the multiple target tissues based on the tissue parameters; acquiring an ablation device for the pulsed electric field ablation, configuring multiple electrodes of the ablation device based on the electric field strength and tissue parameters, and determining multiple electrode implantation sites of the multiple electrodes based on the multi-tissue three-dimensional model; constructing a digestive tract model from the mouth to the duodenum corresponding to the target user, analyzing the ideal transmission path of the multiple electrodes based on the digestive tract model, constructing a mechanical electrode catheter of the multiple electrodes, acquiring the dynamic parameters of the mechanical electrode catheter in real time, and performing path simulation on the mechanical electrode catheter using a preset path simulation algorithm based on the dynamic parameters to obtain a simulated transmission path;The path deviation coefficients of the simulated transmission path and the ideal transmission path are analyzed, and based on the path deviation coefficients, the simulated transmission path is corrected to obtain a corrected simulated transmission path. When the corrected simulated transmission path is consistent with the ideal transmission path, based on the mechanical electrode catheter, after the multi-electrode is transmitted to the multi-electrode implantation site, the multi-electrode is used to achieve pulse electric field ablation of multiple target tissues according to the pulse parameters. 2.The control method of pulsed electric field ablation of the duodenum as described in claim 1 is characterized in that, the registering of the target tissue image to obtain a registered image includes: preprocessing the target tissue image to obtain a preprocessed tissue image; analyzing the key indicators of the preprocessed tissue image, and dividing the preprocessed tissue image into a reference image and a source image based on the key indicators; constructing an image transformation model of the source image, analyzing the transformation parameters of the image transformation model, and transforming the source image based on the transformation parameters to obtain a transformed source image; calculating the registration coefficient of the transformed source image and the reference image; when the registration coefficient is less than a preset registration coefficient threshold, optimizing the transformation parameters to obtain optimized transformation parameters, and when the optimized transformation parameters are greater than or equal to a preset registration coefficient threshold, using the optimized transformation parameters as optimal transformation parameters; based on the optimal transformation parameters, transforming the source image to obtain an optimal transformed original image; aligning the optimal source image and the reference image to obtain a registered image. 3.The control method for pulsed electric field ablation of the duodenum as described in claim 1 is characterized in that the analysis of the electric field strength and pulse parameters required for ablating the multiple target tissues based on the tissue parameters includes: extracting the multiple target tissue radius, multiple target tissue thickness, multiple target tissue dielectric constant and multiple target tissue conductivity of the multiple target tissues based on the tissue parameters; determining the electric field strength required for ablating the multiple target tissues through the multiple target tissue dielectric constant and multiple target tissue conductivity; calculating the pulse voltage required for ablating the multiple target tissues based on the electric field strength, multiple target tissue radius, multiple target tissue thickness, multiple target tissue dielectric constant and multiple target tissue conductivity; identifying the ablation requirements of the multiple target tissues, and determining the pulse width, pulse frequency and pulse sequence of the multiple target tissues based on the ablation requirements and the pulse voltage; The pulse parameters required for ablating the multiple target tissues are determined based on the pulse voltage, pulse width, pulse frequency and pulse sequence. 4.The control method for pulsed electric field ablation of the duodenum as described in claim 1 is characterized in that configuring the multiple electrodes of the ablation device based on the electric field strength and tissue parameters includes: determining the electrode material, electrode size and electrode shape of the ablation device based on the tissue parameters; determining the number of electrodes of the ablation device according to the number of multiple target tissues corresponding to the tissue parameters; configuring the electrode sensor of the ablation device, and monitoring the real-time electric field strength and real-time temperature of the multiple target tissues based on the electrode sensor; constructing an electrode safety mechanism of the ablation device through the real-time electric field strength and real-time temperature; configuring the multiple electrodes of the ablation device according to the electrode material, electrode size, electrode shape, electrode number and electrode safety mechanism. 5.The control method for pulsed electric field ablation of the duodenum as described in claim 1 is characterized in that determining the multi-electrode implantation site of the multi-electrode based on the multi-tissue three-dimensional model includes: analyzing the three-dimensional state of the multi-tissue three-dimensional model; determining the multi-electrode implantation range of the multi-electrode based on the three-dimensional state; identifying the multi-electrode initial implantation point of the multi-electrode implantation range; analyzing the straight-line distance between the multi-electrode initial implantation point and the multi-target tissues corresponding to the multi-tissue three-dimensional model according to the three-dimensional state; calculating the superimposed electric field strength of the multi-electrode based on the straight-line distance; if the superimposed electric field strength does not meet the preset electric field strength threshold, adjusting the implantation point position of the multi-electrode initial implantation point, and returning to the above-mentioned superimposed electromagnetic strength calculation step; if the superimposed electric field strength meets the preset electric field strength threshold, using the multi-electrode initial implantation point as the multi-electrode implantation site. 6.The control method for pulsed electric field ablation of the duodenum as described in claim 1 is characterized in that the construction of the digestive tract model from the mouth to the duodenum of the target user includes: using a preset medical imaging technology to collect image data from the mouth to the duodenum of the target user; performing data cleaning on the image data to obtain cleaned image data; registering the cleaned image data to obtain registered image data; based on the registered image data, extracting geometric structural features from the mouth to the duodenum of the target user; and constructing the digestive tract model from the mouth to the duodenum of the target user according to the geometric structural features. 7.The control method for pulsed electric field ablation of the duodenum as described in claim 1 is characterized in that the construction of the multi-electrode mechanical electrode catheter includes: obtaining simulation test results of the multi-electrode in the corresponding digestive tract model, and analyzing the performance requirements of the mechanical electrode catheter according to the simulation test results; determining the catheter structure for constructing the mechanical electrode catheter based on the performance requirements, wherein the catheter structure includes: a catheter geometric frame and a catheter material; integrating the microcircuit and sensor of the mechanical electrode catheter; constructing a real-time monitoring mechanism of the mechanical electrode catheter based on the microcircuit and the sensor; arranging the transmission device of the mechanical electrode catheter, and constructing a driving mechanism of the mechanical electrode catheter based on the transmission device. Advance mechanism; Construct the multi-electrode electrode mechanical catheter according to the catheter structure, real-time monitoring mechanism and advancement mechanism. 8.The control method for pulsed electric field ablation of the duodenum as described in claim 1 is characterized in that, based on the dynamic parameters, a preset path simulation algorithm is used to perform path simulation on the mechanical electrode catheter to obtain a simulated transmission path, including: extracting the current position vector, current velocity vector and current acceleration vector of the mechanical electrode catheter based on the dynamic parameters; calculating the future position vector of the mechanical electrode catheter based on the current position vector, current velocity vector and current acceleration vector; and performing path simulation on the mechanical electrode catheter based on the previous position vector and the future position vector using a preset path simulation algorithm to obtain a simulated transmission path. 9.The control method for pulsed electric field ablation of the duodenum as described in claim 1 is characterized in that the analysis of the path deviation coefficient of the simulated transmission path and the ideal transmission path includes: taking the same starting point corresponding to the simulated transmission path and the ideal transmission path as the origin of the three-dimensional coordinate system; based on the origin of the three-dimensional coordinate system, integrating the simulated transmission path and the ideal transmission path in the same three-dimensional coordinate system to obtain a three-dimensional curve of the simulated path and a three-dimensional curve of the ideal path; calculating the three-dimensional coordinate difference between the three-dimensional curve of the simulated path and the three-dimensional curve of the ideal path; calculating the sum of the absolute values of the three-dimensional coordinate differences; constructing a position deviation curve of the sum of the absolute values of the differences; analyzing the fluctuation characteristics of the position deviation curve; and calculating the path deviation coefficient of the simulated transmission path and the ideal transmission path based on the fluctuation characteristics. 10.A control system for pulsed electric field ablation of the duodenum, characterized in that it is used to execute the control method for pulsed electric field ablation of the duodenum as described in any one of claims 1 to 9, and the system includes: a tissue model construction module, used to determine the multi-target tissues that need to be subjected to pulsed electric field ablation in the duodenum of the target user, collect target tissue images of the multi-target tissues, register the target tissue images to obtain registered images, and construct a multi-tissue three-dimensional model of the multi-target tissues based on the registered images; an ablation parameter determination module, used to analyze the tissue types of the multi-target tissues, analyze the tissue parameters of the multi-target tissues based on the tissue types, and analyze the electric field strength and pulse parameters required for ablating the multi-target tissues based on the tissue parameters; an electrode implantation position confirmation module, used to obtain the ablation device for the pulsed electric field ablation, configure the multi-electrodes of the ablation device based on the electric field strength and tissue parameters, and determine the multi-electrode implantation sites of the multi-electrodes based on the multi-tissue three-dimensional model; The pulse electric field ablation module is used to construct a digestive tract model from the mouth to the duodenum of the target user, analyze the ideal transmission path of the multi-electrode based on the digestive tract model, and construct a mechanical electrode catheter for the multi-electrode.The dynamic parameters of the mechanical electrode catheter are collected in real time, and based on the dynamic parameters, a preset path simulation algorithm is used to simulate the path of the mechanical electrode catheter to obtain a simulated transmission path, and a path deviation coefficient between the simulated transmission path and the ideal transmission path is analyzed. Based on the path deviation coefficient, the simulated transmission path is corrected to obtain a corrected simulated transmission path. When the corrected simulated transmission path is consistent with the ideal transmission path, after the multi-electrode is transmitted to the multi-electrode implantation site based on the mechanical electrode catheter, the multi-electrode is used to implement the multi-electrode implantation according to the pulse parameters. Pulsed electric field ablation of target tissue. | A duodenal PFA control method that picks multiple target tissues, builds a 3D multi-tissue model from images, and analyzes parameters across them. The model supports staged, parameter-tuned PFA across the duodenum. | R | "analyzing the tissue types of the multiple target tissues, analyzing the tissue parameters of the multiple target tissues based on the tissue types, and analyzing the electric field strength and pulse parameters required for ablating the multiple target tissues based on the tissue parameters" (Claim 1) "extracting the multiple target tissue radius, multiple target tissue thickness, multiple target tissue dielectric constant and multiple target tissue conductivity of the multiple target tissues based on the tissue parameters" (Claim 3) "configuring the electrode sensor of the ablation device, and monitoring the real-time electric field strength and real-time temperature of the multiple target tissues based on the electrode sensor; constructing an electrode safety mechanism" (Claim 4) "calculating the superimposed electric field strength of the multi-electrode based on the straight-line distance; if the superimposed electric field strength does not meet the preset electric field strength threshold, adjusting the implantation point position" (Claim 5) | 1 | [7 empty] | 1 |  |  | 1 | [15 empty] | 1 |  |  |  | UNIV TIANJIN MEDICAL | UNIV TIANJIN MEDICAL | 2024-12-13 | 2024 | 2025-01-21 | 2025 | 2024-12-13 | 2024 | CN202411840143 | [NA]
416 | WO2022182876A1 | TREATMENT PLANNING FOR ATRIAL FIBRILLATION USING PATIENT-SPECIFIC MODELS BASED ON QUANTIFICATION OF FIBRILLATORY AREA | Systems and methods for developing an implementing an ablation plan for treating atrial fibrillation. A 3D digital model of cardiac tissue of a patient is segmented into a plurality of different sections and a fibrillatory area of each section is quantified. The fibrillatory area is based on a surface area of the section and a weighted surface area of each adjacent section, wherein the weighting is based on a proportion of a boundary between the section and the adjacent section that is not occupied by ablative scar tissue. An ablation plan is then defined based on the determining fibrillatory area of each section by adding planned ablation areas on the boundaries between sections until the quantified fibrillatory area of each section is below a maximum fibrillatory area threshold. The ablation plan is then displayed on a display screen to guide an ablation procedure according to the ablation plan. | 1. CLAIMSWhat is claimed is: 1. A method comprising: segmenting a 3D digital model of cardiac tissue of a patient into a plurality of different sections; quantifying, by an electronic controller, a fibrillatory area of each section of the plurality of different sections, wherein the fibrillatory area is quantified based at least in part on a surface area of the section and a weighted surface area of each adjacent section, wherein the weighted surface area of each adjacent section is calculated based on the surface area of the adjacent section and a proportion of a boundary between the section and the adjacent section that is not occupied by ablative scar tissue; defining an ablation plan adding at least one planned ablation area on the boundary between the section and the adjacent section until the quantified fibrillatory area of each section of the plurality of sections is below a maximum fibrillatory area threshold; and displaying the ablation plan overlaid onto the 3D digital model of the cardiac tissue on a display screen during performance of an ablation procedure according to the ablation plan. 2. The method of claim 1, further comprising performing the ablation procedure according to the ablation plan. 3. The method of claim 1, wherein defining the ablation plan further includes: identifying a first section of the plurality of sections with a largest quantified fibrillatory area, identifying a second section of the plurality of sections, wherein the second section is adjacent to the first section and has a larger fibrillatory area than any other section adjacent to the first section, and adding a first planned ablation area to the ablation plan, wherein the first planned ablation area is positioned along a boundary between the first section and the second section at a location where ablative scar tissue is not currently present. 4. The method of claim 3, wherein defining the ablation plan further includes: quantifying a first updated fibrillatory area for the first section, wherein the first updated fibrillatory area is quantified based at least in part on the surface area of the first section and a weighted surface area of each adjacent section, wherein the weighted surface area of the second section is calculated based on the surface area of the second section and a proportion of the boundary between the first section and the second section that is not covered with ablative scar tissue and is not covered by the first planned ablation area; and adding a second planned ablation area to the ablation plan in response to determining that the first updated fibrillatory area for the first section exceeds the maximum fibrillatory area threshold. 5. The method of claim 1, wherein defining the ablation plan further includes iteratively adding additional planned ablation areas on boundaries between the different sections until the quantified fibrillatory area of each section of the plurality of sections is below the maximum fibrillatory area. 6. The method of claim 1, wherein quantifying the fibrillatory area of each section of the plurality of sections included calculating the fibrillatory area of each section as wherein IΆ, is the fibrillatory area of a section wherein EAi is an effective area of the section wherein m is a number of sections adjacent to the section wherein g,, is a sum of lengths of a shared boundary between the section i and an adjacent section j that is not occupied by ablative scar tissue or a planned ablation area, wherein is a total length of the shared boundary between the section i and the adjacent section j, wherein EAj is an effective area of the adjacent section /, and wherein the effective area is a surface area of the section that is not occupied by ablative scar tissue or any planned ablation area. 7. The method of claim 1, further comprising: capturing image data of the cardiac tissue of the patient; generating the 3D digital model of the cardiac tissue of the patient based at least in part on the captured image data; and automatically detecting, by an electronic processor, locations of ablative scar tissue on the 3D digital model. 8. The method of claim 7, wherein capturing the image data includes capturing the image data of the cardiac tissue using late-gadolinium-enhancement (LGE) MRI, and wherein automatically detecting the locations of the ablative scar tissue on the 3D digital model includes automatically detecting locations highlighted by the LGE MRI. 9. The method of claim 7, wherein segmenting the 3D digital model of the cardiac tissue of the patient into the plurality of different sections includes positioning boundaries between the different sections to align with the detected locations of the ablative scar tissue in the 3D digital model. 10. The method of claim 1, further comprising: measuring an effective refractory period (ERP) of the patient; and determining the maximum fibrillatory area threshold based at least in part on the measured effective refractory period. 11. The method of claim 1, further comprising: applying a computer-based simulation of electrical activity to the 3D digital model of the cardiac tissue of the patient, wherein the fibrillatory area of one or more sections are adjusted during the computer-based simulation; and determining the maximum fibrillatory area threshold based at least in part on a minimum fibrillatory area required for observed rotor activity in the computer-based simulation. 12. The method of claim 1, wherein segmenting the 3D digital model of the cardiac tissue of the patient into the plurality of different sections includes segmenting the 3D digital model of a left atrium of the patient based on anatomical landmarks on the 3D model, wherein the plurality of different sections includes a section surrounding the left sided pulmonary veins, a section surrounding the right sided pulmonary veins, a section below the pulmonary veins in a lower part of a posterior wall of the left atrium, a section including an anterior wall of the left atrium, and a section including an appendage of the left atrium. 13. The method of claim 1, wherein no ablative scarring from a prior ablative procedure is present on the cardiac tissue of the patient, and wherein defining the ablation plan includes: defining an initial ablation plan including a plurality of planned ablation areas along the boundaries between the different sections according to a predefined initial ablation plan, and iteratively adding the at least one planned ablation area until the quantified fibrillatory area of each section of the plurality of sections is below the maximum fibrillatory area threshold. 14. The method of claim 13, wherein defining the initial ablation plan includes defining an initial ablation plan that includes planned ablation areas surrounding pulmonary veins of a left atrium of the patient and at least one additional planned ablation area on a border between two other sections of the plurality of different sections. 15. A computer-based system for developing and implementing an ablation plan for an ablation procedure to treat atrial fibrillation, the computer-based system comprising: a display screen; and an electronic controller configured to segment a 3D digital model of cardiac tissue of a patient into a plurality of different sections, quantify a fibrillatory area of each section of the plurality of different sections, wherein the fibrillatory area is quantified based at least in part on a surface area of the section and a weighted surface area of each adjacent section, wherein the weighted surface area of each adjacent section is calculated based on the surface area of the adjacent section and a proportion of a boundary between the section and the adjacent section that is not covered with ablative scar tissue, define an ablation plan adding at least one planned ablation area on the boundary between the section and the adjacent section until the quantified fibrillatory area of each section of the plurality of sections is below a maximum fibrillatory area threshold, and display the ablation plan overlaid onto the 3D digital model of the cardiac tissue on the display screen during performance of an ablation procedure according to the ablation plan | An atrial-fibrillation ablation planning system that segments a 3D cardiac model into sections and quantifies a fibrillatory area in each, weighted by adjacent sections. The fibrillatory-area map produces patient-specific ablation plans. | R | "quantifying, by an electronic controller, a fibrillatory area of each section of the plurality of different sections, wherein the fibrillatory area is quantified based at least in part on a surface area of the section and a weighted surface area of each adjacent section… proportion of a boundary between the section and the adjacent section that is not occupied by ablative scar tissue" (Claim 1) "defining an ablation plan adding at least one planned ablation area on the boundary between the section and the adjacent section until the quantified fibrillatory area of each section of the plurality of sections is below a maximum fibrillatory area threshold" (Claim 1) "capturing the image data of the cardiac tissue using late-gadolinium-enhancement (LGE) MRI… automatically detecting locations highlighted by the LGE MRI" (Claim 8) "applying a computer-based simulation of electrical activity to the 3D digital model… determining the maximum fibrillatory area threshold based at least in part on a minimum fibrillatory area required for observed rotor activity" (Claim 11 | [8 empty] | 1 | [11 empty] | 1 |  |  |  | 1 | [6 empty] | UNIV UTAH RES FOUND | UNIV UTAH RES FOUND | 2022-02-24 | 2022 | 2022-09-01 | 2022 | 2021-02-25 | 2021 | WO2022US17705 | [NA]
417 | DE112024002177T5 | MONITORING ABLATION PROGRESS USING LINEAR EBUS DATA | A system and method are presented for treating targeted tissue using cryoablation. A linear endobronchial ultrasound (EBUS) device is positioned adjacent the targeted tissue and is used to guide a percutaneously inserted cryoprobe into the targeted tissue. The EBUS device is partially rotated in order to create multiple image slices that are combined together. To overcome the acoustic shadow created by the ice ball, a 3D model is created of the ice ball based on the locations identified on visible portion of the ice ball's peripheral surface. These locations are mirrored across an axis defined for the cryoprobe to define an approximate location for the hidden, non-visible periphery. This model is then displayed along with the location of the targeted tissue and an indication of the killing zone of the ice ball defined by a selected isotherm. | 1. Method for treating a patient, comprising: a) placing an ultrasound device in a tracheal tree of the patient in a position to image a target tissue; b) placing a cryoprobe in the target tissue; c) creating an ice ball by cooling the cryoprobe; and d) during the creation of the ice ball: i) rotating the ultrasound device to create multiple 2D image slices of the target tissue, ii) combining the multiple 2D image slices using a computer to identify a visible surface of the ice sphere, iii) creating a 3D model of the ice ball based on the visible surface using the computer, iv) generating an isotherm for the 3D model at a selected temperature identified for effective ablation using the computer, and v) Displaying the 3D model of the ice sphere and the isotherm on a user interface using the computer. 2. Procedure according toClaim 1, where the ultrasound device is a linear endobronchial ultrasound device (EBUS). 3. Procedure according toClaim 2, wherein the linear EBUS device comprises a plurality of ultrasonic transducer elements; further wherein the plurality of ultrasonic transducer elements are of a type selected from a group consisting of PZT-based transducers, pMUT-based transducers and cMUT-based transducers. 4. Procedure according toClaim 1, which further includes displaying a representation of the target tissue relative to the 3D model of the ice sphere and the isotherm on the user interface. 5. Procedure according toClaim 1, further comprising comparing the 3D model of the ice ball on the computer with a known size and shape of the target tissue to identify portions of the target tissue outside of the isotherm; and displaying the identified portions of the target tissue on the user interface using an identifiable distinguishing visual feature. 6. Procedure according toClaim 1, in which a distal end of the ultrasound device further includes electromagnetic sensors that receive electromagnetic signals that locate the distal end in an electromagnetic field, and further comprising using the electromagnetic signals to display the 3D model of the ice ball on a representation of the target tissue on the user interface. 7. Procedure according toClaim 1, where the 3D model of the ice sphere and the isotherm displayed on the interface change over time in response to an increase in the size of the ice sphere. 8. Procedure according toClaim 1, further comprising terminating the generation of the ice sphere before an expected completion time when the user interface indicates that the isotherm includes the target tissue. 9. Procedure according toClaim 1, which further includes continuing to generate the ice ball beyond an expected end time when the user interface indicates that the isotherm does not enclose the target tissue. 10. Procedure according toClaim 1, where creating the 3D model includes: i) identifying an axis for the cryoprobe, ii) identifying 3D positions on the visible surface of the ice sphere, iii) mirroring the identified 3D positions across the axis for the cryoprobe to identify points on a hidden surface of the ice sphere, and iv) Combining the 3D positions on the visible surface with the identified points on the hidden surface. 11. Procedure according toClaim 10, where the reflection of the identified 3D positions across the axis for the cryoprobe occurs perpendicular to the axis and the points are the same distance from the axis as the identified 3D positions. 12. Procedure according toClaim 10, where only a selected number of 3D positions are reflected and the 3D model approximates positions between the selected number of 3D positions on the visible surface and the points on the hidden surface. 13. Procedure according toClaim 1, where generating the isotherm for the 3D model at the selected temperature includes: i) identifying a known distance from a circumference of the ice ball known to be at the selected temperature or colder; ii) Identify the isotherm located at this known distance from the circumference of the ice ball. 14. Procedure according toClaim 13, where the known distance is determined by testing temperatures in cryoablation tests. 15. The procedure afterClaim 13, further requiring that the selected temperature or a colder temperature be maintained for a specified period of time. 16. Procedure according toClaim 15, where the specified period of time is at least three minutes. 17. A system comprising: a) a cryoprobe; b) an ultrasound machine; c) an advertisement; d) a computer system that communicates with the cryoprobe, the ultrasound device and the display, the computer system having a processor operating under programming, the programming causing the computer system to: i) receiving 2D image slices of target tissue in a patient from the ultrasound device as it rotates within the patient, the 2D image slices including representations of an ice ball formed by the cryoprobe, ii) combining the 2D image slices to identify a visible surface of the ice ball, iii) creating a 3D model of the ice ball based on the visible surface, iv) generating an isotherm for the 3D model of the ice ball at a selected temperature identified for effective ablation, and v) Creating a user interface on the display that displays the 3D model of the ice sphere and the isotherm. 18. The system afterClaim 17, wherein the computer system further includes a model of the target tissue registered to the patient, and wherein the user interface displays a representation of the target tissue relative to the 3D model of the ice ball and isotherm. 19. The system afterClaim 18, whereby the programming also causes the computer system compares the 3D model of the ice sphere with a known size and shape of the target tissue to identify portions of the target tissue outside the isotherm; and further that the identified portions of the target tissue are displayed using an identifiable distinguishing visual feature. 20. The system afterClaim 17, where the programming causes the computer system to create the 3D model of the ice ball by causing the computer system to: i) identify an axis for the cryoprobe, ii) identify 3D positions on the visible surface of the ice sphere, iii) mirror the identified 3D positions across the axis for the cryoprobe to identify points on a hidden surface of the ice sphere, and iv) combine the 3D positions on the visible surface with the identified points on the hidden surface. 21. The system afterClaim 20, where the identified 3D positions transverse to the axis of the cryoprobe are mirrored perpendicular to the axis and the points are the same distance from the axis as the identified 3D positions. 22. The system afterClaim 20, where only a selected number of 3D positions are mirrored and the 3D model of the ice ball approximates positions between the selected number of 3D positions on the visible surface and the points on the hidden surface. 23. The system afterClaim 17, where generating the isotherm for the 3D model at the selected temperature includes: i) identifying a known distance from a circumference of the ice ball known to be at the selected temperature or colder; ii) Identify the isotherm located at this known distance from the circumference of the ice ball. 24. The system afterClaim 23, where the known distance is determined by testing temperatures in cryoablation tests. 25. The system afterClaim 23, further requiring that the selected temperature or a colder temperature be maintained for a certain period of time. 26. The system afterClaim 25, where the specified period of time is at least three minutes. 27. A computer system connected to a cryoprobe, an ultrasound device and a display, the computer system comprising: a. a processor; and b. a programming that causes the processor to: i) receiving 2D image slices of a patient's target tissue from the ultrasound device as it rotates within the patient, the 2D image slices containing representations of an ice ball formed by the cryoprobe, ii) combine the 2D image slices to identify a visible surface of the ice sphere, iii) create a 3D model of the ice sphere based on the visible surface, iv) generate an isotherm for the 3D model of the ice sphere at a selected temperature identified for effective ablation, and v) Creating a user interface on the display that shows the 3D model of the ice sphere and the isotherm. | A cryoablation monitoring system that uses a linear EBUS device to guide a cryoprobe and create multiple image slices, then builds a 3D model of the iceball to overcome acoustic shadows. The 3D iceball reconstruction reveals the freeze zone hidden by ultrasound shadows. | R | "rotating the ultrasound device to create multiple 2D image slices of the target tissue, combining the multiple 2D image slices using a computer to identify a visible surface of the ice sphere, creating a 3D model of the ice ball based on the visible surface using the computer, generating an isotherm for the 3D model at a selected temperature identified for effective ablation" (Claim 1) "comparing the 3D model of the ice ball on the computer with a known size and shape of the target tissue to identify portions of the target tissue outside of the isotherm; and displaying the identified portions of the target tissue on the user interface using an identifiable distinguishing visual feature" (Claim 5) "mirroring the identified 3D positions across the axis for the cryoprobe to identify points on a hidden surface of the ice sphere" (Claim 10) "the 3D model of the ice sphere and the isotherm displayed on the interface change over time in response to an increase in the size of the ice sphere" (Claim 7) |  |  | 1 |  |  |  |  | 1 | [6 empty] | 1 |  |  |  |  | 1 | [6 empty] | 1 |  |  |  |  | VERAN MEDICAL TECH INC | VERAN MEDICAL TECH INC | 2024-05-21 | 2024 | 2026-03-12 | 2026 | 2023-05-22 | 2023 | DE20241102177 | CN121398752A¦US12329568B2¦US20240389970A1¦WO2024243177A1
418 | CN120899385A | Compound eye laser ablation catheter, laser ablation device and control method thereof | The invention relates to an ablation catheter in an in-vivo laser ablation system, in particular to a compound eye laser ablation catheter, a laser ablation device and a control method of the laser ablation device. In order to overcome the defects that in an existing laser ablation technology, normal saline for cooling cannot be quantitatively controlled, local blood vessels are prone to being damaged in the ablation process, or an imaging device arranged at the head end of a catheter causes reduction of laser transmission efficiency and energy, and the ablation effect is affected due to the poor imaging monitoring effect, normal saline for cooling cannot be quantitatively controlled. The compound eye laser ablation catheter comprises a catheter body, the catheter body comprises a catheter inner cavity, a light transmission optical fiber layer and a catheter wrapping layer, the light transmission optical fiber layer comprises a plurality of light transmission optical fibers and heat conduction layers filled among the light transmission optical fibers, and a guide wire and a temperature sensor are arranged in the catheter inner cavity; an imaging module is arranged on the catheter wrapping layer and comprises at least two compound eye imaging assemblies, and the compound eye imaging assemblies are used for conducting real-time imaging on the ablation part and transmitting the image to the outside. | 1. L A compound eye laser ablation catheter, comprising a catheter body, the catheter body comprising a guidewire (2), a light-transmitting fiber layer and a catheter sheath (6) disposed outside the light-transmitting fiber layer, the light-transmitting fiber layer being used to transmit pulsed laser, and a catheter lumen (1) for injecting physiological saline being disposed along the central axis of the light-transmitting fiber layer, the guidewire (2) being disposed on the central axis of the catheter lumen (1); characterized in that: it further comprises M temperature sensors (3) and an imaging module, wherein, M22; the light-transmitting fiber layer comprising multiple light-transmitting fibers (4) and a heat-conducting layer (5) filling the spaces between the light-transmitting fibers (4), the distal end of the heat-conducting layer (5) being used to dock with a cold surface for heat conduction; The probes of the M temperature sensors (3) are located at the ablation end of the catheter body and are uniformly arranged in the catheter lumen (1) along the circumference of the guidewire (2); the imaging module is arranged on the catheter sheath (6) and close to the ablation end of the catheter body, including at least two compound eye imaging components (7) uniformly arranged along the circumference of the catheter sheath (6), the compound eye imaging components (7) are used to perform real-time imaging of the ablation area and transmit it to the outside. 2.According to claim 1, the compound eye laser ablation catheter is characterized in that: the compound eye imaging component (7) includes a compound eye lens disposed on the side wall of the catheter sheath (6) and a detector disposed corresponding to the compound eye lens; the compound eye lens is a microlens array disposed along an arc surface, and the microlens array includes multiple microlenses arranged in a hexagonal honeycomb structure; each microlens and its six neighboring microlenses are defined as a monitoring unit, and the microlenses of each monitoring unit are respectively configured with different working wavelengths. 3.The compound eye laser ablation catheter according to claim 1 or 2 is characterized in that: the number N of the optical transmission fibers (4) is calculated by the following formula: N = Stop ¥ Where, 2 is the outer radius of the optical fiber layer, j is the inner radius of the conduit, and are the radii of the optical fiber, and B is the fiber filling factor. 4.According to claim 3, the compound eye laser ablation catheter is characterized in that: the ablation end of the optical fiber layer is provided with a protective glass cover (8); The thermally conductive layer (5) is composed of thermally conductive potting compound injected into the gap between adjacent optical fibers (4); the temperature sensor (3) is a miniature thin-film temperature sensor; multiple optical fibers (4) are arranged concentrically around the central axis, with an optical fiber fill factor of 75.7%, or multiple optical fibers (4) are arranged in hexagonal layers from the inside to the outside around the central axis, with an optical fiber fill factor of 90.69%. 5A laser ablation device, characterized in that it comprises a laser source system, an intelligent temperature control system, the compound eye laser ablation catheter as described in claim 1, and an imaging data processing system; the output end of the laser source system is connected to the compound eye laser ablation catheter to provide a laser source; the intelligent temperature control system is connected to both the temperature sensor (3) and the laser source system, and is used to dynamically adjust the output power and repetition frequency of the laser source system based on the temperature data monitored by the temperature sensor (3); the input end of the imaging data processing system is connected to the output end of the imaging module, and is used to receive and process the image formed by the imaging module, and control the output power and repetition frequency of the laser source system based on the processing result, while simultaneously identifying the corresponding... area. 6.According to claim 5, a laser ablation device is characterized in that: the compound eye imaging component (7) includes a compound eye lens and a detector disposed on the side wall of the catheter envelope (6), the compound eye lens being a microlens array disposed along an arc surface, the microlens array including multiple microlenses arranged in a hexagonal honeycomb structure; each microlens and its six neighboring microlenses are defined as a monitoring unit, and the microlenses of each monitoring unit are respectively set with different working wavelengths to achieve large field-of-view multispectral compound eye imaging; the detector is disposed corresponding to the compound eye lens, used to detect and obtain multispectral images, and transmits the multispectral images of the ablation site to the imaging data processing system via wireless transmission; the imaging data processing system performs spectrochemical analysis on the composition of the ablation site based on the multispectral images. 7.A control method for a laser ablation device, characterized by comprising the following steps: Step 1, constructing the laser ablation device as described in claim 5; Step 2, activating the laser source system, the intelligent temperature control system, and the imaging data processing system; Step 3, the laser source system operates according to a preset output power and repetition frequency, while simultaneously injecting physiological saline into the lumen (1) of the catheter; Step 4, adjusting the output power and repetition frequency of the laser source system, as well as the injection speed of physiological saline, according to data from the intelligent temperature control system and the imaging data processing system; wherein the output power and repetition frequency are simultaneously decreased or simultaneously increased. 8.According to claim 7, a control method for a laser ablation device is characterized in that: in step 1, the imaging module of the compound eye laser ablation catheter in the laser ablation device is used to acquire multispectral images; step 4 specifically involves: detecting the temperature of the ablation site using a temperature sensor (3); adjusting the output power and repetition frequency of the laser source system and the injection speed of saline solution according to the temperature of the ablation site; and the imaging data processing system fitting the real-time spectral curve of the current region based on the multispectral image, comparing it with the spectral curve of the region under normal conditions, and reducing the output power and repetition frequency of the laser source system and increasing the injection speed of saline solution when the deviation is within a set threshold; and adjusting the repetition frequency to 0, stopping the output of the light source, and increasing the injection speed of saline solution to the maximum to achieve cooling when the deviation exceeds the set threshold. 9.According to claim 8, the control method of a laser ablation device is characterized in that, in step 4, the intelligent temperature control system adjusts the output power and repetition frequency of the laser source system according to the temperature of the ablation site, specifically: when the temperature of the ablation site and its surrounding tissue is 50-55°C, the output power is reduced to a lower level and the injection speed of physiological saline is increased; when the temperature of the ablation site and its surrounding tissue is greater than 55°C, the repetition frequency is adjusted to 0, the output of the laser source system is stopped, and physiological saline is injected at the maximum speed to cool down. | A 'compound-eye' laser ablation catheter with a cooling-saline control mechanism and head-end imaging that doesn't reduce laser transmission efficiency. The improved geometry lets the catheter image the site while maintaining laser energy. | R | the compound eye imaging components (7) are used to perform real-time imaging of the ablation area and transmit it to the outside. |  | 1 | [11 empty] | 1 | [16 empty] | 1 | XIAN INST OPTICS & PREC MECHANICS CAS | XIAN INST OPTICS & PREC MECHANICS CAS | 2025-08-22 | 2025 | 2025-11-07 | 2025 | 2025-08-22 | 2025 | CN202511184569 | [NA]
419 | NL2029307B1 | Cryoballoon Catheter for Treating Clinical Pulmonary Hypertension | The invention discloses a cryoballoon catheter for treating clinical pulmonary hypertension, which comprises a handle, a functional base, a balloon catheter and a balloon, wherein the front end of the balloon catheter is connected with an annular 5 open-loop balloon, and the rear end is integrally formed with the functional base and the handle to control the balloon, the main body of the balloon is an annular plastic tube, the annular surface of the annular plastic tube is perpendicular to the aXis of the balloon catheter, a cylindrical balloon is wrapped and installed on the annular plastic tube, and the front end of the annular plastic tube is provided with a 10 reserved hole, with the help of the annular plastic tube and the cylindrical balloon, the ablation is performed with less damage and higher fitting degree after the coolant is injected. Meanwhile, the injected coolant is reﬂuxed and discharged according to the designed reﬂux cycle to enhance the ablation effect, without blocking the blood ﬂow. At the same time, the reserved hole at the front end of the 15 balloon can detect the pulmonary artery pressure, which is reasonable in structural design, convenient and practical, and is suitable for popularization and application in medical institutions. | 1.A cryoballoon catheter for the treatment of clinical pulmonary hypertension, characterized in that the cryoballoon catheter comprises a handle, a functional base, a balloon catheter and a balloon, a leading end of the balloon catheter being tethered with an open-ended annular balloon loop, and a rear end integrally formed with the functional base and handle for operating the balloon; wherein said balloon body is an annular plastic tube pre-formed into an arc shape with an opening; the annular surface of the annular plastic tube is perpendicular to the axis of the balloon catheter, and a cylindrical balloon is wrapped and installed on the annular plastic tube, which cannot reach the fore of the annular plastic tube;the anterior end of the annular plastic tube is provided with a reserved hole, which is connected to the handle at the posterior end by a piezometer installed in the reserved hole and serves to realize monitoring of the pulmonary artery pressure;based on the deformation strength, the cylindrical balloon is divided into an outer expandable portion and an inner expandable portion, and the deformation strength of the inner expandable portion is greater than that of the outer expandable portion;an air inlet and an air outlet are respectively provided in the annular plastic tube at the front and back of the cylindrical balloon, and are respectively connected with an air inlet pipeline and an air outlet pipe; wherein the inside of the aforementioned balloon catheter is divided into four independent cavities, namely, the air inlet pipeline, the air outlet pipeline, a pressure monitor channel connected to the reserved hole, and a functional channel that can reach the front end of the annular plastic tube; the rear end of the balloon catheter is connected to the functional base;the top and bottom of the functional base are respectively provided with an air inlet valve and an air outlet valve; the rear end of the functional base is the handle, and the sensing base at the rear end of the handle is connected to the pressure monitoring unit and the functional unit. 2.A cryoballoon catheter for the treatment of clinical pulmonary hypertension according to claim 1, characterized in that a built-in thermocouple wire extending into the cylindrical balloon in the functional channel,and the functional unit connected to the posterior end of the thermocouple wire is a temperature measurement unit and a built-in pulmonary nerve signal detector is provided in the functional channel to monitor the pulmonary nerve signal and temperature of the cryo-balloon during surgery. 3.A cryoballoon catheter for the treatment of clinical pulmonary hypertension according to claim 1, characterized in that after the cylindrical balloon is filled with coolant, the annular plastic tube leans to the side of the inner expandable portion of the cylindrical balloon and from the axis of the cylindrical balloon. 4.A cryoballoon catheter for the treatment of clinical pulmonary hypertension according to claim 1, characterized in that the air inlet of the cylindrical balloon is provided on one side of the outer expandable portion, and the air outlet is provided on one side of the inner expandable portion of the cylindrical balloon; wherein the air inlet and the air outlet are disposed respectively at the front or rear of the cylindrical balloon. 5.A cryoballoon catheter for the treatment of clinical pulmonary hypertension according to claim 1, characterized in that the thickness of the outer expandable portion of the cylindrical balloon is smaller than that of the inner expandable portion. 6.A cryoballoon catheter for the treatment of clinical pulmonary hypertension flew according to claim 1, characterized in that the main body of the balloon adopts an open-close annular plastic tube, which can be formed into open-close rings of different specifications, and the attached cylindrical balloon while inflating the blood vessel closer to the 5blood vessel wall can be attached. 7.Use of the cryoballoon catheter according to any one of claims 1 to 6 in the treatment of clinical pulmonary hypertension. | A cryoballoon catheter for treating clinical pulmonary hypertension, with an annular open-loop balloon perpendicular to the catheter axis attached to a handle/functional base. The annular shape conforms to pulmonary vessel anatomy for cryotherapy. | R | "a built-in thermocouple wire extending into the cylindrical balloon in the functional channel… and a built-in pulmonary nerve signal detector is provided in the functional channel to monitor the pulmonary nerve signal and temperature of the cryo-balloon during surgery" (Claim 2) "the anterior end of the annular plastic tube is provided with a reserved hole, which is connected to the handle at the posterior end by a piezometer installed in the reserved hole and serves to realize monitoring of the pulmonary artery pressure" (Claim 1) "the cylindrical balloon is divided into an outer expandable portion and an inner expandable portion, and the deformation strength of the inner expandable portion is greater than that of the outer expandable portion" (Claim 1) | 1 |  |  |  | 1 | [9 empty] | 1 | [11 empty] | 1 |  |  |  |  | YANAN AFFILIATED HOSPITAL OF KUNMING MEDICAL UNIV | YANAN AFFILIATED HOSPITAL OF KUNMING MEDICAL UNIV | 2021-10-01 | 2021 | 2023-04-12 | 2023 | 2021-10-01 | 2021 | NL20212029307 | [NA]
420 | CN215384559U | Safe and intelligent ablation host combined with nerve monitoring | A safe and intelligent ablation host combined with nerve monitoring comprises an ablation needle interface, a nerve monitoring interface, a main control module, a signal switching module, an ablation module and a nerve monitoring module, and the output end of a radio frequency output circuit and the output end of a stimulation circuit are connected to the ablation needle interface through the signal switching module. The main control module is respectively connected with the signal switching module, the ablation module and the nerve monitoring module, the signal switching module comprises a relay, a switching signal end of the relay is respectively connected with the ablation module and the nerve monitoring module, and a control signal of the relay is connected with the main control module. The ablation electrode of the radiofrequency ablation needle is used as a stimulation electrode for nerve monitoring, a stimulation electrode does not need to be additionally arranged independently, a special ablation needle does not need to be produced independently, and the monitoring effect can be greatly improved through time-division intermittent work of radiofrequency ablation and nerve monitoring. If nerves are about to be damaged in the puncture and ablation process, an alarm can be given in advance, and damage is prevented. | 1 . A safe and intelligent ablation host combined with nerve monitoring, characterized in that it comprises an ablation needle interface, a nerve monitoring interface, a main control module, a signal switching module, an ablation module and a nerve monitoring module, wherein the ablation module includes a radio frequency output circuit, The nerve monitoring module includes a stimulation circuit, the radio frequency output circuit and the output end of the stimulation circuit are connected to the ablation needle connector through a signal switching module, and the main control module is respectively connected to the signal switching module, the ablation module and the nerve monitoring module, so The signal switching module includes a relay, the switching signal terminals of the relay are respectively connected to the ablation module and the nerve monitoring module, and the control signal of the relay is connected to the main control module. 2 . The safe and intelligent ablation host combined with nerve monitoring according to claim 1, wherein the nerve monitoring module comprises a stimulation circuit and a signal receiving and conditioning circuit, and the stimulation circuit is connected to the ablation needle connector through the signal switching module, so that the The ablation needle interface is connected to the ablation needle ablation electrode circuit, and the signal receiving and conditioning circuit is connected to an external nerve monitoring recording electrode through the nerve monitoring receiving interface. 3 . The safe and intelligent ablation host combined with nerve monitoring according to claim 2, wherein the signal receiving and conditioning circuit comprises a signal filtering and amplifying circuit, an AD conversion circuit and an algorithm monitoring circuit connected in sequence, and the algorithm monitoring circuit outputs The main control module is connected, and the signal filtering and amplifying circuit, the AD conversion circuit and the algorithm monitoring circuit are respectively connected to the main control module. 4 . The safe and intelligent ablation host combined with nerve monitoring according to claim 1, wherein the ablation module comprises a radio frequency output circuit and a temperature detection circuit, and the radio frequency output circuit comprises an external DC power input, a DC- A DC voltage regulation conversion circuit and a DC-AC radio frequency output circuit, the DC-AC radio frequency output circuit is connected to the ablation needle interface through a signal switching module, the ablation needle interface is connected to the ablation needle ablation electrode circuit, and the temperature detection circuit does not pass through The signal switch is directly connected to the thermocouple interface of the ablation needle interface. 5 . The safe and intelligent ablation host combined with nerve monitoring according to claim 4, wherein the ablation module further comprises an impedance monitoring circuit, and the output of the impedance monitoring circuit is connected to the main control module. 6 . The safe and intelligent ablation host combined with nerve monitoring according to claim 5, wherein the main control module is further connected with an alarm circuit. | A utility-model RF ablation host that combines neural monitoring: a signal switching module connects RF/stimulation circuits to a needle interface, with a separate neural monitoring interface and module. The setup integrates ablation and nerve safety monitoring in one host. | R | "the radio frequency output circuit and the output end of the stimulation circuit are connected to the ablation needle connector through a signal switching module… the signal switching module includes a relay" (Claim 1) "The ablation electrode of the radiofrequency ablation needle is used as a stimulation electrode for nerve monitoring, a stimulation electrode does not need to be additionally arranged independently" (Abstract) "If nerves are about to be damaged in the puncture and ablation process, an alarm can be given in advance, and damage is prevented" (Abstract) "the ablation module further comprises an impedance monitoring circuit, and the output of the impedance monitoring circuit is connected to the main control module" (Claim 5) | 1 |  |  |  | 1 | [5 empty] | 1 | [18 empty] | 1 |  | ZHEJIANG CURAWAY MEDICAL TECH CO LTD | ZHEJIANG CURAWAY MEDICAL TECH CO LTD | 2020-11-13 | 2020 | 2022-01-04 | 2022 | 2020-11-13 | 2020 | CN202022622766U | [NA]
421 | CN215899864U | Radio frequency ablation electrode needle with multi-dimensional nerve monitoring function | A radio frequency ablation electrode needle with a multi-dimensional nerve monitoring function comprises a nerve monitoring flexible electrode, a radio frequency thermal coagulation electrode and three knob modules used for adjusting the angle of the nerve monitoring flexible electrode, and an X-axis rotating knob, a Y-axis rotating knob and a Z-axis telescopic knob are arranged on a triangular fixing base. The X-axis rotary knob, the Y-axis rotary knob and the Z-axis telescopic knob are connected with the nerve monitoring flexible electrode through control stay wires respectively, the front end of the triangular fixing seat is fixedly connected with the radio frequency thermal coagulation electrode, and the rear end of the triangular fixing seat is in butt joint with a cable. Compared with the prior art, the electrode tip stretching in multiple dimensions penetrates out of the head end of the radio frequency thermal coagulation electrode, front-back and left-right deflection in the X/Y direction and stretching in the Z direction can be achieved, and the detection position of the flexible electrode tip can be controlled through the X/Y/Z rotary knob; the monitoring electrode disclosed by the utility model can be used for monitoring an ablated nerve focus part in all directions, so that the problem of incomplete nerve damage or excessive nerve damage is avoided to a certain extent. | 1. A radiofrequency ablation electrode needle with multi-dimensional nerve monitoring, characterized in that it includes a nerve monitoring flexible electrode, a radiofrequency thermocoagulation electrode, a knob module for adjusting the angle of the nerve monitoring flexible electrode, and a cable, and the knob module includes a triangular A fixed seat, an X-axis rotation knob, a y-axis rotation knob and a Z-axis telescopic knob, the X-axis rotation knob, the y-axis rotation knob and the Z-axis telescopic knob are respectively arranged on the triangular fixed seat, the X-axis rotation knob, the y-axis rotation knob, The shaft rotation knob and the Z-axis telescopic knob are respectively connected to the nerve monitoring flexible electrodes through control cables, the front end of the triangular fixing base is fixedly connected to the radio frequency thermocoagulation electrode, and the rear end is connected to the cable. 2. The radiofrequency ablation electrode needle with multi-dimensional nerve monitoring according to claim 1, wherein a guide rail device is provided inside the radiofrequency thermocoagulation electrode, and the nerve monitoring flexible electrode is installed on the radiofrequency thermocoagulation electrode through the guide rail device. Internally and isolated from it, the nerve monitoring flexible electrode protrudes from the front end of the radiofrequency thermocoagulation electrode. 3. The radiofrequency ablation electrode needle with multi-dimensional nerve monitoring according to claim 2, wherein the control wire comprises an X-axis wire, a y-axis wire and a Z-axis wire, and the front ends of the flexible electrodes for nerve monitoring are respectively provided with There are X-axis cable fixing points, X-axis corner buckles, y-axis cable fixing points, and y-axis corner buckles. One end of the X-axis cable is wound on the X-axis rotating knob, and the other end is fixed on the X-axis cable through the X-axis corner buckle. On the fixed point, one end of the y-axis cable is wound on the y-axis rotating knob, the other end is fixed on the y-axis cable fixing point through the y-axis corner buckle, and one end of the Z-axis cable is wound on the Z-axis telescopic knob, and the other end is fixed on the y-axis cable fixing point. One end is connected to the side of the rear end of the flexible electrode for nerve monitoring, the radio frequency thermocoagulation electrode is provided with a wire tube inside, and the X-axis pulling wire, the y-axis pulling wire and the Z-axis pulling wire are arranged through the wire tube. 4The radiofrequency ablation electrode needle with multi-dimensional nerve monitoring according to claim 3, wherein a thermocouple is arranged at the inner front end of the radiofrequency thermocoagulation electrode. 5. The radiofrequency ablation electrode needle with multi-dimensional nerve monitoring according to claim 3, wherein the bottom surface of the triangular holder is an equilateral triangle, and the triangular holder comprises an electrode communication socket, a triangular holder body and a A vertebral body with a 6% Luer taper, the X-axis rotation knob, the y-axis rotation knob and the Z-axis telescopic knob are respectively located at the top three corners of the bottom surface of the triangular fixed seat, and the electrode communication socket is located on the triangular fixed seat The center position of the triangular fixation seat, there is a cavity inside the triangular fixation seat, the top of the triangular fixation seat is provided with a vertebral body with a 6% Luer taper, and the radiofrequency thermocoagulation electrode is inserted into the vertebral body with a 6% Luer taper, connected to Assemble in a sealed box with glue. 6. The radio frequency ablation electrode needle with multi-dimensional nerve monitoring according to claim 5, wherein the two ends of the cable are respectively provided with an electrode communication plug and a radio frequency host connection plug, and the electrode communication plug is used to connect the triangle The electrode communication socket of the fixed seat, the radio frequency host connection plug is used to connect the radio frequency host. 7The radiofrequency ablation electrode needle with multi-dimensional nerve monitoring according to claim 6, wherein the electrode communication plug is in the shape of a meniscus, and the electrode communication socket is in the shape of a meniscus. | A utility-model RF ablation electrode needle with multi-dimensional nerve monitoring, including a flexible nerve monitoring electrode plus X/Y rotational and Z-axis telescopic knobs for angle adjustment. Operators can fine-tune the nerve electrode position relative to the RF tip. | R | "a nerve monitoring flexible electrode, a radio frequency thermocoagulation electrode and three knob modules used for adjusting the angle of the nerve monitoring flexible electrode" (Abstract) "the electrode tip stretching in multiple dimensions penetrates out of the head end of the radio frequency thermocoagulation electrode, front-back and left-right deflection in the X/Y direction and stretching in the Z direction can be achieved" (Abstract) "the monitoring electrode disclosed by the utility model can be used for monitoring an ablated nerve focus part in all directions, so that the problem of incomplete nerve damage or excessive nerve damage is avoided to a certain extent" (Abstract) "a thermocouple is arranged at the inner front end of the radiofrequency thermocoagulation electrode" (Claim 4) | 1 |  |  |  | 1 | [5 empty] | 1 | [18 empty] | 1 |  | ZHEJIANG CURAWAY MEDICAL TECH CO LTD | ZHEJIANG CURAWAY MEDICAL TECH CO LTD | 2020-12-29 | 2020 | 2022-02-25 | 2022 | 2020-12-29 | 2020 | CN202023259802U | [NA]
422 | CN117731246A | Ablation point position and temperature field monitoring device based on thin film optical fiber sensing array | The invention discloses an ablation point position and temperature field monitoring device based on a thin-film optical fiber sensing array, and the device comprises the thin-film optical fiber sensing array which is used for being attached to the surface of a human body for monitoring a body surface temperature field; the signal demodulation module is used for demodulating the optical signals in the thin film optical fiber sensing array into temperature data; the upper computer is used for obtaining ablation point positions and temperature field distribution by utilizing an ablation point position and temperature field prediction algorithm according to the temperature data, and the ablation point position and temperature field prediction algorithm is constructed through the following processes: generating a training data set according to array arrangement information sensed by the thin film optical fiber and an ablation point temperature field prediction model; training a neural network by using the training data set; and acquiring temperature field information data of the thin-film optical fiber sensing array obtained by the calibration device, so that the trained neural network is corrected and calibrated to obtain an ablation point position and a temperature field prediction algorithm. | 1· An ablation point position and temperature field monitoring device based on a thin film optical fiber sensing array, which is characterized by including: a thin film optical fiber sensing array, the thin film optical fiber sensing array is used to attach to the surface of the human body to monitor the body surface temperature field Monitoring; a signal demodulation module, the signal demodulation module is used to demodulate the optical signal in the thin film optical fiber sensing array into temperature data; a host computer, the host computer is used to use ablation according to the temperature data The point position and temperature field prediction algorithm obtains the ablation point position and temperature field distribution. The ablation point position and temperature field prediction algorithm is constructed through the following process: According to the array arrangement information of the thin film optical fiber sensing and the ablation point temperature field prediction model Generate a training data set, and use the training data set to train the neural network; obtain the temperature field information data of the thin film optical fiber sensing array obtained through the calibration device, thereby correcting and calibrating the trained neural network to obtain the ablation point Position and temperature field prediction algorithm. 2. The device according to claim 1, wherein the thin film optical fiber sensing array includes a flexible film and an optical fiber sensing array, and the optical fiber sensing array is based on the topological design of an equally spaced parallel grid and is shaped by a mold. The optical fiber arrangement is adhered to the heat shrink film and fixed on the surface of the flexible film. 3. The device according to claim 2, wherein the optical fiber sensing array includes a temperature sensing segment optical fiber and a connecting segment optical fiber, and the temperature sensing segment optical fiber is arranged on the flexible film in an array form, It is used for temperature field monitoring, and the connecting section optical fiber is used to connect adjacent temperature sensing section optical fibers. 4. The device according to claim 1, characterized in that the signal demodulation module is connected to the thin film optical fiber sensing array through an optical fiber jumper. 5. The device according to claim 1, characterized in that the signal demodulation module is connected to the host computer through a communication optical cable. 6. The device according to claim 1, wherein the neural network is a long short-term memory network, a bidirectional long short-term memory network or a gated recurrent unit neural network. 7. The device according to claim 1, characterized in that generating a training data set based on the array arrangement information of the thin film optical fiber sensing and the ablation point temperature field prediction model includes: based on the temperature in the thin film optical fiber sensing array. Construct a temperature measuring point information matrix based on the position of the measuring point; according to the thin film optical fiber sensing array temperature measuring point position information matrix and the preset ablation point position and temperature information, a thin film optical fiber sensor is generated through the classic ablation point temperature field prediction model The array temperature field distribution data and the major diameter and short diameter of the ablation temperature affected area are then used to obtain a training data set for neural network training; wherein, the temperature field distribution data of the thin film optical fiber sensing array is the training input in the training data set. , the ablation point position, temperature and the major axis and minor axis of the ablation temperature affected area are the training outputs in the training data set. 8. The device according to claim 1, characterized in that the ablation point temperature field prediction model is: - except not^πk^\d~ +x wherein, % is the ablation point ablation temperature, k is the human body at the ablation site Thermal conductivity of tissue, d is the vertical distance from the ablation point to a position near the ablation point, quality. 9. The device according to claim 1, characterized in that, the thin film optical fiber sensing array obtained by the calibration device is obtained. A series of temperature field information data is obtained to correct and calibrate the trained neural network, including: obtaining the temperature field information data of the thin film optical fiber sensing array obtained through the calibration device, wherein the temperature field information data of the thin film optical fiber sensing array is obtained through the calibration device. The position movement and temperature adjustment of the point-like heat source inside the calibration device are obtained by simulating different ablation positions and temperature conditions; the data are used to correct and calibrate the trained neural network. 10. The device according to claim 1, wherein the calibration device includes simulated human tissue, a fiber grating array and a point heat source, and the point heat source simulates different ablation positions and temperatures by moving the position and adjusting the temperature, so The fiber grating array is arranged on the simulated human tissue in vertical and horizontal directions for temperature calibration. | A surface-mounted thin-film fiber-optic sensing array that monitors body-surface temperature and ablation-point position, plus a signal demodulator and host computer. The array reads temperature/position non-invasively during ablation. | R | "a thin film optical fiber sensing array, the thin film optical fiber sensing array is used to attach to the surface of the human body to monitor the body surface temperature field; a signal demodulation module… is used to demodulate the optical signal in the thin film optical fiber sensing array into temperature data" (Claim 1) "the host computer is used to use ablation point position and temperature field prediction algorithm obtains the ablation point position and temperature field distribution… train the neural network… correcting and calibrating the trained neural network" (Claim 1) "the temperature field distribution data of the thin film optical fiber sensing array is the training input… the ablation point position, temperature and the major axis and minor axis of the ablation temperature affected area are the training outputs" (Claim 7) "the neural network is a long short-term memory network, a bidirectional long short-term memory network or a gated recurrent unit neural network" (Claim 6) |  | 1 | [19 empty] | 1 | [9 empty] | ZHEJIANG LAB | ZHEJIANG LAB | 2023-12-18 | 2023 | 2024-03-22 | 2024 | 2023-12-18 | 2023 | CN202311738770 | [NA]
423 | WO2022067146A3 | SYSTEMS, DEVICES AND METHODS FOR TREATING LUNG TUMORS WITH A ROBOTICALLY DELIVERED CATHETER | A system for treatment of a target region of lung tissue including: a flow regulator configured to be interposed between a conductive fluid source and a conductive fluid outlet positionable at or in proximity of the target region of lung tissue, the flow regulator being further configured for controlling a flow rate or a bolus quantity of conductive fluid coming from the fluid source and delivered to the conductive fluid outlet; and a controller communicatively connectable with said flow regulator and with at least one sensor, with the at least one sensor being configured for detecting values taken by at least one control parameter representative of a physical property, wherein the physical property is one of temperature (T), pressure (p), electric impedance (Z), or electric conductivity (C) of material present at or in proximity of the target region of lung tissue. | 1. CLAIMS: The invention is: 1. A system for treatment of a target region of lung tissue, the system comprising: at least one a flow regulator configured to be interposed between a conductive fluid source and a conductive fluid outlet positionable at or in proximity of the target region of lung tissue, the flow regulator being further configured for controlling a flow rate or a bolus quantity of conductive fluid coming from the fluid source and delivered to the conductive fluid outlet; a controller communicatively connectable with said flow regulator and with at least one sensor, with the at least one sensor being configured for detecting values taken by at least one control parameter representative of a physical property, wherein the physical property is one of temperature (T), pressure (p), electric impedance (Z), or electric conductivity (C) of material present at or in proximity of the target region of lung tissue; wherein the controller is configured for: receiving from said sensor signals representative of detected values of the control parameter; controlling the flow regulator based on one or more detected values of the control parameter, wherein controlling the flow regulator comprises executing a control cycle including: o controlling the flow regulator in a high delivery mode, wherein in the high delivery mode: ■ the flow rate of conductive fluid delivered to the conductive fluid outlet is equal or above a set high flow rate, or■ the bolus quantity of conductive fluid delivered to the conductive fluid outlet is equal above a set high bolus quantity, o controlling the flow regulator in a low delivery mode, wherein in the low delivery mode: ■ the flow rate of conductive fluid delivered to the conductive fluid outlet is equal or below a set low flow rate smaller than the set high flow rate, or ■ the bolus quantity of conductive fluid delivered to the conductive fluid outlet is equal or below a set low bolus quantity smaller than the set high bolus quantity. 2. A system according to claim 1, wherein in the low delivery mode: ■ the flow rate of conductive fluid delivered to the conductive fluid outlet is equal or below a set low flow rate smaller than 50% of the set high flow rate, or■ the bolus quantity of conductive fluid delivered to the conductive fluid outlet is equal or below a set low bolus quantity smaller than 50% of the set high bolus quantity. 3. A system according to claim 1 or 2, wherein in the low delivery mode the set low flow rate is between 0 and 5 ml/min or wherein the set low bolus quantity is between 0 and 10 ml. 4. A system according to claim 1 or 2 or 3, wherein in the high delivery mode the set high flow rate is between 2 and 16 ml/min or the set high bolus quantity is between 0.3 and 60 ml. 5. A system according to any one of the preceding claims, wherein controlling the flow regulator comprises repeatedly executing said control cycle. 6. A system according to any one of the preceding claims, comprising at least one ablation element positionable at the target region of the lung tissue and connectable to an ablation source. 7. A system according to claim 6, comprising at least one flexible shaft configured to advance through an airway passage of a lung and having an active portion positionable at the target region of the lung tissue and including the at least one ablation element. 8. A system according to any one of the preceding claims, comprising the at least one sensor, the sensor being configured to be positionable at the target region of the lung tissue. 9. A system according to claim 7 in combination with any one of claims 1-6 and 8, wherein the at least one sensor is carried by the active portion of said flexible shaft. 10. A system according to claim 7 in combination with any one of claims 1-6 and 8, wherein the at least one sensor is configured to be positioned in correspondence of a volume surrounding the active portion of said flexible shaft. 11. A system according to claim 9 or 10, wherein the at least one sensor is configured for sensing values taken by the at least one control parameter, and wherein the physical property is one of temperature, pressure, electric impedance, or electric conductivity of material present in a volume surrounding the active portion. 12. A system according to any one of the preceding claims comprising the conductive fluid outlet which is configured to be placed in fluid communication with the conductive fluid source. 13. A system according to claim 7 and 12, wherein the conductive fluid outlet is carried by the flexible shaft active portion. 14. A system according to claim 7 and 12, wherein the conductive fluid outlet is configured to be positioned in correspondence of said volume surrounding the active portion. 15. A system according to any one of the preceding claims 6-14, wherein the controller is connectable with said ablation source and configured for controlling the ablation energy source to deliver ablation energy to the at least one ablation element. 16. A system according to claim 15, further wherein the controller is configured for executing said steps of: receiving from said sensor signals representative of sensed values of the control parameter, and105 controlling the flow regulator based on one or more sensed values of the control parameter and executing, optionally repeatedly executing, said control cycle, while the controller commands the ablation energy source to deliver ablation energy to the at least one ablation element. 17. A system according to any one of the preceding claims 7-16, comprising an electrical connector carried by the flexible shaft and adapted to electrically connect the at least one ablation element to the ablation source. 18. A system according to any one of the preceding claims, wherein the control cycle includes: verifying if one or more sensed values of the control parameter fall below a set low threshold (T Low), and wherein said controlling the flow regulator to low delivery mode is executed if the one or more sensed values of the control parameter fall below the set low threshold (T Low). 19. A system according to any one of the preceding claims, wherein the control cycle includes: verifying if one or more sensed values of the control parameter exceed a set high threshold (T High, Z High), and wherein said controlling the flow regulator to high delivery mode is executed if the one or more sensed values of the control parameter exceed the set high threshold (T High, Z High). 20. A system according to any one of the preceding claims, wherein the control cycle includes: periodically verifying if one or more sensed values of the control parameter fall below a set low threshold (T Low),106 switching the flow regulator from high delivery mode to low delivery mode when the one or more sensed values of the control parameter fall below the set low threshold (T Low); optionally, wherein said step of periodically verifying is executed at least 10 times per second. 21. A system according to any one of the preceding claims, wherein the control cycle includes: periodically verifying if one or more sensed values of the control parameter exceed a set high threshold (T High, Z High), switching the flow regulator from low delivery mode to high delivery mode when the one or more sensed values of the control parameter exceed the set high threshold (T High, Z High); optionally, wherein said step of periodically verifying is executed at least 10 times per second. 22. A system according to any one of the preceding claims wherein the controller is configured for repeating the control cycle a plurality of times during a same treatment session. 23. A system according to claim 22, wherein the controller is configured to control the flow regulator in high delivery mode or in low delivery mode for a respective time interval, and wherein a duration of said respective time intervals is either predetermined or determined by detection of a triggering event. 24. A system according to claim 23, wherein the controller is configured to determine duration of said time intervals by detection of a triggering event, wherein detection of the triggering event comprises one or more of: detection that one or more values of the sensed parameter exceeds a set very high threshold (T Overheat),107 detection that one or more values of the sensed parameter exceeds said set high threshold (T High, Z High), detection that one or more values of the sensed parameter falls below a set low threshold (T Low). 25. A system according to claim 22 or 23 or 24, wherein the controller is configured to execute the same treatment session which includes a plurality of time intervals where the flow regulator is adjusted to low delivery mode intercalated by time intervals where the flow regulator is adjusted to high delivery mode, thereby reducing the overall amount of conductive fluid delivered over said treatment session while maintaining under control the detected values of the parameter. 26. A system according to any one of the preceding claims, wherein the step of controlling the flow regulator to low delivery mode comprises: adjusting the flow regulator to maintain the flow rate of conductive fluid to the conductive fluid outlet equal or below said set low flow rate during a low delivery time interval (Flow Low Time), in particular comprised between 1 to 10 seconds; or adjusting the flow regulator to deliver to the conductive fluid outlet the bolus quantity of conductive fluid equal or below said set low bolus quantity within a low delivery time interval (Flow Low Time), in particular comprised between 1 to 10 seconds. 27. A system according to claim 26, wherein the cycle comprises a sub-routine optionally executed after expiration of said low delivery time interval, said sub-routine including: a further step of verifying if one or more values of the sensed parameter falls below or above the set low threshold (T Low), in case one or more values of the parameter sensed in the further step of verifying falls below the set low threshold (T Low), assigning a decreased value to the set low flow rate or to the set low bolus quantity, and108 repeating controlling the flow regulator to low delivery mode using the decreased value of the set low flow rate or the decreased value of set low bolus quantity. 28. A system according to any one of the preceding claims, wherein the step of controlling the flow regulator to high delivery mode comprises: adjusting the flow regulator to maintain the flow rate of conductive fluid to the conductive fluid outlet equal or above said set high flow rate during a high delivery time interval (Flow High Time), in particular comprised between 1 to 30 seconds; or adjusting the flow regulator to deliver to the conductive fluid outlet the bolus quantity of conductive fluid equal or above said set high bolus quantity within a high delivery time interval (Flow High Time), in particular comprised between 1 to 30 seconds. 29. A system according to claim 28, wherein the cycle comprises a sub-routine optionally executed after expiration of said high delivery time interval, said sub-routine including: a further step of verifying if one or more values of the sensed parameter falls below or above the set low threshold (T Low), in case one or more values of the parameter sensed in the further step of verifying remains above the set low threshold (T Low), assigning an increased value to the set high flow rate or to the set high bolus quantity, and repeating controlling the flow regulator to high delivery mode using the increased value of the set high flow rate or the increased value of set high bolus quantity. 30. A system according to claim 29, wherein the cycle provides for repeating the subroutine of claim 29 until when the further step of verifying if one or more values of the sensed falls below the set low-threshold (T Low) is positively passed. 31. A system according to claim 16 and 30, wherein the controller is configured to interrupt or reduce delivery of ablation energy to the at least one ablation element if it has109 determined that, after a predetermined number of repetitions of the subroutine of claim 29, the step of verifying if one or more values of the sensed fall below the set low threshold (T Low) is not positively passed. 32. A system according to any one of the preceding claims wherein the cycle comprises: determining occurrence of a safety relevant event if the one or more parameter values are above a set over-high threshold (T Over High; Z Over High), which is greater than said high threshold (T High; Z High); if a safety relevant condition is determined, then: o temporarily adjust down power supplied to the ablation energy source and/or o control the flow regulator to a very high delivery mode, wherein in the very high delivery mode the flow rate of conductive fluid delivered to the conductive fluid outlet is equal or above a set very high flow rate greater than the set high flow rate, or the bolus quantity of conductive fluid delivered to the conductive fluid outlet is equal or above a very set high bolus quantity greater than the high bolus quantity. 33. A system according to any one of the preceding claims 6 to 32, wherein the controller is configured for maintaining power supplied by the ablation energy source in a range comprised between 20 and 200 W over a major portion of the treatment session, optionally over the entire treatment session. 34. A system according to any one of preceding claims 6-33, wherein the controller is configured to increase power supplied by the ablation energy source from an initial value to a regimen value, during an initial portion of the treatment session optionally lasting between 10% and 30% of the entire treatment session. 11035. A system according claim 34, wherein the controller is configured to maintain power supplied by the ablation energy source at the regimen value during a major portion of the treatment session following said initial portion of the treatment session. 36. A system according to claim 35, wherein the initial value is comprised between 20 W and 80 W and wherein the regimen value is comprised between 40 W and 200 W, further wherein the initial value smaller than 80% of the regimen value, optionally smaller than 50% of the regimen value. 37. A system according to any one of the preceding claims 22-36 wherein the treatment session has a total treatment duration comprised between 30 seconds and 30 minutes. 38. A system according to any one of preceding claims 6-37, wherein the controller is configured to automatically stop delivery of power from the ablation energy source and automatically command the flow regulator to stop delivery of conductive fluid when the treatment duration has expired. 39. A system according to any one of the preceding claims 22-38, wherein the controller is configured to control the flow regulator to impose that: a maximum volume of conductive fluid delivered during the treatment session is comprised between 0.3 ml and 60 ml, and/or an average flow rate of conductive fluid maintained during the treatment session is of 0.1 to 15 ml/min, in particular wherein the controller is configured to automatically stop delivery of power from the ablation energy source and/or automatically command the flow regulator to stop delivery of conductive fluid when said maximum of conductive fluid delivered has been reached. 40. The system of any one of the preceding claims in combination with one of claims 18- 21, wherein the set high threshold (T High) is greater than the set low threshold (T Low). 11141. The system of claim 40 in combination with claim 24, wherein the set very high threshold (T Overheat) is greater than the set high threshold (T High). 42. A system according to any one of the preceding claims, wherein the physical property is the temperature of material present at the target region, in particular when this claim also depends upon claim 11 the physical property is the temperature of material present in the volume surrounding the active portion. 43. The system of claims 40 and 42, wherein said set low threshold (T Low) is 60 to 95°C. 44. The system of claim 40 in combination with one of claims 42 or 43, wherein said set high threshold (T High) is from above 75°C to 105°C. 45. The system of claim 41 in combination with one of claims 42 or 43 or 44, wherein said set very high threshold (T Overheat) is between 85 to 115°C. 46. The system of any one of the preceding claims, wherein the ablation energy source is a radiofrequency generator and wherein the controller is configured to control the radiofrequency generator to deliver RF, having a power in a range of 1 to 200 W, in particular comprised between 20 and 200 W, for a duration of 30 seconds to 30 minutes. 47. The system of any one of the preceding claims, wherein the flow regulator comprises: - a pump, optionally a syringe pump or a peristaltic pump or an infusion pump, or- a valve. 48. The system of any one of the preceding claims 7-47, comprising: - a conductive fluid source configured to deliver an hypertonic saline solution; - a fluid port connectable to the conductive fluid source and in fluid communication with the conductive fluid outlet,112 optionally, wherein the hypertonic saline solution includes a reverse phase transition polymer and water, which may transition to higher viscosity when transitioned from below body temperature to body temperature. 49. The system of claim 48, wherein the hypertonic saline solution comprises one or more physiologically acceptable solutes and has a theoretical Osmolarity between 0.8 and 15 Osm/L, calculated according to the formulaOsmolarity = (molarity X ri) Each solute in which n is the number of particles that dissociate from each solute molecule; optionally wherein the hypertonic saline solution is a solution according to any one of claims from 123 to 146. 50. The system of claim 48 or 49, wherein the hypertonic saline solution comprises sodium chloride (NaCl) at a concentration of between 3% to 30% (w/v). 51. The system of any one of the preceding claims 7-50, wherein the flexible shaft is the flexible shaft of an ablation catheter. 52. The system of claim 51, wherein the ablation catheter has a/the fluid port is at a proximal end of the flexible shaft and is in fluid communication with the conductive fluid outlet which is located at the active portion of the flexible shaft. 53. The system of any one of the preceding claims 7-52, wherein the active portion is a distal end portion of the flexible shaft. 54. The system of any one of the preceding claims 7-53 comprising at least one space occluder operative at or proximate to the flexible shaft active portion, in particular at or proximate to the flexible shaft distal end portion. 55. The system of claim 54, wherein the space occluder is one of a deployable balloon, a deployable valve, a deployable stent, or a tapered shaft section. 56. The system of claim 54 or 55, wherein the space occluder comprises a deployable occlusion balloon having a first cross section width of 1 to 30 mm, a length in a range of 5 to 30 mm, and wherein the occlusion balloon is configured to expand to occlude a portion of the airway. 57. The system of claim 56, wherein the first cross section width is at a proximal region of the deployable occlusion balloon, a second cross section width in a range of 1 to 30 mm is at a distal region of the balloon, and a cross section width between the first and second cross section width is less than both the first and second cross section width. 58. The system of claim 56, wherein the first cross section width is at a proximal region of the deployable occlusion balloon, and a second cross section width in a range of 1 to 20 mm and less than the first cross section width is at a distal region of the balloon. 59. The system of any one of the preceding claims 7-58, comprising a tubular sheath or bronchoscope receiving said flexible shaft, wherein at least the active portion, in particular the distal end portion, of the flexible shaft is configured to emerge from the tubular sheath or bronchoscope. 60. The system of claim 59, in combination with any one of claims 54 to 58, wherein the space occluder is carried by the tubular sheath or bronchoscope. 61. The system of any one of claims 54 to 58, wherein the at least one space occluder is directly carried by the flexible shaft or by another shaft. 62. The system of any one of claims 54 to 61, further comprising an inflating lumen extending through the flexible shaft and having a proximal end connectable to a source of a fluid, optionally a source of liquid or a source of gas, and a distal end in fluid communication with inside of the balloon. 63. The system of any one of the preceding claims 7 to 62, wherein the flexible shaft comprises depth markers at least on 5 cm of the proximal region and 5 cm of the distal region. 64. The system of any one of the preceding claims 7 to 62, comprising at least one suction opening at the flexible shaft distal end portion configured to be placed in fluid communication with a vacuum source to aspirate air from a lung volume surrounding the distal end portion of the shaft. 65. The system of claim 64, in combination with any one of claims 54 to 63, wherein the at least one suction opening is positioned distal with respect to the space occluder. 66. The system of claim 64, in combination with any one of claims 54 to 63, comprising an additional space occluder operative at or proximate to the shaft distal end portion, in particular wherein the additional space occluder is one of a deployable balloon, a deployable valve, a deployable stent, and wherein the at least one suction opening is positioned between the space occluder and the additional space occluder. 67. The system of claim 65 or 66, wherein the at least one conductive fluid outlet is positioned distal with respect to the space occluder or between the space occluder and the additional space occluder. 68. The system of any one of claims 54 to 67, wherein the at least one sensor is positioned distal to the space occluder or between the space occluder and the additional space occluder. 69. The system of claim 67, wherein the at least one sensor is carried by the flexible shaft distal end portion and wherein the physical property is of one of temperature, pressure, electric impedance, or electric conductivity of material surrounding the distal end portion of the flexible shaft. 11570. The system of claim 68 or 69, wherein the at least one sensor comprises a first sensor positioned proximal to the ablation element and a second sensor positioned distal to the ablation element. 71. The system of any one of the preceding claims 7-70, wherein the at least one ablation element comprises an ablation element having a rounded distal tip positioned at a distal tip of the flexible shaft. 72. The system of claim 64, in combination with any of claims 54 to 63 and 65 to 71, comprising: a common lumen extending through the flexible shaft and having a proximal end, selectively connectable to at least one of the source of the conductive liquid and the vacuum source, and a distal end, forming a common opening defining said at least one outlet and said at least one suction opening; or a dedicated irrigation lumen and a dedicated air suction lumen, with the irrigation lumen connected to the at least one outlet and extending through the catheter flexible shaft, the irrigation lumen having an inlet port configured to be connected to the source of conductive fluid, and with the air suction lumen connected to the at least one air suction opening and extending through the catheter flexible shaft, the air suction lumen having a suction port configured to be connected to the source of vacuum. 73. The system of any one of the preceding claims 7-72, wherein the flexible shaft has an outer diameter less than or equal to 2 mm. 74. The system of any one of the preceding claims 7-73, wherein at least a portion of the flexible shaft is capable of turning such that a bend in the shaft has a radius of curvature of at least 7 mm. 75. The system of any one of the preceding claims 7-74, wherein the flexible shaft has a length of at least 50 cm. 11676. The system of any one of the preceding claims 7-75, wherein the elongated shaft has a guidewire lumen in the distal end portion configured to receive a guidewire. 77. The system of claims 72 and 76, wherein the suction lumen and guidewire lumen are formed by a common lumen. 78. The system of any one of the preceding claims 6-77, wherein the ablation element comprises at least one electrode characterized by one or more of the following features: total surface area not greater than 120 mm2; diameter in a range of 0.5 to 2 mm; length in a range of 3 to 20 mm. 79. The system of any one of the preceding claims 6-78, wherein the at least one ablation element includes at least two electrodes, and wherein a separation between electrodes is between five to fifteen mm. 80. The system of any one of the preceding claims including an interface component connectable with said at least one sensor and at least communicatively connectable with the controller to transfer to the controller the detected values of said at least one control parameter detected by the sensor. 81. The system of any one of the preceding claims, wherein the controller is configured for: processing said sensed values, and based on one or more of said sensed values, generating at least one output signal which comprises one or more of: o a user identifiable output, optionally the user identifiable output comprising an audible signal, a visual signal or a vibratory signal signaling to the user to deploy at least one space occluder operative at or proximate to the shaft distal end portion,117 o a status output, indicative of the degree of air volume reduction of a lung portion located at a/the catheter distal end portion, o an output command automatically deploying at least one space occluder operative at or proximate to a/the shaft distal end portion, o a temperature output providing an indication of the temperature of material surrounding a/the distal end portion of a/the flexible shaft, o an electric property output providing an indication of the impedance or conductivity of material surrounding a/the distal end portion of the shaft, o a pressure output providing an indication of the pressure of material surrounding a/the distal end portion of a/the flexible shaft. 82. The system of any one of the preceding claims, wherein the controller is configured to: receive signals from the at least one sensor, said sensor being a temperature sensor configured to: monitor temperature at said target region, and control the conductivity or the composition of the conductive fluid delivered through said at least one outlet based on the monitored temperature to maintain the temperature values detected by the temperature sensor within a determined temperature range or above a certain temperature threshold. 83. The system of any one of the preceding claims, wherein controller configured to: receive signals from the at least one sensor, said sensor being a temperature sensor, in particular when this claim depends upon claim 7 said sensor being configured for detecting values of temperature of material surrounding the distal end portion of the flexible shaft, monitor temperature at the target region, adjust the ablation energy power output by the energy source to maintain the temperature values detected by the temperature sensor within a determined temperature range or above a certain temperature threshold. 11884. The system of claim 82 or 83, wherein the determined temperature range is between 60 and 115 °C and the certain temperature threshold is at least 80 °C. 85. The system of any one of the preceding claims 1-84, further comprising a navigation sensor, such as a three-dimensional navigation sensor, or a shape sensor, such as a Fiber Bragg Grating sensor, on at least the distal end region, in particular wherein the navigation sensor is one or more of an electromagnetic sensor, a 3D electromagnetic sensor, shape sensor, FBG sensor, a 3D ultrasound sensor, and an impedance tracking for 3D navigation. 86. The system of any one of the preceding claims 7-85, further comprising a perforation element at a distal tip of the flexible shaft configured to advance through a tumor, wherein the perforation element is selected from a list comprising a needle, a deployable needle, and an RF perforation electrode. 87. The system of any one of the preceding claims 6-87 in combination with claim 54, wherein a distance between the space occluder and the ablation element is in a range of 1 mm to 40 mm. 88. An ablation catheter comprising: a flexible shaft configured to advance through airway passages of a lung; at least one ablation element, positioned at a distal end portion of the flexible shaft and electrically connectable to an ablation energy source; a fluid port connectable to a conductive fluid source; at least one outlet for conductive fluid, the outlet being located at said distal end portion and being in fluid communication with the fluid port. 89. The catheter of claim 88 further comprising at least one space occluder operative at or proximate to the shaft distal end portion, in particular wherein the space occluder is one of a tapered shaft section, a deployable balloon, a deployable valve, or a deployable stent. 11990. The catheter of claim 89, wherein the occluder comprises a deployable occlusion balloon having a first cross section width of 1 to 30 mm, a length in a range of 5 to 30 mm, and wherein the occlusion balloon is configured to expand to occlude a portion of the airway. 91. The catheter of claim 90, wherein the first cross section width is at a proximal region of the deployable occlusion balloon, a second cross section width in a range of 1 to 30 mm is at a distal region of the balloon, and a cross section width between the first and second cross section width is less than both the first and second cross section width. 92. The catheter of claim 90, wherein the first cross section width is at a proximal region of the deployable occlusion balloon, and a second cross section width in a range of 1 to 20 mm and less than the first cross section width is at a distal region of the balloon. 93. The catheter of claim 88 to 92, comprising a tubular sheath or bronchoscope receiving said shaft, wherein at least the distal end portion of the flexible shaft is configured to emerge from the tubular sheath or bronchoscope. 94. The catheter of claim 93 in combination with claim 90, wherein the space occluder is carried by the tubular sheath or bronchoscope or wherein the at least one space occluder is directly carried by the shaft or by a different shaft. 95. The catheter of claim 88 to 94 in combination with claim 90, further comprising an inflating lumen extending through the flexible shaft and having a proximal end connectable to a source of a fluid, optionally a source of liquid or a source of gas, and a distal end in fluid communication with inside of the balloon. 96. The catheter of any of claims 88 to 95, wherein the flexible shaft comprises depth markers at least on 5 cm of the proximal region and 5 cm of the distal region. 12097. The catheter of any of claims 88 to 96, comprising at least one suction opening at the shaft distal end portion configured to be placed in fluid communication with a vacuum source to aspirate air from a lung volume surrounding the distal end portion of the shaft. 98. The catheter of claim 97 in combination with claim 90, wherein the at least one suction opening is positioned distal with respect to the space occluder. 99. The catheter of claim 98 comprising an additional sp<TRUNCATED> | A robotically delivered catheter system for treating lung tumors that uses a flow regulator between a conductive-fluid source and an outlet at the target. Controlled fluid delivery (rate or bolus) supports the ablation procedure. |  | "at least one sensor being configured for detecting values taken by at least one control parameter representative of a physical property, wherein the physical property is one of temperature (T), pressure (p), electric impedance (Z), or electric conductivity (C) of material present at or in proximity of the target region of lung tissue" (Claim 1) "controlling the flow regulator in a high delivery mode… controlling the flow regulator in a low delivery mode… executing a control cycle" (Claim 1) "wherein said set low threshold (T Low) is 60 to 95°C… said set high threshold (T High) is from above 75°C to 105°C… said set very high threshold (T Overheat) is between 85 to 115°C" (Claims 43–45) "a navigation sensor, such as a three-dimensional navigation sensor, or a shape sensor, such as a Fiber Bragg Grating sensor, on at least the distal end region" (Claim 85) |  |  |  |  | 1 | [5 empty] | 1 | [15 empty] | 1 |  |  |  |  | ZIDAN MEDICAL INC | ZIDAN MEDICAL INC | 2021-09-27 | 2021 | 2022-05-12 | 2022 | 2020-09-28 | 2020 | WO2021US52121 | AU2021350113A1¦CA3192769A1¦CN116744872A¦EP4216857A2¦JP2023544302A¦KR20230097033A¦US20230355300A1¦WO2022067146A2